## IQCP Statement of Record
Individualized Quality Control Plan (IQCP): Blood Gas Systems: RAPIDPoint® 400/405/500 and RAPIDLab® 1200 Series Systems
Quality Control Plan (QCP): The IQCP is following exactly what is written in manufacturer’s instructions for doing the test, and in addition the laboratory follows their practices, resources, and procedures as stated in the lab’s manual assuring accurate and reliable test results at all times. The laboratory relies on internal built in controls of the test system assuring the test system is properly working. For quality control, the laboratory is following the manufacturer’s recommended quality control procedure by testing external quality control material for each shipment of a new kit lot number. The laboratory has quality control data supporting the manufacturer’s QC recommendation.
Quality Assessment (QA): All IQCP QA monitoring is part of the lab’s overall QA plan and is an ongoing review process for the IQCP. The lab has a review system for the ongoing monitoring of the effectiveness of the lab’s IQCP. The monitoring includes, but not limited to specimen, environment, testing personnel, reagent, and test system. In addition, if there is a complaint on a test result and or unacceptable QC and proficiency test result, the lab investigates and takes corrective action. It may be necessary to change the IQCP.
Laboratory Director Signature: ________________________________________ Date: __________________________
I, the Laboratory Director delegate to: _______________________________________in preparation of this document which I am ultimately responsible for. Date: _______________________________

## Instructions for Use
Individualized Quality Control Plan (IQCP): Blood Gas Systems: RAPIDPoint® 400/405/500 and RAPIDLab® 1200 Series Systems
Instruction for Use
Phase | Defined as Pre-Analytical, Analytical and Post-Analytical
Process Step | Defined in test instruction for use
Identified Risk - Source of Potential Risk | Defined in test instruction for use and institute practices
Frequency of Occurrence | This should be defined by individual site risk assessments for previously documented events
Severity of Harm | Can use risk asssesment guide to determine category
Risk Mitigation: Manage/Prevent Errors from Occurring | This is provided in instructions for use, operators guides, current SOP's and various regulatory instructions
Risk Level Acceptable: Yes or No | Answer according to level of risk being mitigated
Reference | References from manufacturer, regulatory bodies and individual institute SOP's

## Specimen
Individualized Quality Control Plan (IQCP): Blood Gas Systems: RAPIDPoint® 400/405/500 and RAPIDLab® 1200 Series Systems
Specimen
Phase | Process Step | Identified Risk - Source of Potential Risk | Frequency of Occurrence | Severity of Harm | Risk Mitigation: Manage/Prevent Errors from Occurring | Risk Level Acceptable: Yes or No | Reference
Pre-Analytical
Pre-Analytical | Patient Identification | Not following hospital protocol to properly verify the patient identity and to properly record patient information. |  |  | Operators should be trained in the hospital's patient identification protocol and to use the hospital system barcode to ensure proper data entry of patient identification number and to leverage custom-definable data-entry fields, if applicable. Follow established SOP XXX.XXX. |  | CLSI document # C46-A2
Pre-Analytical | Specimen Collection | Improper selection and preparation of a suitable site for the withdrawal of the whole blood sample; |  |  | Operators should be trained to follow hospital SOPs for the site preparation and the anaerobic collection of whole blood specimens for blood gas analysis. The whole blood sample should be kept at room temperature and analyzed within 30 minutes of collection. Follow established SOP XXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Specimen Collection Devices | Optimal analyzer performance requires the use of properly heparinized whole blood samples collected in either syringes or capillaries. The use of blood samples in sampling devices without the appropriate anticoagulant will result in clots and system fluidic errors, and has also been shown to potentially impact sensor performance. |  |  | The use of appropriate levels of electrolyte-balanced Lithium Heparin as anticoagulant, for arterial puncture or line draws as well as capillary samples, will minimize clot formation and will not adversely affect the system sensors. Follow established SOP XXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Specimen Labeling | Improper labeling of the specimen for positive identification. |  |  | Operators should be trained to follow hospital instructions/directions for printing and applying sample barcode labels to specimen collection devices to ensure proper/positive patient-to-sample identification. Follow established SOP XXX.XXX. |  | CLSI document # C46-A2
Pre-Analytical | Specimen Handling | Mis-handling of the specimen such as, not removing the air bubble, inadequate mixing, delay in analysis, and hemolysis, will adversely affect patient results. |  |  | Operators should be trained in the correct procedures to eliminate the air bubble, appropriate mixing of the specimen both immediately after collection and prior to analysis, and analysis within 30 minutes of collection. Follow established SOP XXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Specimen Mixing | Not properly mixing the specimen immediately after collection and prior to introducing the specimen into the analyzer for measurement, may result in insufficient heparinzation of the blood and lack of specimen homogeneity. |  |  | Operators should be trained in the proper mixing technique for both syringe and capillary specimens. Follow established SOP XXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Specimen Transport | Improper storage and handling of specimen during transport from the patient to the analyzer, in terms of elapsed time and storage temperature. |  |  | Ensure that operators are trained appropriately on the time allowance from collection to analysis for blood gas testing. Procedures for proper transport should be followed. Follow established SOP XXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Infection Prevention | Improper handling of whole blood specimen. |  |  | Operators should be trained on proper sample handling technique to minimize exposure to infection. Follow hospital site instructions for proper sample handling as per SOP XXX.XXX. Operators should be trained to follow the system cleaning instructions found in the operator's manual. Follow established SOP XXX.XXX for infection prevention. |  | CLSI documents # GP05-A3,M29-A3, GP17-A3
Analytical
Analytical | Specimen Integrity: Lipemia | Testing whole blood specimens from hyperlipemic patients. |  |  | RAPIDPoint/RAPIDLab 1200 systems are not affected with lipid concentrations 5% or less. Follow SOP XXXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Analytical | Specimen Integrity: Hemolysis | Testing whole blood specimens that are hemolyzed. |  |  | There is no direct test for hemolysis in whole blood samples. Operators should be trained on pre-analytical steps to avoid hemolysis and to review the K+ results which can be falsely elevated in the event of hemolysis. Follow SOP XXXX.XXX to minimize causing hemolysis during sampling. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Analytical | Specimen Integrity: Interfering Substances | Testing whole blood specimens that may contain known interferents. |  |  | Operators should be familiar with the RAPIDPoint/RAPIDLab 1200 list of interfering substances. The systems perform auto-checks designed to identify select interferents and flag the sample with the appropriate message. Follow SOP XXXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Analytical | Specimen Integrity: Clotting | Testing whole blood specimens that may contain clots. |  |  | Operators should be trained on pre-analytical procedures to minimize/eliminate clot formation. In the event of aspirating a clot(s), the sample port is designed to trap the majority of clots and avoid the likelihood of the clot from getting into the cartridge. Follow manufacturer's instructions to replace the sample port. Systems have clot detection checks and identification. Follow SOP XXXX.XXX |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Analytical | Specimen Integrity: Incorrect collection device | Use of improper syringe or capillary sampling device. |  |  | The use of appropriate levels of electrolyte-balanced Lithium Heparin as anticoagulant, for arterial puncture or line draws as well as capillary samples, will minimize clot formation and will not adversely affect the system sensors. Follow SOP XXXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Analytical | Specimen Integrity: Air Bubbles | Not properly expelling the air bubble from the syringe or capping the capillary tubes to minimize exposure to air. |  |  | Operators should be trained in the correct procedures to eliminate the air bubble, appropriate mixing of the specimen both immediately after collection and prior to analysis, and analysis within 30 minutes of collection. Systems have bubble detection and identification checks. Follow established SOP XXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Analytical | Specimen Integrity: Inadequate Volume | Not filling the syringe/capillary to the appropriate patient sample fill volume. |  |  | Operator should be trained on the appropriate fill volumes based on sampling device size. RAPIDPoint system detects short samples and aborts the analysis. RAPIDLab 1200 system detects short samples and shifts to microsample mode prior to analysis. Follow SOP XXXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Post-Analytical
Post-Analytical | Specimen Disposal | Improper disposal of specimen after testing. |  |  | Operator should be trained on the hospital policy for the disposal of biohazard waste and treat the syringes and capillaries as such. Follow hospital instructions for proper disposal of biohazard waste as per SOP XXX.XXX. |  | CLSI documents # GP05-A3,M29-A3, GP17-A3
Post-Analytical | Infection Prevention | Improper handling of whole blood specimen after analysis. |  |  | All whole blood and reagent waste products are fully contained within the RAPIDPoint waste/wash cartridge to minimize operator exposure; The RAPIDLab 1200 waste bottle contains the whole blood and reagent waste. Follow hospital instructions for proper disposal of biohazard waste as per SOP XXX.XXX. Operators should be trained to follow the system cleaning instructions found in the operator's manual. Follow established SOP XXX.XXX for infection prevention. |  | CLSI documents # GP05-A3,M29-A3, GP17-A3
***Disclaimer: This template is devised to be utilized as a sample of an IQCP document for the Siemens Blood Gas products. Each facility needs to complete these categories based on their own risk assessment which should include institute SOP’s, local, state and all other regulatory agency requirements along with manufacturer’s instructions for use.

## Testing Personnel
Individualized Quality Control Plan (IQCP): Blood Gas Systems: RAPIDPoint® 400/405/500 and RAPIDLab® 1200 Series Systems
Operator
Phase | Process Step | Identified Risk - Source of Potential Risk | Frequency of Occurrence | Severity of Harm | Risk Mitigation: Manage/Prevent Errors from Occurring | Risk Level Acceptable: Yes or No | Reference
Pre-Analytical
Pre-Analytical | Patient Identification | Not following hospital protocol to properly verify the patient identity and to properly record patient information. |  |  | Operators should be trained in the hospital's patient identification protocol and to use the hospital system barcode to ensure proper data entry of patient identification number and to leverage custom-definable data-entry fields, if applicable. Follow established SOP XXX.XXX. |  | CLSI document # C46-A2
Pre-Analytical | Specimen Collection | Improper selection and preparation of a suitable site for the withdrawal of the whole blood sample; |  |  | Operators should be trained to follow hospital SOPs for the site preparation and the anaerobic collection of whole blood specimens for blood gas analysis. The whole blood sample should be kept at room temperature and analyzed within 30 minutes of collection. Follow established SOP XXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Specimen Collection Devices | Optimal analyzer performance requires the use of properly heparinized whole blood samples collected in either syringes or capillaries. The use of blood samples in sampling devices without the appropriate anticoagulant will result in clots and system fluidic errors, and has also been shown to potentially impact sensor performance. |  |  | The use of appropriate levels of electrolyte-balanced Lithium Heparin as anticoagulant, for arterial puncture or line draws as well as capillary samples, will minimize clot formation and will not adversely affect the system sensors. Follow established SOP XXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Specimen Labeling | Improper labeling of the specimen for positive identification. |  |  | Operators should be trained to follow hospital instructions/directions for printing and applying sample barcode labels to specimen collection devices to ensure proper/positive patient-to-sample identification. Follow established SOP XXX.XXX. |  | CLSI document # C46-A2
Pre-Analytical | Specimen Handling | Mis-handling of the specimen such as, not removing the air bubble, inadequate mixing, delay in analysis, and hemolysis, will adversely affect patient results. |  |  | Operators should be trained in the correct procedures to eliminate the air bubble, appropriate mixing of the specimen both immediately after collection and prior to analysis, and analysis within 30 minutes of collection. Follow established SOP XXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Specimen Mixing | Not properly mixing the specimen immediately after collection and prior to introducing the specimen into the analyzer for measurement, may result in insufficient heparinzation of the blood and lack of specimen homogeneity. |  |  | Operators should be trained in the proper mixing technique for both syringe and capillary specimens. Follow established SOP XXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Specimen Transport | Improper storage and handling of specimen during transport from the patient to the analyzer, in terms of elapsed time and storage temperature. |  |  | Ensure that operators are trained appropriately on the time allowance from collection to analysis for blood gas testing. Procedures for proper transport should be followed. Follow established SOP XXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Infection Prevention | Improper handling of whole blood specimen. |  |  | Operators should be trained on proper sample handling technique to minimize exposure to infection. Follow hospital site instructions for proper sample handling as per SOP XXX.XXX. Operators should be trained to follow the system cleaning instructions found in the operator's manual. Follow established SOP XXX.XXX for infection prevention. |  | CLSI documents # GP05-A3,M29-A3, GP17-A3
Analytical
Analytical | Specimen Integrity: Lipemia | Testing whole blood specimens from hyperlipemic patients. |  |  | RAPIDPoint/RAPIDLab 1200 systems are not affected with lipid concentrations 5% or less. Follow SOP XXXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Analytical | Specimen Integrity: Hemolysis | Testing whole blood specimens that are hemolyzed. |  |  | There is no direct test for hemolysis in whole blood samples. Operators should be trained on pre-analytical steps to avoid hemolysis and to review the K+ results which can be falsely elevated in the event of hemolysis. Follow SOP XXXX.XXX to minimize causing hemolysis during sampling. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Analytical | Specimen Integrity: Interfering Substances | Testing whole blood specimens that may contain known interferents. |  |  | Operators should be familiar with the RAPIDPoint/RAPIDLab 1200 list of interfering substances. The systems perform auto-checks designed to identify select interferents and flag the sample with the appropriate message. Follow SOP XXXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Analytical | Specimen Integrity: Clotting | Testing whole blood specimens that may contain clots. |  |  | Operators should be trained on pre-analytical procedures to minimize/eliminate clot formation. In the event of aspirating a clot(s), the sample port is designed to trap the majority of clots and avoid the likelihood of the clot from getting into the cartridge. Follow manufacturer's instructions to replace the sample port. Systems have clot detection checks and identification. Follow SOP XXXX.XXX |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Analytical | Specimen Integrity: Incorrect collection device | Use of improper syringe or capillary sampling device. |  |  | The use of appropriate levels of electrolyte-balanced Lithium Heparin as anticoagulant, for arterial puncture or line draws as well as capillary samples, will minimize clot formation and will not adversely affect the system sensors. Follow SOP XXXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Analytical | Specimen Integrity: Air Bubbles | Not properly expelling the air bubble from the syringe or capping the capillary tubes to minimize exposure to air. |  |  | Operators should be trained in the correct procedures to eliminate the air bubble, appropriate mixing of the specimen both immediately after collection and prior to analysis, and analysis within 30 minutes of collection. Systems have bubble detection and identification checks. Follow established SOP XXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Analytical | Specimen Integrity: Inadequate Volume | Not filling the syringe/capillary to the appropriate patient sample fill volume. |  |  | Operator should be trained on the appropriate fill volumes based on sampling device size. RAPIDPoint system detects short samples and aborts the analysis. RAPIDLab 1200 system detects short samples and shifts to microsample mode prior to analysis. Follow SOP XXXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Post-Analytical
Post-Analytical | Specimen Disposal | Improper disposal of specimen after testing. |  |  | Operator should be trained on the hospital policy for the disposal of biohazard waste and treat the syringes and capillaries as such. Follow hospital instructions for proper disposal of biohazard waste as per SOP XXX.XXX. |  | CLSI documents # GP05-A3,M29-A3, GP17-A3
Post-Analytical | Infection Prevention | Improper handling of whole blood specimen after analysis. |  |  | All whole blood and reagent waste products are fully contained within the RAPIDPoint waste/wash cartridge to minimize operator exposure; The RAPIDLab 1200 waste bottle contains the whole blood and reagent waste. Follow hospital instructions for proper disposal of biohazard waste as per SOP XXX.XXX. Operators should be trained to follow the system cleaning instructions found in the operator's manual. Follow established SOP XXX.XXX for infection prevention. |  | CLSI documents # GP05-A3,M29-A3, GP17-A3
Pre-Analytical/Analytical/Post-Analytical
Pre-Analytical, Analytical, Post-Analytical | Operator Training | Improper or inadequate operator training. |  |  | The systems can be configured to manage operator access and to lock-out unauthorized personnel that are not adequately trained. If interfaced to RAPIDComm Data Management System, operator competencies can be managed via RAPIDComm and operators with expired training can be locked out of system access and operation. PEP and POCkit are available training tools. On-board system videos help to supplement training and optimize system use. Operators should be trained via SOP XXXX.XXX. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical, Analytical, Post-Analytical | Operator Competency Assessment | Operators may have expired training records or may have not been re-certified. Proper testing technique, sample handling, reagent handling, system maintenance may not be accurately followed. |  |  | The systems can be configured to manage operator access and to lock-out unauthorized personnel that are not adequately trained. If interfaced to RAPIDComm Data Management System, operator competencies can be managed via RAPIDComm and operators with expired training can be locked out of system access and operation. PEP and POCkit are available training tools. On-board system videos help to supplement training and optimize system use. Operators should be trained via SOP XXXX.XXX. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical, Analytical, Post-Analytical | Proficiency Testing | Errors in proficiency testing performance. |  |  | Operators should perform and review proficiency testing samples and reports. Follow proficiency agency instructions. Operators should be trained via SOP XXXX.XXX. |  | Proficiency agency instructions RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical, Analytical, Post-Analytical | Un-trained Operator Access | Un-trained operators may be running the system or unauthorized operators may have access to the system. |  |  | The systems can be configured to manage operator access and to lock-out unauthorized personnel that are not adequately trained. Operators should be trained via SOP XXXX.XXX. |  | CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical, Analytical, Post-Analytical | Adequate Staffing | Staffing levels may be inadequate to perform testing without errors. Possibly, untrained operators may be performing the testing. |  |  | The blood gas systems are designed for ease-of-use and require minimal operator interaction, allowing for more time with the patient. Calibrations and AQC routines are automatically performed. Sample aspiration is hands-free and automatic. Both systems (RAPIDPoint/RAPIDLab 1200) require minimal to no maintenance. Operators should be training via SOP XXXX.XXX. |  | Siemens: Minimizing Pre-Analytical Error in Blood Gas Testing Document (A91DX-100341-GC1-4A00), CLSI document # C46-A2 RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical, Analytical, Post-Analytical | Infection Control | Improper handling of specimen, reagents or test system. |  |  | Operators should be trained on proper sample handling technique to minimize exposure to infection. Follow hospital site instructions for proper sample handling as per SOP XXX.XXX. Operators should be trained to follow the system cleaning instructions found in the operator's manual. Follow established SOP XXX.XXX for infection prevention. |  | CLSI documents # GP05-A3,M29-A3, GP17-A3
***Disclaimer: This template is devised to be utilized as a sample of an IQCP document for the Siemens Blood Gas products. Each facility needs to complete these categories based on their own risk assessment which should include institute SOP’s, local, state and all other regulatory agency requirements along with manufacturer’s instructions for use.

## Reagent
Individualized Quality Control Plan (IQCP): Blood Gas Systems: RAPIDPoint® 400/405/500 and RAPIDLab® 1200 Series Systems
Reagent
Phase | Process Step | Identified Risk - Source of Potential Risk | Frequency of Occurrence | Severity of Harm | Risk Mitigation: Manage/Prevent Errors from Occurring | Risk Level Acceptable: Yes or No | Reference
Pre-Analytical
Pre-Analytical | Reagent Storage | Improper storage of various cartridges (refrigerated vs room temperature) before installing the cartridges on the system may affect assay performance. Improper storage of select sensors. |  |  | Only 3 cartridges require refrigeration: RAPIDPoint measurement cartridge, RAPIDLab 1200 reagent cartridge and the AQC cartridge. Only 2 RAPIDLab 1200 sensors need refrigeration: Glucose and lactate. Ensure that operators are aware of the appropriate storage conditions by reviewing product labeling. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual and Product Labeling
Pre-Analytical | Reagent Expiration Date | Reagent used past expiration date or use-by date. |  |  | The cartridge install date and expiration (by date or number of tests) is monitored by the instrument and the operator is notified when the cartridges are depleted. If cartridges are expired or depleted, no testing can be done. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual and Product Labeling
Pre-Analytical | Reagent Contamination | Improper use of damaged/contaminated cartridges. |  |  | Ensure operators are trained to not install visibly damaged cartridges. Calibration routines are performed automatically and parameters are disabled if they exceed performance limits. All cartridges are fully self-contained to minimize reagent contamination. Operators should be trained to discard the cartridge if it is visibly damaged. Follow established SOP XXX.XXX for maintaining the system. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual and Product Labeling
Pre-Analytical | Reagent Preparation | N/A - All reagents/cartridges are ready-to-use requiring no reagent preparation. Exception is the deproteinizer that is used on the RAPIDLab 1200, which needs to be reconstituted prior to use. |  |  | System cartridges do not require any reagent preparation. Operators should be trained on proper handling and installation of the cartridges. Ensure operators follow deproteinizer reconstitution direction as described with the packaging. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual and Product Labeling
Pre-Analytical | Reagent Handling | Improper storage of measurement or reagent cartridges before installing on the analyzer. |  |  | Only 3 cartridges require refrigeration: RAPIDPoint measurement cartridge, RAPIDLab 1200 reagent cartridge and the AQC cartridge. Only 2 RAPIDLab 1200 sensors need refrigeration: Glucose and lactate. Ensure that operators are aware of the appropriate storage conditions by reviewing product labeling. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual and Product Labeling
Pre-Analytical | Lot-to-lot variation of AQC | Not validating/verifying new AQC cartridges, when received. |  |  | AQC value assignments are performed against a master lot and programmed into the cartridge, eliminating batch change-over issues. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual, AQC White Paper(A91DX-POC-151051-GC1-4A00) and AQC Customer Presentation (A91DX-POC-150527-GC1-4A00)
Pre-Analytical | Infection Prevention | Improper handling of the system cartridges or individual sensors. |  |  | Operators should be trained on the proper handling and disposal of the cartridges/sensors to minimize exposure to infection. All cartridges/sensors should be treated as bio-hazard. Follow hospital site instructions for proper bio-hazard handling as per SOP XXX.XXX. Follow established SOP XXX.XXX for infection prevention. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual CLSI documents # GP05-A3,M29-A3, GP17-A3
Analytical
Analytical | Reagent (Measurement/Reagent/Wash/Waste/AQC Cartridge) Installation | Measurement/reagent cartridges or wash/waste cartridges or AQC cartridge are improperly installed on the system. |  |  | Operators should be trained to follow the proper procedure for installing the cartridges. On-board system video guidance helps to supplement training and optimize system use. Built-in fault detection and identification. Follow SOP XXXX.XXX for appropriate operator training. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual and Product Labeling
Analytical | Quality Control | Improper frequency of running and monitoring QC. |  |  | Develop procedures for analyzing the number of QC samples and levels of QC material for each day of testing. Ensure that your procedures adhere to federal, state, and local regulatory agency requirements for performing quality control. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual, AQC White Paper(A91DX-POC-151051-GC1-4A00) and AQC Customer Presentation (A91DX-POC-150527-GC1-4A00)
Post-Analytical
Post-Analytical | Result Interpretation | Improperly recording results. |  |  | Results can be set to automatically transmit to middleware/LIS/HIS after meeting pre-determined acceptance criteria. Follow hospital SOP XXXX.XXX for result acceptance and reporting. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Post-Analytical | Cartridge Removal | Improper removal of measurement/reagent cartridge from the system. |  |  | Operators should follow manufacturer's direction for proper cartridge removal as specified in the operators manual to avoid injury or possible damage to the system. Video guidance sequences are provided to aid in the removal of the cartridges. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Post-Analytical | Cartridge Disposal | Improper disposal of the measurement/reagent cartridge. |  |  | Operators should be trained on the proper handling and disposal of the cartridges/sensors to minimize exposure to infection. All cartridges/sensors should be treated as bio-hazard. Follow hospital site instructions for proper bio-hazard handling as per SOP XXX.XXX. Follow established SOP XXX.XXX for infection prevention. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual CLSI documents # GP05-A3,M29-A3, GP17-A3
Post-Analytical | Infection prevention | Improper handling of the system cartridges or individual sensors. |  |  | Operators should be trained on the proper handling and disposal of the cartridges/sensors to minimize exposure to infection. All cartridges/sensors should be treated as bio-hazard. Follow hospital site instructions for proper bio-hazard handling as per SOP XXX.XXX. Follow established SOP XXX.XXX for infection prevention. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual CLSI documents # GP05-A3,M29-A3, GP17-A3
***Disclaimer: This template is devised to be utilized as a sample of an IQCP document for the Siemens Blood Gas products. Each facility needs to complete these categories based on their own risk assessment which should include institute SOP’s, local, state and all other regulatory agency requirements along with manufacturer’s instructions for use.

## Environment
Individualized Quality Control Plan (IQCP): Blood Gas Systems: RAPIDPoint® 400/405/500 and RAPIDLab® 1200 Series Systems
Environment
Phase | Process Step | Identified Risk - Source of Potential Risk | Frequency of Occurrence | Severity of Harm | Risk Mitigation: Manage/Prevent Errors from Occurring | Risk Level Acceptable: Yes or No | Reference
Pre-Analytical
Pre-Analytical | Temperature: Instrument | Laboratory temperature is outside the operating temperature range of the system. |  |  | Operators should be trained to ensure that the laboratory testing area temperature is adequate to meet the systems' operating temperature as specified in the operator manual. System performs operating temperature checks and patient results will not be generated if the operating temperature is outside of the acceptable limit. Follow SOP XXXX.XXX for monitoring and recording ambient laboratory temperature. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Temperature: Cartridges | Improper storage of measurement/reagent/wash waste/AQC cartridges or metabolite biosensors before installing on the system. |  |  | Only 3 cartridges require refrigeration: RAPIDPoint measurement cartridge, RAPIDLab 1200 reagent cartridge and the AQC cartridge. Only 2 RAPIDLab 1200 sensors need refrigeration: Glucose and lactate. Ensure that operators are aware of the appropriate storage conditions. Systems have internal checks that detect when optimal temperature conditions in the measurement chamber are not reached and will trigger an error when temperature is out of range, Follow SOP XXXX.XXX for monitoring and recording laboratory refrigerator and testing area temperature. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Ventilation: Instrument | May cause temperature errors if airflow is restricted. Air filter not changed as per maintenance recommendation. |  |  | Ensure that there is adequate space around the system for sufficient airflow and that the fan filter is changed at the maintenance interval as described in the operators manual. On-board thermal control/checks detects over-heating. Follow SOP XXXX.XX |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Humidity: Instrument/Cartridges | Humidity is outside of the system testing requirements and may affect analyzer and sensor performance. |  |  | Operators should be trained to ensure the system is used within the relative humidity tolerances of 10% to 90% (non-condensing) as recommended in the operators manual. Follow SOP XXXX.XXX for monitoring and recording humidity levels in the laboratory. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Dust | Dust in the air filter can affect airflow and system performance. |  |  | Ensure that operators clean or replace the air filter at regularly scheduled intervals as described in the operator manual. The self-contained cartridges and closed measurement chamber prevent dust from getting into the measurement system. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Barometric Pressure | Incorrect calibration/recording of the barometric pressure can result in incorrect calculations of some derived parameters. |  |  | Ensure that the systems operate within acceptable ambient operating barometric pressure (523 - 800 mmHg) and that the internal barometer on the RAPIDLab 1200 is calibrated with the laboratory barometric pressure. Follow SOP XXXX.XXX for appropriate operator training on the calibration of the barometric pressure. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Utilities (electrical power disruption) | Power interruption can cause power surges and affect system operation. |  |  | System should always be plugged into a grounded outlet and the use of an UPS will allow for continuous power and system operation. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Space (around instrument) | May cause temperature errors if airflow is restricted. Air filter not changed as per maintenance recommendation. |  |  | Ensure that there is adequate space around the system for sufficient airflow and that the fan filter is changed at the maintenance interval as described in the operators manual. Follow SOP XXXX.XX |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Analytical
Analytical | Temperature: Analyzer | Laboratory temperature is outside the operating temperature range of the system. |  |  | Operators should be trained to ensure that the laboratory testing area temperature is adequate to meet the systems' operating temperature as specified in the operator manual. System performs operating temperature checks and patient results will not be generated if the operating temperature is outside of the acceptable limit. Follow SOP XXXX.XXX for monitoring and recording ambient laboratory temperature. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Post-Analytical
Post-Analytical | Infection prevention | Improper handling of specimen, reagents or test system. |  |  | Operators should be trained on proper sample handling technique to minimize exposure to infection. Follow hospital site instructions for proper sample handling as per SOP XXX.XXX. Operators should be trained to follow the system cleaning instructions found in the operator's manual. Follow established SOP XXX.XXX for infection prevention. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual CLSI documents # GP05-A3,M29-A3, GP17-A3
***Disclaimer: This template is devised to be utilized as a sample of an IQCP document for the Siemens Blood Gas products. Each facility needs to complete these categories based on their own risk assessment which should include institute SOP’s, local, state and all other regulatory agency requirements along with manufacturer’s instructions for use.

## Test System
Individualized Quality Control Plan (IQCP): Blood Gas Systems: RAPIDPoint® 400/405/500 and RAPIDLab® 1200 Series Systems
Test System
Phase | Process Step | Identified Risk - Source of Potential Risk | Frequency of Occurrence | Severity of Harm | Risk Mitigation: Manage/Prevent Errors from Occurring | Risk Level Acceptable: Yes or No | Reference
Pre-Analytical
Pre-Analytical | Proper Specimen Sampling | Mis-handling of the specimen such as, not removing the air bubble, inadequate mixing, delay in analysis, and hemolysis, will adversely affect patient results. |  |  | Operators should be trained in the correct procedures to eliminate the air bubble, appropriate mixing of the specimen both immediately after collection and prior to analysis, and analysis within 30 minutes of collection. Follow established SOP XXX.XXX. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Calibration Issues | Failure to calibrate successfully will results in delays in test reporting. |  |  | 1-point, 2-point and full calibration sequences occur automatically, and if unsuccessful, affected parameter is deactivated and patient results cannot be reported. Calibrations may be manually initiated, if necessary. Follow SOP XXXX. XXX for calibration sequences. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Software | Incorrect version of software. |  |  | Ensure operator installs most current version of software according to manufacturer's updates.. Follow SOP XXXX.XXX |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Barcode Data Entry | Transcription error of manually-entered data such as patient information, control data and other manually entered data. |  |  | Systems support barcode data entry to ensure proper sample identification, positive patient ID and operator ID and QC mean and ranges for all parameters. Operators should be trained on the barcode capabilities and the symbologies supported by the analyzer listed in the operators manual. Follow SOP XXXX.XXX |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Quality Control | Improper installation of AQC cartridge. Improper sample handling of the ampuled quality control materials. Improper storage of the QC materials and Calibration Verification Material (CVM). |  |  | Operators should be trained to properly install the AQC cartridge and perform ampuled QC as per the Instructions for Use or via hospital SOP XXXX.XXX. System can be set-up to run AQC automatically according to pre-set schedule and deactivate parameter if QC fails. Follow hospital SOP XXXX.XXX. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Pre-Analytical | Maintenance | Improper maintenance of the system. |  |  | System performs continuous checks, including: mechanical and electronic checks and generates error codes when limits are not met and prevents testing of samples under certain conditions. Operators should be trained to review error codes and instructions for next steps in the system operator manual. Operators should be trained to perform maintenance tasks. System can be set-up to remind users when to perform maintenance. Follow hospital SOP XXXX.XXX. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Analytical
Analytical | Connectivity to LIS/HIS | Unable to transmit data to middleware or LIS/HIS due to improper system connectivity set-up or lost connection |  |  | System automatically re-transmits results in holding queue due to connectivity drop. System generates error codes related to software and hardware and/or server errors. Operators should be trained to back-up and restore database, if necessary. System settings, patient and QC data can be backed-up and restored. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Analytical | Instrument Failure | Improper maintenance of the system. Improper use of the system. |  |  | System performs continuous checks, including: mechanical and electronic checks and generates error codes when limits are not met and prevents testing of samples under certain conditions. Operators should be trained to review error codes and instructions for next steps in the system operator manual. Operators should be trained to perform maintenance tasks. System can be set-up to remind users when to perform maintenance. Follow hospital SOP XXXX.XXX. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
Post-Analytical
Post-Analytical | Infection Prevention | Improper handling of specimen, reagents or test system. |  |  | Operators should be trained on proper sample handling technique to minimize exposure to infection. Follow hospital site instructions for proper sample handling as per SOP XXX.XXX. Operators should be trained to follow the system cleaning instructions found in the operator's manual. Follow established SOP XXX.XXX for infection prevention. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual CLSI documents # GP05-A3,M29-A3, GP17-A3
Post-Analytical | Result Reporting | Inaccurate transcription of results to patient record. |  |  | System can be connected (serially or networked) to transmit results directly to middleware data management system and/or LIS/HIS. Follow hospital SOP XXXX.XXX for data recording. |  | RAPIDPOINT 400/405/500 RAPIDLab 1200 Product Operator's Manual
***Disclaimer: This template is devised to be utilized as a sample of an IQCP document for the Siemens Blood Gas products. Each facility needs to complete these categories based on their own risk assessment which should include institute SOP’s, local, state and all other regulatory agency requirements along with manufacturer’s instructions for use.

## Risk Assessment Guide
Risk Assessment Guide* Hospital/Test System residual risk
Probability of Harm
Frequent = once/week
Probable = once/month
Occasional = once/year
Remote = once/few years
Improbable = once/lifetime
Severity of Harm
Negligible = inconvenient or temporary discomfort
Minor = temporary injury or impairment not requiring professional medical intervention
Serious = injury or impairment requiring professional medical intervention
Critical = permanent impairment or life-threatening injury
Catastrophic = patient death
Probability of Harm | Negligible | Minor | Serious | Critical | Catastrophic
Frequent | Unacceptable | Unacceptable | Unacceptable | Unacceptable | Unacceptable
Probable | Acceptable | Unacceptable | Unacceptable | Unacceptable | Unacceptable
Occasional | Acceptable | Acceptable | Perform Risk-Benefit Analysis | Unacceptable | Unacceptable
Remote | Acceptable | Acceptable | Perform Risk-Benefit Analysis | Perform Risk-Benefit Analysis | Unacceptable
Improbable | Acceptable | Acceptable | Perform Risk-Benefit Analysis | Perform Risk-Benefit Analysis | Perform Risk-Benefit Analysis
*ISO 14971: Medical devices- Application of risk management to medical devices. www.iso.org
NOTE: This is an example only. Risk Gradient tables must be established in adherence to site-specific requirements and risk criteria.