## IQCP Statement of Record
Individualized Quality Control Plan (IQCP): DCA Systems™ Microalbumin Creatinine Reagent Test
Quality Control Plan (QCP): The IQCP is following exactly what is written in manufacturer’s instructions for doing the test, and in addition the laboratory follows their practices, resources, and procedures as stated in the lab’s manual assuring accurate and reliable test results at all times. The laboratory relies on internal built in controls of the test system assuring the test system is properly working. For quality control, the laboratory is following the manufacturer’s recommended quality control procedure by testing external quality control material for each shipment of a new kit lot number. The laboratory has quality control data supporting the manufacturer’s QC recommendation.
Quality Assessment (QA): All IQCP QA monitoring is part of the lab’s overall QA plan and is an ongoing review process for the IQCP. The lab has a review system for the ongoing monitoring of the effectiveness of the lab’s IQCP. The monitoring includes, but not limited to specimen, environment, testing personnel, reagent, and test system. In addition, if there is a complaint on a test result and or unacceptable QC and proficiency test result, the lab investigates and takes corrective action. It may be necessary to change the IQCP.
Laboratory Director Signature: ________________________________________ Date: __________________________
I, the Laboratory Director delegate to: _______________________________________in preparation of this document which I am ultimately responsible for. Date: _______________________________
33782 Rev A, 11 2015

## Instruction for Use
Individualized Quality Control Plan (IQCP): Stratus® Acute Care Diagnostic System
Instruction for Use
Phase | Defined as Pre-Analytical, Analytical and Post-Analytical
Process Step | Defined in test instruction for use
Identified Risk - Source of Potential Risk | Defined in test instruction for use and institute practices
Frequency of Occurrence | This should be defined by individual site risk assessments for previously documented events
Severity of Harm | Can use risk asssesment guide to determine category
Risk Mitigation: Manage/Prevent Errors from Occurring | This is provided in instructions for use, operators guides, current SOP's and various regulatory instructions
Risk Level Acceptable: Yes or No | Answer according to level of risk being mitigated
Reference | References from manufacturer, regulatory bodies and individual institute SOP's

## Specimen
Individualized Quality Control Plan (IQCP): DCA Systems™ Microalbumin Creatinine - Specimen
Specimen
Phase | Process Step | Identified Risk - Source of Potential Risk | Risk Mitigation: Manage/Prevent Errors from Occurring | Risk Level Acceptable: Yes or No | Reference
Pre-Analytic
Pre-Analytic | Patient Identification | Did not follow protocol to properly verify patient identity and properly record patient information. | Operators should be trained to use system barcode to ensure proper data entry of patient identification number and leverage offered custom-definable data-entry fields, if applicable. Follow established SOP XXX.XXX.
Pre-Analytic | Specimen Collection | Did not follow protocol for proper urine sample collection. May have added preservatives or frozen samples. May have collected samples of patients with suspected UTI. May have collected specimen of patients who have undergone strenuous activity. | Operators should be trained to follow site SOP for collecting manufacturer's recommended urine sample types: random, overnight or 24-hour urine specimens. As per manufacture's instructions, samples should not be frozen nor should preservatives be used. Specimens of patients with or suspected of having UTI or patients following strenuous activity should not be used as per the manufacturer's instructions. Follow established SOP XXX.XXX.
Pre-Analytic | Specimen Labeling | Did not properly label specimen. | Operators should be trained to follow site instructions for printing and applying sample barcode labels to specimen collection container to ensure proper patient-to-sample identification. Follow established SOP XXX.XXX.
Pre-Analytic | Specimen Transport | Improper storage and handling of specimen; recommended storage temperature may not have been followed or improper temperature storage during transportation. | Follow established SOP XXX.XXX.
Pre-Analytic | Infection Prevention | Improper handling of urine specimen. | All urine and reagent waste products are fully contained within the test cartridges and minimize operator exposure. Follow site instructions for proper disposal of test cartridge as per SOP XXX.XXX. Operators should be trained to follow the system cleaning instructions found in the operators manual. Follow established SOP XXX.XXX for infection prevention.
Pre-Analytic | Specimen Storage | Samples not properly stored by either freezing or addition of preservatives. | Operators should be trained to follow the manufacturer's instruction for urine specimen storage at 2-8°C (36-46°F) for up to two weeks. Do not freeze. Follow established SOP XXX.XXX.
Pre-Analytic | Specimen Preservation | Samples not properly stored by using preservatives. | Operators should be trained to follow manufacturer's instructions for urine specimen preservation. The use of preservatives is not recommended. Follow established SOP XXX.XXX.
Pre-Analytic | Identification of patient with gross proteinuria: albumin > 2000 mg/L | Patient history or trend of extremely high albumin levels unknown. | If proteinuria > 2000 mg/L suspected, use an alternative method or dilute the sample as per SOP XXX.XXX.
Analytic
Analytic | Specimen Integrity: Albuminuria | Urine specimens tested that may be diluted or highly concentrated. | System automatically corrects for patient urine concentration by measuring creatinine and calculating albumin-to-creatinine ratio. The system automatically identifies samples with ratio >2000 mg/g or 226 mg/mmol. Operators should be trained to understand the albumin-to-creatinine ratio flagging as found in the manufacturer's instructions for use.
Analytic | Specimen Integrity: Hemolysis or Highly Pigmented Urine | Urine specimens that are visibly hemolyzed or highly pigmented are used. | Operators should be trained to inspect sample integrity and follow SOP XXXX.XXX and follow manufacturer's instructions to not use visibly hemolyzed or highly pigmented urines. System may detect turbid or highly pigmented urines and may produce an error.
Analytic | Specimen Integrity: Interfering Substances | N/A | Operators should be familiar with the manufacturer's list of substances that do not interfere with the test Follow SOP XXXX.XXX.
Analytic | Specimen Integrity: UTI | Testing specimen of patients with known or suspected UTI. | Operators should be trained to follow manufacturer's instructions to not test if patient is know or suspected to have a UTI. Follow SOP XXXX.XXX
Analytic | Specimen Integrity: Refrigerated Sample | Improper warming of refrigerated specimen to room temperature. | Operator should be trained to follow the manufacturer's instructions in the operators manual or quick reference guide. Specimens should be brought to room temperature before testing.
Analytic | Specimen Integrity: Sample Bubbles | Not properly inspecting the glass capillary for presence of bubbles. | Operator should be trained to follow sample inspection instruction in the operators manual. Operator must check for bubbles in the glass capillary holder and repeat collection if bubbles are visible. Follow SOP XXXX.XXX.
Analytic | Specimen Integrity: Inadequate Volume | Do not follow manufacturers instruction for filling the urine capillary and using the visible cue to indicate proper sample volume.
Post-Analytic
Post-Analytic | Specimen Disposal | Improper disposal of specimen after testing. | Sample capillary is snapped into the fully contained test cartridge and minimize operator exposure. Follow site instructions for proper disposal of test cartridge as per SOP XXX.XXX.
33782 Rev A, 11 2015

## Operator
Individualized Quality Control Plan (IQCP): DCA Systems™ Microalbumin Creatinine - Operator
Operator
Testing Phase | Process Step | Identified Risk - Source of Potential Risk | Risk Mitigation: Manage/Prevent Errors from Occurring | Risk Level Acceptable: Yes or No | Reference
Pre-Analytic/Analytic/Post-Analytic
Pre-Analytic, Analytic, Post-Analytic | Operator Training | Improper operator training or access of unauthorized operators. | System can be configured to manage operator access and lock-out unauthorized personnel that are not adequately trained. If interfaced to RapidComm Data Management System operator competencies can be managed via RapidComm and operators with expired training can be locked out of accessing and operating the system. Operators should be trained via SOP XXXX.XXX.
Pre-Analytic, Analytic, Post-Analytic | Operator Competency | Operators may have expired training records or have not been recertified. Proper testing technique, sample handling, reagent handling, system maintenance are not being accurately followed. | System can be configured to manage operator access and lock-out unauthorized operators that are not adequately trained. If interfaced to RapidComm Data Management System operator competencies can be managed via RapidComm and operators with expired training can be locked out of accessing and operating the system. Operators competency should be followed. SOP XXXX.XXX.
Pre-Analytic, Analytic, Post-Analytic | Adequate Staffing | Untrained operators may be running the testing. | System and testing procedure is simple to use. There is no sample or reagent preparation required to run the test. Test calibration is simple and requires a single card barcode swipe with every new lot of reagent tested. If operator is not trained or unauthorized, the system can be programmed to lock-out users. Operators should be training via SOP XXXX.XXX.
Pre-Analytic, Analytic, Post-Analytic | Untrained Operator Access | Untrained operators may be running the testing. | System and testing procedure is simple to use. There is no sample or reagent preparation required to run the test. Test calibration is simple and requires a single card barcode swipe with every new lot of reagent tested. If operator is not trained or unauthorized, the system can be programmed to lock-out users. Operators should be training via SOP XXXX.XXX.
Pre-Analytic, Analytic, Post-Analytic | Infection Control | Untrained operators may be running the testing. Improper handling of specimen, reagent or system. | All urine and reagent waste products are fully contained within the test cartridge and minimize operator exposure. Follow site instructions for proper disposal of test cartridge as per SOP XXX.XXX. Operators should be trained to follow the system cleaning instructions found in the operators manual. Follow established SOP XXX.XXX for infection prevention.
33782 Rev A, 11 2015

## Reagent
Individualized Quality Control Plan (IQCP): DCA Systems™ Microalbumin Creatinine - Reagent
Reagent
Phase | Process Step | Identified Risk - Source of Potential Risk | Risk Mitigation: Manage/Prevent Errors from Occurring | Risk Level Acceptable: Yes or No | Reference
Pre-Analytic
Pre-Analytic | Reagent Contamination | Improper use of damage reagent test cartridges. Not following maintenance schedule. | Ensure operators are trained to perform quality control checks with every new lot of reagent or as per the product Instructions for Use or via site SOP XXXX.XXX. System can be configured to automatically alert users when new lot of reagents is introduced and force QC to be run. All reagent cartridges are fully contained and minimize reagent contamination. Operators should be trained to discard the reagent cartridge if it is visibly damaged or meets any of the other criteria listed in the Instructions for Use. Operators should be trained to follow the system cleaning instructions found in the operators manual to keep the reagent compartment clean. Follow established SOP XXX.XXX for maintaining the system.
Pre-Analytic | Reagent Preparation | N/A for reagent test cartridge. Improper preparation of lyophilized quality control reagents. | System does not require any reagent preparation to run test. Operators should be trained to properly handle the reagent cartridges. If using Siemens quality controls, follow control material preparation steps as per the quality control Instructions for Use.
Pre-Analytic | Reagent Handling | Improper storage of reagent test cartridges. Not brining reagent test cartridges to room temperature before testing. | Operators should be trained to refrigerate reagents if they wish to use reagents up to the expiration dating on the kit box and store at 2-8°C (36-46°F). Reagents can be stored for up to 3 months at room temperature and before the expiration date. If reagents are refrigerated, they must be brought to room temperature for 15 minutes prior to using on system. Reagent test cartridges must be used within 10 minutes after opening the foil pouch. System has internal checks that can detect when optimal temperature conditions in the measurement chamber are not reached and may trigger an error that may be due to reagent cartridge not being brought to room temperature after refrigeration. Follow SOP XXXX.XXX
Pre-Analytic | Infection Prevention | Improper handling of reagent test cartridge. | Reagent test cartridges are fully contained to minimize operator exposure. Follow site instructions for proper disposal of test cartridge as per SOP XXX.XXX. Operators should be trained to follow the system cleaning instructions found in the operators manual. Follow established SOP XXX.XXX for infection prevention.
Pre-Analytic | Specimen Delivery | Sample collection device improperly loaded into reagent test cartridge. | Sample collection device can only be inserted into the reagent test cartridge in one particular direction. Operators should be trained to follow the proper loading of sample collection device into regent test cartridge. Follow instructions in the operators manual to properly load the device. Follow SOP XXXX.XXX for appropriate operator training.
Analytic
Analytic | Reagent Calibration | Lot calibration card is lost or damaged; barcode illegible. | Operator must be trained to manually scan the calibration card once with each new lot of reagent. System will prevent any patient, control or proficiency testing until new lot calibration card is scanned. System automatically rejects calibration for lots of reagents that have expired. Follow SOP XXXX. XXX if lot calibration card is unable to be scanned.
Analytic | Lot and Expiration Dating | Reagent test cartridge barcode is damaged or illegible. | Operator must be trained to manually scan each reagent test cartridge in order for the test to initiate. System will prevent any patient, control or proficiency testing if reagent lot calibration is not previously scanned or if the reagent lot has expired. Follow SOP XXXX. XXX if reagent cartridge barcode is damaged or illegible.
Analytic | Reagent Cartridge Loading | Reagent cartridge is improperly loaded. | Reagent test cartridges can only be inserted into the reagent compartment in one particular manner. Operators should be trained to follow the proper loading of reagent test cartridge and pulling of reagent foil tab to release reagents into fully enclosed test cartridge for testing. Follow SOP XXXX.XXX for appropriate operator training.
Post-Analytic
Post-Analytic | Result Interpretation | Improperly recording results. Improperly interpreting test results. | System automatically corrects for patient urine concentration by measuring creatinine and calculating albumin-to-creatinine ratio. The system automatically identifies samples with ratio >2000 mg/g or 226 mg/mmol. Operators should be trained to understand the albumin-to-creatinine ratio flagging as found in the manufacturer's Instructions for Use. Follow SOP XXXX.XXX for result interpretation and reporting.
Post-Analytic | Reagent Cartridge Removal | Improper removal of reagent test cartridge. | Operators should follow manufacturer's instruction for cartridge removal as specified in the operators manual to avoid injury and damage to the system.
Post-Analytic | Reagent Cartridge Disposal | Improper reagent test cartridge disposal. | Operators should follow manufacturer's instruction for cartridge removal as specified in the operators manual to avoid injury and damage to the system. Follow site instructions for proper disposal of test cartridge as per SOP XXX.XXX.
Post-Analytic | Infection prevention | Improper handling of reagent test cartridge. | Reagent test cartridges are fully contained to minimize operator exposure. Follow site instructions for proper disposal of test cartridge as per SOP XXX.XXX. Operators should be trained to follow the system cleaning instructions found in the operators manual. Follow established SOP XXX.XXX for infection prevention.
33782 Rev A, 11 2015

## Environment
Individualized Quality Control Plan (IQCP): DCA Systems™ Microalbumin Creatinine - Environment
Environment
Phase | Process Step | Identified Risk - Source of Potential Risk | Risk Mitigation: Manage/Prevent Errors from Occurring | Risk Level Acceptable: Yes or No | Reference
Pre-Analytic
Pre-Analytic | Temperature: Analyzer | Lab temperature outside the operating temperature of the system. | Operators should be trained to ensure that the lab testing area temperature is adequate to meet the system operating temperature of 18-30°C as per requirements in the operators manual. System performs reagent compartment temperature checks and may generate an error code if outside the acceptable limit. Follow SOP XXXX.XXX for monitoring and lab temperature.
Pre-Analytic | Temperature: Cartridges | Improper storage of reagent test cartridges. Not brining reagent test cartridges to room temperature before testing. | Operators should be trained to refrigerate reagents if they wish to use reagents up to the expiration dating on the kit box and store at 2-8°C (36-46°F). Reagents can be stored for up to 3 months at room temperature and before the expiration date. If reagents are refrigerated, they must be brought to room temperature for 15 minutes prior to using on system. Reagent test cartridges must be used within 10 minutes after opening the foil pouch. System has internal checks that can detect when optimal temperature conditions in the measurement chamber are not reached and may trigger an error that may be due to reagent cartridge not being used at room temperature. Follow SOP XXXX.XXX for monitoring and recording lab refrigerator and testing area temperature.
Pre-Analytic | Ambient Environment requirements | Ambient environment is outside of the system testing requirements. | Operators should be trained to ensure the system is used in ambient temperature environment that meet the recommendations of 18-30° (64-86° F) in the operators manual. Follow SOP XXXX.XXX for monitoring and recording ambient temperature in the laboratory.
Pre-Analytic | Airflow/ventilation Instrument | May cause errors if airflow is restricted. Air filter not changed as per maintenance recommendations. | System performs automatic checks that may cause errors if airflow is restricted. Operators should be trained to follow the maintenance schedule for changing air filter in the operators manual every quarter. Follow SOP XXXX.XX
Pre-Analytic | Humidity: Instrument/Cartridge | Humidity is outside of the system testing requirements. | Operators should be trained to ensure the system is used within the relative humidity tolerances recommended in the operators manual of 10% to 90% (non-condensing, actively controlled). Follow SOP XXXX.XXX for monitoring and recording humidity measurements in the laboratory.
Pre-Analytic | Dust | Dust on the barcode window and reagent compartment. | System uses self-contained reagent test cartridges to prevent dust introduction into measurement system. Operators should be trained to change the air filter as per the required quarterly maintenance schedule as well as perform the routine maintenance schedule to keep the analyzer and the reagent compartment free from dust.
Pre-Analytic | Water | N/A | System does not require any external sources for water.
Pre-Analytic | Utilities (electrical power disruption) | N/A | System cannot be operated during power interruptions and should always be plugged into a grounded outlet.
Pre-Analytic | Infection prevention | Improper handling of reagent test cartridge. | Reagent test cartridges are fully contained to minimize operator exposure. Follow site instructions for proper disposal of test cartridge as per SOP XXX.XXX. Operators should be trained to follow the system cleaning instructions found in the operators manual. Follow established SOP XXX.XXX for infection prevention.
Analytic
Analytic | Temperature: Analyzer | Lab temperature outside the operating temperature of the system. | Operators should be trained to ensure that the lab testing area temperature is adequate to meet the system operating temperature requirements in the operators manual. System performs reagent compartment temperature checks and may generate an error code if outside the acceptable limit. Follow SOP XXXX.XXX for monitoring and lab temperature.
Analytic | Temperature: Cartridges | Improper storage of reagent test cartridges. Not brining reagent test cartridges to room temperature before testing. | System has internal checks that can detect when optimal temperature conditions in the measurement chamber are not reached and may trigger an error that may be due to reagent cartridge not being used at room temperature. Follow SOP XXXX.XXX for monitoring and recording lab refrigerator and testing area temperature.
Post-Analytic
Post-Analytic | Infection prevention | Improper handling of reagent test cartridge. | Reagent test cartridges are fully contained to minimize operator exposure. Follow site instructions for proper disposal of test cartridge as per SOP XXX.XXX. Operators should be trained to follow the system cleaning instructions found in the operators manual. Follow established SOP XXX.XXX for infection prevention.
33782 Rev A, 11 2015

## Test System
Individualized Quality Control Plan (IQCP): DCA Systems™ Microalbumin Creatinine - Test System
Test System
Phase | Process Step | Identified Risk - Source of Potential Risk | Risk Mitigation: Manage/Prevent Errors from Occurring | Risk Level Acceptable: Yes or No | Reference
Pre-Analytic
Pre-Analytic | Inadequate Volume/Sampling | Do not follow manufacturers instruction for filling the urine capillary and using the visible cue to indicate proper sample volume. | Operator should be trained to follow sample collection instructions in the operators manual and ensure the sample reaches the starch plug, approximately 5 seconds. Operator must allow sample to flow to reach the starch plug. This will ensure that the 40 µL of urine needed for testing is obtained. Follow SOP XXXX.XXX.
Pre-Analytic | Calibration Issues | Lot calibration card is lost or damaged; barcode illegible. | Operator must be trained to manually scan the calibration card once with each new lot of reagent. System will prevent any patient, control or proficiency testing until a new lot calibration card is scanned. System automatically rejects calibration for lots of reagents that have expired. Follow SOP XXXX. XXX if lot calibration card is unable to be scanned.
Pre-Analytic | Optics | Results that do not match clinical picture. Improper maintenance of system. | System performs checks of mechanical components, optical components, electronic components and instrument heater block. Any failure will generate analyzer errors. Operators should be trained on how to retrieve the error logs and report any recurrent errors to manufacturer. Operators should be trained to perform routine maintenance and leverage the maintenance scheduling and reminder set-up on the analyzer. Follow SOP XXXX.XXX.
Pre-Analytic | Barcode Data Entry | Transcription error of patient, control and other manually entered data. | System supports an optional, external barcode reader to ensure proper sample identification, configurable definition of specific prefix and identification for patient ID. Additional configurable positive patient ID fields include: Sample ID and 8 configurable demographic fields. Operators should be trained about barcode capabilities and the symbologies supported by the analyzer listed in the operators manual. Follow SOP XXXX.XXX
Pre-Analytic | Quality Control | Improper preparation of quality control materials. Improper test reagent storage and handling. Improper testing technique. | Operators should be trained to perform quality control checks with every new lot of reagent or as per the product Instructions for Use or via site SOP XXXX.XXX. System can be configured to automatically alert users when new lot of reagents is introduced and force QC to be run. System can be set-up to remind or force QC regimen to be followed and lock out QC if regimen is not followed or QC fails. Operators should be trained to follow quality control material preparation and storage instructions found in the QC control kit Instructions for Use. Follow SOP XXXX.XXX
Pre-Analytic | Temperature monitor | Lab temperature outside the operating temperature of the system. | Operators should be trained to ensure that the lab testing area temperature is adequate to meet the system operating temperature of 18-30°C (64-86° F) requirements in the operators manual. System performs reagent compartment temperature checks and may generate an error code if temperature is outside the operating range. Follow SOP XXXX.XXX for monitoring lab temperature.
Analytic
Analytic | Connectivity to LIS/HIS | Unable to transmit data to middleware or LIS/HIS do to improper system connectivity set up or lost connection | System automatically re-transmits results in holding queue due to connectivity drop. System generates error codes related to software and hardware and/or server errors. Operators should be trained to back up and restore database if necessary. System settings, patient and QC data can be backed up and restored. As good lab practice, update SOP # XXX.XXX to include system tests prior to "go-live" and after software upgrades.
Analytic | Instrument Failure | Improper maintenance of analyzer. Improper use of analyzer. | System performs continuous checks, including: mechanical checks, heater-block temperature, and electronic checks and generates errors when limits are not met and prevents testing of patient samples and quality control materials. Operators should be trained to review error codes and instructions for next steps in the system operator manual. Operators should be trained to follow routine maintenance schedule. System can be set-up to remind users when to perform maintenance. Follow SOP XXXX.XXX.
Post-Analytic
Post-Analytic | Infection Prevention | Improper maintenance of analyzer. | Reagent test cartridges are fully contained to minimize operator exposure. Operators should be trained to follow the system cleaning instructions found in the operators manual. System can be set-up to remind users when to perform maintenance. Follow established SOP XXX.XXX for infection prevention.
Post-Analytic | Result reporting | Inaccurate transcription of results to patient record. | System can be connected to transmit results directly to middleware data management and or LIS/HIS. Follow established SOP XXXX.XXX for data recording.
33782 Rev A, 11 2015

## Risk Assessment Guide
Risk Assessment Guide* Hospital/Test System residual risk
Probability of Harm
Frequent = once/week
Probable = once/month
Occasional = once/year
Remote = once/few years
Improbable = once/lifetime
Severity of Harm
Negligible = inconvenient or temporary discomfort
Minor = temporary injury or impairment not requiring professional medical intervention
Serious = injury or impairment requiring professional medical intervention
Critical = permanent impairment or life-threatening injury
Catastrophic = patient death
Probability of Harm | Negligible | Minor | Serious | Critical | Catastrophic
Frequent | Unacceptable | Unacceptable | Unacceptable | Unacceptable | Unacceptable
Probable | Acceptable | Unacceptable | Unacceptable | Unacceptable | Unacceptable
Occasional | Acceptable | Acceptable | Perform Risk-Benefit Analysis | Unacceptable | Unacceptable
Remote | Acceptable | Acceptable | Perform Risk-Benefit Analysis | Perform Risk-Benefit Analysis | Unacceptable
Improbable | Acceptable | Acceptable | Perform Risk-Benefit Analysis | Perform Risk-Benefit Analysis | Perform Risk-Benefit Analysis
*ISO 14971: Medical devices- Application of risk management to medical devices. www.iso.org
NOTE: This is an example only. Risk Gradient tables must be established in adherence to site-specific requirements and risk criteria.