Urinalysis CLINITEK Status Connect System with Siemens MULTISTIX Family of Reagent Strips Prepared By ____________________________________Date: _______________ Approved By ____________________________________Date: ________________ Effective Date ________________________________________ Discontinued Date: _______________________________(retain this procedure for at least two years) Supersedes an Earlier Procedure: _____________(Y or N) Earlier Procedure Discontinuance Date: ___________________

The medical/laboratory Director or the Director’s designee will review all copies of this procedure at least once a year. The director should keep a log of the copies being maintained.

| Reviewed By: | Date: | Comments: |
| --- | --- | --- |

# Urinalysis: CLINITEK Status Connect System with Siemens MULTISTIX Family of Test Strips

# Principle of the Test

Siemens MULTISTIX® Urinalysis Strips are read instrumentally by the CLINITEK Status® Connect System. Depending on the product being used, Siemens MULTISTIX Urinalysis Strips provide tests for:

- Glucose

- Bilirubin

- Ketone (acetoacetic acid)

- Specific gravity

- Blood

- pH

- Protein

- Urobilinogen

- Nitrite

- Leukocytes in urine

Edit the above list and the following table to omit tests that your laboratory does not perform

Chemical Principles of Siemens MULTISTIX Urinalysis Strips

| Test Name | Chemical Principle |
| --- | --- |
| Glucose | Glucose oxidase catalyzes the formation of gluconic acid and hydrogen peroxide from the oxidation of glucose. Peroxidase catalyzes the reaction of hydrogen peroxide with a potassium iodide chromogen to oxidize the chromogen to colors ranging from green to brown. |
| Bilirubin | Bilirubin couples with diazotized dichloroaniline in a strongly acid medium. Colors range through various shades of tan. |
| Ketone | Acetoacetic acid reacts with nitroprusside. Colors range from buff-pink, for a negative reading, to maroon for a positive reading. |
| Specific Gravity | pKa changes occur for certain pretreated polyelectrolytes in relation to ionic concentration. In the presence of an indicator, colors range from deep blue-green in urine of low ionic concentration through green and yellow-green in urines of increasing ionic concentration. |
| Blood | Hemoglobin catalyzes the reaction of diisopropylbenzene dihydroperoxide and 3,3',5,5’-tetramethylbenzidine. Colors range from orange through green; very high levels of blood may cause the color development to continue to blue. |
| pH | The double indicator principle gives a broad range of colors covering the entire urinary pH range. Colors range from orange through yellow and green to blue. |
| Protein | At a constant pH, the development of any green color is due to the presence of protein (protein error-of-indicators principle). Colors range from yellow for “Negative” through yellow-green and green to green-blue for “Positive” reactions. |
| Urobilinogen | In a modified Ehrlich reaction, p-diethylaminobenzaldehyde in conjunction with a color enhancer reacts with urobilinogen in a strongly acid medium to produce a pink-red color. |
| Nitrite | Nitrate (derived from the diet) is converted to nitrite by the action of Gram negative bacteria in the urine. At the acid pH of the reagent area, nitrite in the urine reacts with p-arsanilic acid to form a diazonium compound. This diazonium compound couples with 1,2,3,4-tetrahydrobenzo(h)quinolin-3-ol to produce a pink color. |
| Leukocytes | Esterases in granulocytic leukocytes catalyze the hydrolysis of the derivatized pyrrole amino acid ester to liberate 3-hydroxy-5-phenyl pyrrole. This pyrrole then reacts with a diazonium salt to produce a purple product. |

# Clinical Application and Usefulness

Siemens MULTISTIX Urinalysis Strips are for in vitro diagnostic use. Urinalysis can provide the physician with important information regarding the status of a patient’s health. Test results may provide information regarding the status of:

Carbohydrate metabolism

Kidney function and liver function

Acid-base balance

Urinary tract infection

NOTE:  As with all laboratory tests, definitive diagnostic or therapeutic decisions should not be based on any single result or method.

# Specimen Collection and Handling

Specimen Collection and Storage

|  |  | BIOHAZARD All products or objects that come in contact with human or animal body fluids should be handled, before and after cleaning, as if capable of transmitting infectious diseases. Wear facial protection, gloves, and protective clothing. |
| --- | --- | --- |

Obtain a fresh urine specimen. A first-morning specimen is preferred but random collections are acceptable. Specimens should be at room temperature for less than two hours before testing.

Collect the urine in a clean, dry, covered container.

The urine specimen should be well mixed and un-centrifuged. If the urine depth in its container is less than 3 inches pour the specimen into a narrow tube, such as a URIN-TEK® Specimen Tube.

To avoid contamination of the specimen, any cultures ordered on the same specimen should be set up before performing urinalysis.

If testing is delayed (>2 hour after collection), specimen should be refrigerated (2 to 8°C) for preservation. Allow urine specimen to return to room temperature before testing.

Specimen Rejection Criteria

Add your laboratory-specific rejection criteria here

Do not accept the following:

Specimens that have remained at room temperature for longer than two hours

Specimens with urine preservatives

Specimens that arrive in homemade containers (glass jars, pill bottles, etc.)

Leaking specimen containers

If an unacceptable specimen is received, note the reason for rejection on the test requisition and request a new, acceptable specimen from the patient. Acceptable collection container and collection instructions should be provided to the patient.

WARNING: Some urine specimens may have been collected during a critical procedure or by means of an invasive procedure; therefore, it is important to never dispose of an unacceptable specimen until the caregiver has been notified.

Specimen Referral Criteria

Add your laboratory-specific referral criteria here

# Reagents

Storage and Stability

- Store Siemens MULTISTIX Urinalysis Strips at room temperature, 15 - 30°C.

- Do not store the reagent strips in direct sunlight. Protection from exposure to light, heat and ambient moisture is mandatory to guard against altered reagent reactivity.

- Store the unused reagent strips in the original bottle. Transferring unused reagent strips to other containers may cause the strips to deteriorate and become un-reactive.

- Do not remove desiccants from bottle.

- Do not use reagent strips beyond the expiration date.

- Initial and date the reagent bottle when you first open it.

- Do not remove the strip from the bottle until immediately before use.  Replace cap immediately and tightly after removing the strip.

- Avoid touching the test areas of the reagent strip.

- Discoloration or darkening of the reagent areas may indicate deterioration. If this happens, confirm the expiration date and/or check performance with known negative and positive controls. If acceptable results are not obtained, discard the deteriorated strips and retest using a new, unopened bottle of reagent strips

- Due to the nature of the urobilinogen and leukocytes reagents found on the strips, these two results may be decreased at temperatures below 22(C and increased at temperatures above 26(C.

Reagent Ingredients

Reagent ingredients for Siemens MULTISTIX Urinalysis Strips are as follows:

Edit this list to omit the tests your laboratory does not perform

| Test | Ingredients |
| --- | --- |
| Glucose | 2.2% w/w glucose oxidase (microbial, 1.3 IU); 1.0% w/w peroxidase (horseradish, 3300 IU); 8.1% w/w potassium iodide; 69.8% w/w buffer; 18.9% w/w nonreactive ingredients |
| Bilirubin | 0.4%w/w 2,4-dichloroaniline diazonium salt; 37.3%w/w buffer; 62.3% w/w nonreactive ingredients |
| Ketone | 7.1 % w/w sodium nitroprusside; 92.9% w/w buffer |
| Specific Gravity | 2.8% w/w bromthymol blue; 68.8% w/w poly (methyl vinyl ether/maleic anhydride); 28.4% w/w sodium hydroxide |
| Blood | 6.8% w/w diisopropylbenzene dihydroperoxide; 4.0% w/w 3,3',5,5'-tetramethylbenzidine; 48.0% w/w buffer; 41.2% w/w nonreactive ingredients |
| pH | 0.2% w/w methyl red; 2.8% w/w bromthymol blue; 97.0% w/w nonreactive ingredients |
| Protein | 0.3% w/w tetrabromphenol blue; 97.3% w/w buffer; 2.4% w/w nonreactive ingredients |
| Urobilinogen | 0.2% w/w p-diethylaminobenzaldehyde; 99.8% w/w nonreactive ingredients |
| Nitrite | 1.4% w/w p-arsanilic acid; 1.3% w/w 1,2,3,4-tetrahydrobenzo(h) quinolin-3-ol; 10.8% w/w buffer; 86.5% w/w nonreactive ingredients |
| Leukocytes | 0.4% w/w derivatized pyrrole amino acid ester; 0.2% w/w diazonium salt; 40.9% w/w buffer; 58.5% w/w nonreactive ingredients |

The strips have been determined to be nonhazardous under the guidelines issued by OSHA in 29 CFR 1910.1200(d).

Reagent Special Preparation

No special preparation for reagent strips is required.

# Instrument Operation and System Description

The CLINITEK Status Connector is intended for use with the CLINITEK Status®+ Analyzer.  Together, the two units comprise the CLINITEK Status Connect System.  The Connector also supports importing certain information using an optional bar code scanner.  It is for in vitro diagnostic use in:

The semi-quantitative or qualitative detection of bilirubin, blood (occult), glucose, ketone (acetoacetic acid), leukocytes, nitrite, pH, protein, specific gravity and urobilinogen in urine samples. The tests reported depend on the type of Siemens MULTISTIX Urinalysis Strip used.

The optical system consists of six light emitting diodes, a light guide, a mirror, a lens and a detector. Light from the LEDs travels along the light guide and is reflected off the calibration bar, strip or cassette onto the mirror. It is then directed through an aperture on the lens, from where it is focused onto the detector. The light intensity detected is converted into electrical impulses, which are processed by the instrument’s microprocessor and converted into clinically meaningful results.

# System Start-up and Maintenance

The test table insert and the test table should be kept clean if the analyzer is to operate properly.

WARNING: Do not autoclave the test table or test table insert.

WARNING: Care should be taken not to scratch the white calibration bar. If it is scratched or scuffed, obtain a new test table. Solvents of any kind must not be used to clean the bar.

Refer to your CLINITEK Status+ Analyzer Operator’s Manual for detailed cleaning and maintenance instructions.

System Start-up

Power on the system by pressing the on/off button located at the front of the instrument. The analyzer automatically runs a system diagnostic check during which it performs a series of electronic, signal and memory checks. The Select Ready screen displays after system initialization.

Calibration

The CLINITEK Status Connect System performs a “self-test” and calibration each time you power on the system. The analyzer performs an automatic calibration each time you run a test. The white calibration bar (on the test table) provides NIST traceable calibration.

Enabling the Connector to Communicate with the Analyzer

- At the Select Ready screen, select Instrument Set up.  The Choose Settings screen displays.

- Use the arrow keys to select Instrument Settings.

- Choose Select.  The Instrument Settings screen displays.

- Use the arrow keys to select Connectivity.

- Choose Select.  The Connectivity screen 1 of 2 displays.

- To enable communication between the instrument and the connector, select Enabled.  To prevent communication, select Disabled.

- Select Previous.

- Select Done twice to return to the Select Ready screen.

# Managing Operators List

To manage the Operator’s List, you can enable or disable authorized operator options, and add operator’s IDs.

# Setting Operators IDs

- At the Select Ready screen, select Instrument Set Up.  The Choose Settings screen displays.

- Use the arrow keys to select Instrument Settings.

- Choose Select.  The Instrument Settings screen displays.

- Use the arrow keys to select Authorized Operator.

- Choose Select.  The Authorized Operator screen displays.

- To permit access only by authorized operators, select Enabled.  To allow all operators access to the system, select Disabled.

- For Enabled, see Adding Operator IDs in Operator Manual to add at least one operator.  If Disabled is selected, select Done 3 times to return to the Select Ready screen.

# Scanning Patient Information

To require patient information entry only by scanning:

- At the Select Ready screen, select Instrument Set Up.  The Choose Settings screen displays.

- Use the arrow keys to select Operator and Patient Information.

- Choose Select.  The Input Settings screen displays.

- Select Custom Set Up.

- Select Next.  The Custom Settings-Operator screen 1 of 5 displays.

- Select Next.  The Custom Settings-Patient Information screen 2 of 5 displays.

- Select Bar Code Reader Settings.  The Patient Information screen displays.

- To require entry of patient information by either bar code scanner or manually, enter Disabled.

- Select Done.  The Customs Settings Patient Information screen 2 of 5 displays.

- Select Next 4 times.

- Select Done twice to return to the Select Ready screen.

# Managing Urine Colors

To include instrument-provided colors, perform the following steps:

- At the Select Ready screen, select Instrument Set Up.  The Choose Settings screen displays.

- Use the arrow keys to select Operator and Patient Information.

- Choose Select.  The Input Settings screen displays.

- Select Custom Set Up.

- Select Next.  The Custom Settings-Operator screen 1 of 5 displays.

- Select Next 3 times.  The Custom Settings-Sample Appearance screen 4 of 5 displays.

- Select Edit Colors.  The Sample Appearance-Select Colors screen 1 of 3 displays.

- To choose colors, select the button for the color you want.  To remove a selected color, select that color button again.

- Select Next.  The Sample Appearance-Select Colors screen 2 of 3 displays.

- To choose colors, select the button for the color you want.

- Select Next.  The Sample Appearance-Select Colors screen 3 of 3 displays.

- Select Next 3 times.  The Input Settings-Confirmation screen displays.

- Select Done twice to return to the Select Ready screen.

# Adding Customized Colors (up to 4 custom colors can be added)

- At the Sample Appearance-Select Colors screen 3 of 3, select Enter Custom Color 1 (2, 3, or 4) corresponding to each custom color.

- Enter the custom color using the alpha keyboard to enter text.  To enter numeric text, select 123.

- Select Enter.  The Sample Appearance-Select colors screen 3 of 3 displays.  If a custom color exists, the Sample Appearance screen displays.  Select Yes, to edit that custom color and delete all records.  Select No, to return to the Sample Appearance Select Colors screen 3 of 3.

- Select Next 4 times.  The Input Settings-Confirmation screen displays.

- Select Done twice to return to the Select Ready  screen

# Entering Urine Color and Clarity During Test Cycle (in Full Test or Custom mode only)

During a patient test, the Select Appearance-Test In Progress screen displays.

- To choose the standard option, select Yellow and Clear.  Go to Step 8.  To choose a urine color and clarity, select Other.  The Select Appearance-Test In Progress screen 2 of 4 displays.

- Select a color for the urine sample from the options displayed.

- Select Next.  The Select Appearance-Test In Progress screen 3 of 4 displays.

- If you want a different color for the urine sample, select a color from the options displayed.

- Select Next.  The Select Appearance-Test In Progress screen 4 of 4 displays.

- Select the clarity of the urine sample.

- Select Next.  The Analyzing-In-Progress screen displays while the test is in progress followed by the Results screen.

- Select Done to return to the Select Ready screen.

# Managing Strip Lot Number and Expiration Date

Setting Strip Information Prompt

- At the Select Ready screen, select Instrument Set Up.  The Choose Settings screen displays.

- Use the arrow keys to select Instrument Settings.

- Select Enter.  The Instrument Settings screen displays.

- Use the arrow keys to select Urinalysis Test Settings.

- Choose Select.  The Urinalysis Test Settings screen displays.

- Select Next.  The Urinalysis Test screen displays.

- To prompt for strip information before each test, select Enabled.  To bypass a prompt to enter strip information before each test, select Disabled.

- Select Done 3 times to return to the Select Ready screen.

# Entering the Strip Lot Number and Expiration Date

- At the Select Ready screen select Strip Test.  The Strip Screen displays.

- To use the last strip number and begin the test, select Use Last Lot.  To enter new strip data, select Enter new lot and expiration.  The Strip Lot screen displays.

- Enter the strip lot number.  Use the alpha keyboard to enter text.  To enter numeric text, select 123.

- Select Enter.  The Strip Expiration screen displays.

- Use the arrow keys to indicate the strip expiration date.

- Select Enter.  The Prepare Test screen displays.

- Select Start.

# Setting Up QC Strip Tests

Accessing QC Strip Set Up

- At the Select Ready screen, select Instrument Set Up.  The Choose Settings screen displays.

- Use the arrow keys to select Instrument Settings.

- Choose Select.  The Instrument Settings screen displays.

- Use the arrow keys to select QC Settings.

- Choose Select.  The QC Settings screen displays.

- For QC strip test, select Set Up.  The QC Settings-Strip test screen 1 of 3 displays.

Setting QC Strip Tests

- To allow QC prompts, select Enabled.  To prevent QC prompts, select Disabled.

- To require QC tests when they are due and not permit patient tests, select Required. To permit patient tests even if a QC test is due, select Reminder.

- Select Next.  The QC Settings-Strip Test screen 2 of 3 displays.

- To have the instrument determine QC pass or fail, select Instrument.  To have the operator determine QC pass of fail, select Operator.

- To permit patient tests if the QC test fails, select Yes.  To prevent patient tests if the QC test fails, select No.

- Select Next.  The QC Settings-Strip test screen 3 of 3 displays.

# Defining QC Strip Test Controls

- To set the number of levels to complete a QC test, select 1, 2 or 3.

- For each control set in Step 1, select the corresponding Control Level, 1, 2, or 3.

- Select Enter Name of Control.  The Name of Control screen displays.

- Enter the name of this control.  Use the alpha keyboard to enter text.  To enter numeric text, select 123.

- Select Enter.  The Control Level 1-Strip test screen displays.

- Select Enter Name of Level.  The Name of Level screen displays.

- Enter the name of this level.  Use the alpha keyboard to enter text.  To enter numeric text, select 123.

- Select Enter.  The Control Level 1-Strip test screen displays.

- Choose Select Strip Type.  The QC Strip Type-Select Strip for this QC level screen displays.

- Use the arrow keys to select the strip type.

- Choose Select.  The Control Level 1-Strip Test screen displays.

# Setting QC Strip Test Pass/Fail Ranges

- Select Set Pass Ranges.  The QC Level 1-Tests and Pass Ranges screen 1 of 3 displays.

- To set the glucose pass range, select the GLU radio button.

- Select the corresponding Pass Range button.  The QC Level 1-Set Pass range for GLU screen displays.

- Use the arrow keys to indicate the lowest acceptable clinical values for glucose.

- Use the arrow keys to indicate the highest acceptable clinical values for glucose.

- Select Set.  The QC Level 1-Tests and Pass Ranges screen 1 of 3 displays.

- Enter the pass ranges for glucose, ketone, blood, Bilirubin, specific gravity, pH, protein, nitrite, urobilinogen, and leukocytes as directed in steps 2 – 6.  Select Next to advance through all 3 screens.  When pass ranges for all tests are entered, the QC Level 1-Tests and Pass Ranges screen 3 of 3 displays.

- Select Done.  The Control Level 1-Strip Test screen displays.

- Select Done.  The QC Settings-Strip Test screen 3 of 3 displays.

- If necessary, repeat defining QC strip test controls for the next control level.

- When all QC levels are defined select Done.  The QC Settings screen displays.

# Setting the QC Strip Schedule

- For QC strip test, select Interval.  The QC Settings-Set QC Schedule for Strip Tests screen displays.

- To schedule QC tests by hours, select the Hours radio button and use the arrow keys to indicate the number of hours between QC tests.  Go to step 7.  To schedule QC tests by days, select the Days radio button and use the arrow keys to indicate the number of days between QC tests.

- Select Set QC Time.  The QC Settings-Set QC Time For Strip Tests screen displays.

- Select the 1, 2, or 3 radio button for the number of QC tests per day.

- Use the arrow keys to indicate time of day for each test.

- Select Set.  The QC Settings-Set QC Schedule for Strip Test screen displays.

- Select Done.  The QC Settings screen displays.

- Select Done 3 times to return to the Select Ready screen.

# Quality Control (QC)

QC Materials

Use Chek-Stix® Positive and Negative controls or equivalent quality control material from a third party vendor to monitor performance.  Prepare fresh control solutions as instructed in the control product insert.

Water should not be used as a negative control.

QC Frequency

Test known negative and positive specimens or controls whenever a new bottle of reagent strips is first opened, whenever you have questionable test results and when training new operators.

Analyze positive and negative control solutions on a periodic basis, following each laboratory’s established quality control procedures for ensuring compliance to regulatory requirements.

Add your laboratory-specific quality control frequency requirements here

Running a QC Strip Test

  - At the Select Ready screen, select QC Test Strip Due.  The QC Test screen displays.

  - Select QC Strip Test Required.

  - If the instrument is set to determine pass/fail, the Control Lot screen displays.  Got to step 8.  If the operator is set to determine pass/fail, the Name of Control screen displays.

  - Enter the control name.  Use the alpha keyboard to enter text.  To enter numeric text, select 123.

  - Select Enter.  The Name of Control Level screen displays.

  - Enter the control level.  Use the alpha keyboard to enter text.  To enter numeric text, select 123.

  - Select Enter.  The Control Lot screen displays.

  - Enter the control lot.  Use the alpha keyboard to enter text.  To enter numeric text, select 123.

  - Select Enter.  The Control Expiration screen displays.

  - Use the arrow keys to indicate the control lot expiration date.

  - Select Enter.  The Strip Lot screen displays.

  - Use the alpha keyboard to enter text.  To enter numeric text, select 123.

  - Select Enter.  The Strip Expiration screen displays.

  - Use the arrow keys to indicate the strip expiration date.

  - Select Enter.  The Prepare Test screen displays.

  - Select Start.  The Results screen displays.

  - Select Print.

  - Select Done.  The QC Test-Select PASS or FAIL screen displays.

  - To pass the test, select QC PASS.  To fail the test, select FAIL.

  - Select Done.  The QC Test-Results Summary screen displays.

  - Select Repeat failed QC test to repeat a failed QC test.

  - Select Done to return to the Select Ready screen.

Troubleshooting Out-of-Range QC Values

A QC run is acceptable when all values fall within the expected ranges.

If the QC results do not fall within the defined ranges then the run is rejected, and you must take the following corrective actions:

Review instructions to ensure that the test was performed according to the procedures recommended by Siemens HealthCare Diagnostics;

Verify that the reagent strips and control materials are not expired;

If necessary, re-run the quality control samples or contact Siemens HealthCare Diagnostics for more assistance.

# Routine Analysis

The CLINITEK Status+ Analyzer can be operated in two modes, Quick Test or Full Test. By default the analyzer will run in Quick Test mode. Refer to your CLINITEK Status+ Analyzer Operator’s Manual  for a complete listing of other default settings.

Quick Test Mode

In the Quick Test mode, testing begins as soon as Strip Test is selected on the main menu screen.

WARNING: Bring the patient sample to 15 - 30°C and mix well prior to testing. WARNING: Do not remove the strip from the bottle until immediately before use. Replace cap immediately and tightly after removing the strip.

At the main Select screen, touch Strip Test. A Prepare Test screen will appear.  
Follow these steps to prepare the test:

- Make sure the test table insert has the reagent strip holder facing upward.

- Have the test strip, urine sample and paper towel ready.

Touch START. Another Prepare Test screen will appear. Follow these steps to continue preparing the test:

WARNING: Once you touch the START button you have eight seconds to treat the reagent strip with the urine sample and place the strip on the test table.

  - Dip the reagent strip into the urine sample, wetting all pads. Immediately remove the strip from the urine.

  - Drag the edge of the strip against the side of the sample container as you remove it.

  - Blot by touching the edge of the strip to the paper towel to remove excess urine.

WARNING: Do not lay the pads on the paper towel or cover the pads by the paper towel.

  - Place the reagent strip in the channel of the test table with the test pads facing up. Slide the strip to the end of the channel.

WARNING: Do not push or pull the test table.

  - At the end of the eight second countdown, the test table and strip will automatically be pulled into the instrument.

The analyzer will perform an automatic calibration and begin analyzing the strip.

WARNING: Do not move or bump the table while the instrument is calibrating.

When the analysis is complete, the Results screen will display the first page of results.

Touch More to display the second page of results.

Remove the used urinalysis strip from the test table and dispose of it according to your standard laboratory procedures.

Touch Done to complete the test and return to the main Select screen.

Full Test Mode

To access the Full Test mode you must customize your settings.

- At the main Select screen, touch Instrument Set Up. The Choose Settings screen will appear.

- Use the down arrow to highlight Operator and Patient Information. Touch Select. The Input Settings screen will appear.

- Touch the circle in the Full Test box and touch Next. An Input Settings confirmation screen will appear.

- Touch Done and you will return to the main Select screen.

For complete information on Instrument Set Up and customizing your test sequence, see your CLINITEK Status+ Analyzer Operator’s Manual.

In the Full Test mode, you are prompted to enter an Operator ID, Patient Name, and/or Patient ID prior to running a test.

- At the main Select screen, touch Strip Test. The Operator ID screen will appear.

- If you were the last operator to enter an ID on the analyzer, touch Last Operator. The Patient Information screen will appear. Go to step 6.

- If you are a new operator, touch Enter New Operator ID. The Enter Operator ID screen will appear.

- Use the keypad to enter your ID using a maximum of 13 characters.

- Touch Enter. The Patient Information screen will appear.

- If the test sample is from a previous patient, touch Recall Patient. A list of up to 200 patients will appear on the screen.

- Use the up and down arrows to scroll through the patient names. Once the patient you are looking for has been highlighted, touch Select. A Prepare Test screen will appear. Go to step 13.

- If the test sample is from a new patient, touch Enter New Patient. The Enter Patient Name screen will appear.

- Use the keypad to enter the patient’s name using a maximum of 20 characters.

- Touch Enter. The Patient ID screen will appear.

- Use the keypad to enter the patient’s ID using a maximum of 13 characters.

- Touch Enter. A Prepare Test screen will appear.

WARNING: Bring the patient sample to room temperature (15 - 30°C) and mix well prior to       testing.

WARNING: Do not remove the strip from the bottle until immediately before use. Replace cap immediately and tightly after removing the strip.

- Follow these steps to continuing preparing the test:

- Make sure the test table insert has the reagent strip holder facing upward.

- Have the test strip, urine sample and paper towel ready.

Touch START. Another Prepare Test screen will appear displaying the next 4 steps:

WARNING: Once you touch the START button you have eight seconds to treat the reagent strip with the urine sample and place the strip on the test table.

  - Dip the reagent strip into the urine sample, wetting all pads. Immediately remove the strip from the urine.

  - Drag the edge of the strip against the side of the sample container as you remove it.

  - Blot by touching the edge of the strip to the paper towel to remove excess urine.

WARNING: Do not lay the pads on the paper towel or cover the pads by the paper towel.

  - Place the reagent strip in the channel of the test table with the test pads facing up. Slide the strip to the end of the channel.

WARNING: Do not push or pull the test table.

  - At the end of the eight second countdown, the test table and strip will automatically be pulled into the instrument.

The analyzer will perform an automatic calibration and begin analyzing the strip.

WARNING: Do not move or bump the table while the instrument is calibrating.

While the strip is being analyzed, a Select Appearance screen will be displayed. The urine sample must be visually observed and then the appropriate color and clarity must be selected.

- If the urine sample is yellow and clear, touch the Yellow and Clear button.

- If the urine sample is not yellow and clear, touch the Other button for more choices.

- If you touched the Other button, then select the appropriate color by touching the circle button that corresponds to the correct description.

- Select the clarity by touching the circle that corresponds to the correct description. Then touch Next.

When the analysis is complete, the Results screen will display the first page of results.

Touch More to display the second page of results.

Remove the used urinalysis strip from the test table and dispose of it according to your standard laboratory procedures.

Touch Done to complete the test and return to the main Select screen.

STAT Samples

STAT samples are run like all other samples. Follow directions as outlined above in either Quick Test Mode or Full Test Mode.

# Reporting Results

As with all in vitro diagnostic assays, each laboratory should determine its own reference range(s) for the diagnostic evaluation of patient results.

Reference Interval

Add your laboratory-specific expected values here

| Glucose | Negative |
| --- | --- |
| Bilirubin | Negative |
| Ketone | Negative |
| Specific Gravity | 1.001 – 1.035 |
| Occult Blood | Negative |
| pH | 5 – 9 |
| Protein | Negative |
| Urobilinogen | (1.0 mg/dL (E.U./dL) |
| Nitrite | Negative |
| Leukocytes | Negative |
| Color | Yellow or Amber |
| Clarity | Clear |

Critical Values

Add your laboratory-specific critical values here

Reporting Protocol for Critical Values

Add your laboratory-specific protocol for reporting results here

Units for Reporting Results

The CLINITEK Status+ Analyzer reports results in conventional units.

Add your laboratory-specific reporting units for each test in the table below:

| Test | Reporting Unit |
| --- | --- |
| Glucose | mg/dL |
| Bilirubin | N/A |
| Ketone | mg/dL |
| Specific Gravity | (1.005 – (1.030 (in 0.005 increments) |
| Occult Blood | N/A |
| pH | 5.0 – (9.0 (in 0.5 increments) |
| Protein | mg/dL |
| Urobilinogen | mg/dL |
| Nitrite | N/A |
| Leukocytes | N/A |
| Color | N/A |
| Clarity | N/A |

Detectable Range

| Reagent Area | Sensitivity |
| --- | --- |
| Glucose | 75 – 125 mg/dL glucose |
| Bilirubin | 0.4 – 0.8 mg/dL bilirubin |
| Ketone | 5 – 10 mg/dL acetoacetic acid |
| Blood | 0.015 – 0.062 mg/dL hemoglobin |
| Protein | 15 – 30 mg/dL albumin |
| Nitrite | 0.06 – 0.1 mg/dL nitrite ion |
| Leukocytes | 5 – 15 cells/hpf in clinical urine |

Acceptable Results

Add your laboratory-specific acceptance criteria here

Patient test results are acceptable and may be reported when:

Corrective Action

Add your laboratory-specific repeat and corrective action protocol here

Patient test results must be repeated and corrective action taken when:

# Procedure Notes

Specimens used for Nitrite Testing

Using a first morning specimen or one that has incubated in the bladder for four hours or more optimizes nitrite test results.

Specimens used for Bilirubin and Urobilinogen Testing

It is especially important to use fresh urine to obtain optimal results with the tests for bilirubin and urobilinogen, as these compounds are very unstable when exposed to room temperature and light.

Disposal

Dispose of hazardous or biologically contaminated materials according to the practices of your institution. Discard all materials in a safe and acceptable manner, and in compliance with all federal, state, and local requirements.

# Method Limitations

Substances that cause abnormal urine color may affect the readability of test pads on urinalysis reagent strips. These substances include:

Drugs containing azo dyes (e.g., PYRIDIUM, Azo GANTRISIN, AZO GANTANOL)

NITROFURATOIN (MACRODANTIN, FURADANTIN)

Riboflavin

Visible levels of blood or bilirubin

For additional information on method limitations and performance characteristics, see the product information in the Siemens MULTISTIX Urinalysis Strips Instructions for Use.

# Equipment and Supplies

Add the product-specific name of the reagent strip(s) used in your laboratory.

- Siemens MULTISTIX Urinalysis Strips

CLINITEK Status+ Analyzer

CLINITEK Status Connector

Specimen collection container

Paper towels

# References

- Siemens MULTISTIX Urinalysis Strips Package Insert.

- Siemens Healthcare Diagnostics CLINITEK Status Analyzer Operator’s Manual.

- Siemens Healthcare Diagnostics CLINITEK STATUS Connect System Operator’s Manual.

- National Committee for Clinical Laboratory Standards (NCCLS), Clinical Laboratory Procedure Manuals, Third Edition (GP2-A3), 1996

*NCCLS is now known as: Clinical and Laboratory Standards Institute (CLSI).

Trademark Information

CLINITEK, MULTISTIX, CLINITEK Status+, Chek-Stix and URIN-TEK are trademarks of, or licensed to, Siemens Healthcare Diagnostics Inc.  All other trademarks and brands are the property of their respective owners.

Technical Assistance

Siemens Healthcare Technical Solutions Center: 1-877-229-3711

Serial Number:

Customer Account Number:

Siemens Healthcare Diagnostics

- Benedict Avenue

Tarrytown, NY  10591-6147

© 2018 Siemens Healthcare Diagnostics