Rt Dist Main RA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist SFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist SFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist T-P Trunk | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist UE Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt DPA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Epididymal A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Inf Thyroid A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Interlobar A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Mid Interlobar A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Sup Interlobar A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt ICA Siphon | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt LE Dist Anast 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Prox Anast 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Dist Graft 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Mid Graft 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Prox Graft 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Dist Inflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Mid Inflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Prox Inflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Dist Outflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Mid Outflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Prox Outflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Dist Anast 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Prox Anast 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Dist Graft 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Mid Graft 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Prox Graft 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Dist Inflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Mid Inflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Prox Inflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Dist Outflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Mid Outflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Prox Outflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Mid ARA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Axillary A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Mid CIA (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid Deep Brach A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Mid RHA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid MCA M1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ophthalmic A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop A Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid PTA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid DFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid DFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid Radial A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Main RA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid SFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid SFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid T-P Trunk | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid UE Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Pop A Aneurysm | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Post-inj Cavernosal A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt PCA P1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCoA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Pre-inj Cavernosal A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt DFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Prox ARA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Axillary A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Prox CIA (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox Deep Brach A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Prox RHA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox MCA M1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ophthalmic A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop A Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox PTA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox DFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox DFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox Radial A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Main RA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox SFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox SFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox T-P Trunk | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox UE Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Main RA Origin (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt RA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt Inf Segmental A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Mid Segmental A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Sup Segmental A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Subclav A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Sup Thyroid A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Terminal ICA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Testicular A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Art Anast | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Dist AVF | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid AVF | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox AVF | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist AVF Inflow A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid AVF Inflow A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox AVF Inflow A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist AVF Outflow V | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid AVF Outflow V | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox AVF Outflow V | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt AVF Venous Anast | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt UE Dist Anast | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Prox Anast | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Dist Graft | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Mid Graft | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Prox Graft | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Dist Inflow Vessel | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Mid Inflow Vessel | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Prox Inflow Vessel | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Dist Outflow Vessel | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Mid Outflow Vessel | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Prox Outflow Vessel | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Splenic A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA Origin (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin, Postprandial | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Suprarenal Ao (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Dist Donor A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Donor A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Donor A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Inf Arcuate A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Mid Arcuate A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Sup Arcuate A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Inf Arcuate A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Mid Arcuate A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Sup Arcuate A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Art Anast | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Dist RA (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid RA (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox RA (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist RA (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid RA (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Prox RA (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Donor A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Donor A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Donor A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Donor A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Donor A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Donor A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Inf Interlobar A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Mid Interlobar A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Sup Interlobar A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Inf Interlobar A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Mid Interlobar A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Sup Interlobar A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Inf Segmental A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Mid Segmental A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Sup Segmental A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Inf Segmental A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Mid Segmental A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Sup Segmental A (Tx) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116

## Vascular & GI Calculations
EnglishLabel | Description | DetailEnglishLabel | DefaultUnits | DefaultPrecision | MaxInstances | ConceptCoding | ValueType | ParentContainer | ParentLaterality | SRSection | SRSectionLaterality | VesselBranchModifier | AdditionalModifiers | AdditionalModifiers2
Liver | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000
Rt Kidney | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Kidney;SRT;T-71000 |  | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Kidney | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Kidney;SRT;T-71000 |  | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Kidney (Tx) | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Transplanted Kidney;SRT;T-7100A |  | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Kidney (Tx) | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Transplanted Kidney;SRT;T-7100A |  | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
GB | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Gallbladder;SRT;T-63000 |  | Finding Site;SRT;G-C0E3=GallBladder;SRT;T-74000;SRT;T-63000
Spleen | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Spleen;SRT;T-C3000 |  | Finding Site;SRT;G-C0E3=Spleen;SRT;T-C3000
Lt RA/Suprarenal Ao | Lt [Origin, Prox, Mid, Dist] Main RA/Suprarenal Ao PSV Ratio |  | unitless | 2 | 1 | Renal Artery/Suprarenal Aorta velocity ratio;99SIEMENS;RAOverSuprAoVel | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt RA/Suprarenal Ao | Rt [Origin, Prox, Mid, Dist] Main RA/Suprarenal Ao PSV Ratio |  | unitless | 2 | 1 | Renal Artery/Suprarenal Aorta velocity ratio;99SIEMENS;RAOverSuprAoVel | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt ARA/Suprarenal Ao | Lt [Origin, Prox, Mid, Dist] Accessory RA/Suprarenal Ao PSV Ratio |  | unitless | 2 | 1 | Accessory Renal Artery/Suprarenal Aorta velocity ratio;99SIEMENS;ARAOverSuprAoVel | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt ARA/Suprarenal Ao | Rt [Origin, Prox, Mid, Dist] Accessory RA/Suprarenal Ao PSV Ratio |  | unitless | 2 | 1 | Accessory Renal Artery/Suprarenal Aorta velocity ratio;99SIEMENS;ARAOverSuprAoVel | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100
PreVoid Bladder | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Prior to voiding;DCM;109134 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
PostVoid Bladder | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Post voiding;DCM;109135 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
Micturated Vol | Micturated Volume |  | cm³ | 2 | 1 | Micturated Volume;99SIEMENS;MicturatedVol | NUM | Bladder;SRT;T-74000 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
Rt ICA/CCA | Rt [Prox, Mid, or Dist] ICA PSV / Rt [Prox, Mid, or Dist] CCA PSV |  | unitless | 1 | 1 | ICA/CCA velocity ratio;LN;33868-1 | NUM | Artery of neck;SRT;T-45005 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt ICA/CCA | Lt [Prox, Mid, or Dist] ICA PSV / Lt [Prox, Mid, or Dist] CCA PSV |  | unitless | 1 | 1 | ICA/CCA velocity ratio;LN;33868-1 | NUM | Artery of neck;SRT;T-45005 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Prox MCA M2/Rt Dist MCA M1 | Rt Prox M2 MCA PSV/Rt Dist M1 MCA PSV Ratio |  | unitless | 2 | 1 | Rt Prox MCA M2 PSV/Dist MCA M1 PSV;99SIEMENS;RPMca2OverDMca1P | NUM | Blood Vessel of Head;SRT;T-40501 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Prox MCA M2/Lt Dist MCA M1 | Lt Prox M2 MCA PSV/Lt Dist M1 MCA PSV Ratio |  | unitless | 2 | 1 | Lt Prox MCA M2 PSV/Dist MCA M1 PSV;99SIEMENS;LPMca2OverDMca1P | NUM | Blood Vessel of Head;SRT;T-40501 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Prostate | Volume | Vol | cm³ | 2 | 1 | Prostate Volume derived from height, width and length;LN;15308-0 | NUM | Prostate;SRT;T-92000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000
Prostate | Prostate Weight | Weight | g | 1 | 1 | Organ Weight;SRT;M-02850 | NUM | Prostate;SRT;T-92000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000
Prostate Rt Mass 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Identifier;DCM;125010=1
Prostate Rt Mass 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Identifier;DCM;125010=2
Prostate Rt Mass 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Identifier;DCM;125010=3
Prostate Rt Mass 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Identifier;DCM;125010=4
Prostate Rt Mass 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Identifier;DCM;125010=5
Prostate Lt Mass 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Identifier;DCM;125010=1
Prostate Lt Mass 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Identifier;DCM;125010=2
Prostate Lt Mass 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Identifier;DCM;125010=3
Prostate Lt Mass 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Identifier;DCM;125010=4
Prostate Lt Mass 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Identifier;DCM;125010=5
Rt Testicle | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Testis;SRT;T-94000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Testicle | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Testis;SRT;T-94000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Epididymis | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Epididymis;SRT;T-95000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Epididymis | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Epididymis;SRT;T-95000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Testis Mass 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Identifier;DCM;125010=1
Rt Testis Mass 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Identifier;DCM;125010=2
Rt Testis Mass 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Identifier;DCM;125010=3
Lt Testis Mass 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Identifier;DCM;125010=1
Lt Testis Mass 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Identifier;DCM;125010=2
Lt Testis Mass 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Identifier;DCM;125010=3
Rt Thyroid | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Lobe of Thyroid Gland;SRT;T-B6070 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Thyroid | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Lobe of Thyroid Gland;SRT;T-B6070 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Thyroid Mass 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1
Rt Thyroid Mass 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2
Rt Thyroid Mass 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3
Rt Thyroid Mass 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4
Rt Thyroid Mass 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5
Lt Thyroid Mass 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1
Lt Thyroid Mass 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2
Lt Thyroid Mass 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3
Lt Thyroid Mass 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4
Lt Thyroid Mass 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5
Isthmus Mass 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1
Isthmus Mass 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2
Isthmus Mass 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3
Isthmus Mass 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4
Isthmus Mass 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5
Liver Mass 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1
Liver Mass 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2
Liver Mass 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3
Liver Mass 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4
Liver Mass 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5
Liver Mass 6 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6
Liver Mass 7 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7
Liver Mass 8 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8
Liver Mass 9 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9
Liver Mass 10 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10
Portal V Anast/Pre-Anast Vel Ratio | Portal V Anastomosis/Pre-Anastomosis Vmax Ratio |  | unitless | 2 | 1 | Portal Vein Anastomosis/Pre-Anastomosis Peak Velocity Ratio;99SIEMENS;PortlVnVmaxRatio | NUM | Transplantation of Liver;SRT;P1-5BD10 |  | Finding Site;SRT;G-C0E3=Transplantation of Liver;SRT;P1-5BD10
Rt Neck Mass 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Neck Mass 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Neck Mass 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Neck Mass 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Neck Mass 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Neck Mass 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Neck Mass 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Neck Mass 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Neck Mass 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Neck Mass 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt MCA/ICA | Rt [Prox, Mid, or Dist] MCA Mean (TCD) / Rt Dist ICA-skull entrance Mean (TCD) |  | unitless | 1 | 1 | MCA/ICA Mean (TCD);99SIEMENS;MCAOverICAMnTCD | NUM | Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt MCA/ICA | Lt [Prox, Mid, or Dist] MCA Mean (TCD) / Lt Dist ICA-skull entrance Mean (TCD) |  | unitless | 1 | 1 | MCA/ICA Mean (TCD);99SIEMENS;MCAOverICAMnTCD | NUM | Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 1 | Liver Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 1 | Liver Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 1 | Liver Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 1 | Liver Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 1 | Liver Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 1 | Liver Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 1 | Liver Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 1 | Liver Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 1 | Liver Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 1 | Liver Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 1 | Liver Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 1 | Liver Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 1 | Liver Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 1 | Liver Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 1 | Liver Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 2 | Liver Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 2 | Liver Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 2 | Liver Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 2 | Liver Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 2 | Liver Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 2 | Liver Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 2 | Liver Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 2 | Liver Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 2 | Liver Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 2 | Liver Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 2 | Liver Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 2 | Liver Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 2 | Liver Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 2 | Liver Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 2 | Liver Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 3 | Liver Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 3 | Liver Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 3 | Liver Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 3 | Liver Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 3 | Liver Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 3 | Liver Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 3 | Liver Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 3 | Liver Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 3 | Liver Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 3 | Liver Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 3 | Liver Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 3 | Liver Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 3 | Liver Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 3 | Liver Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 3 | Liver Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 4 | Liver Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 4 | Liver Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 4 | Liver Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 4 | Liver Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 4 | Liver Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 4 | Liver Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 4 | Liver Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 4 | Liver Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 4 | Liver Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 4 | Liver Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 4 | Liver Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 4 | Liver Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 4 | Liver Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 4 | Liver Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 4 | Liver Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 5 | Liver Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 5 | Liver Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 5 | Liver Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 5 | Liver Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 5 | Liver Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 5 | Liver Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 5 | Liver Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 5 | Liver Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 5 | Liver Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 5 | Liver Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 5 | Liver Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 5 | Liver Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 5 | Liver Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 5 | Liver Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 5 | Liver Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 6 | Liver Mass 6 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 6 | Liver Mass 6 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 6 | Liver Mass 6 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 6 | Liver Mass 6 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 6 | Liver Mass 6 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 6 | Liver Mass 6 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 6 | Liver Mass 6 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 6 | Liver Mass 6 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 6 | Liver Mass 6 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 6 | Liver Mass 6 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 6 | Liver Mass 6 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 6 | Liver Mass 6 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 6 | Liver Mass 6 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 6 | Liver Mass 6 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 6 | Liver Mass 6 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 7 | Liver Mass 7 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 7 | Liver Mass 7 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 7 | Liver Mass 7 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 7 | Liver Mass 7 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 7 | Liver Mass 7 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 7 | Liver Mass 7 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 7 | Liver Mass 7 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 7 | Liver Mass 7 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 7 | Liver Mass 7 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 7 | Liver Mass 7 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 7 | Liver Mass 7 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 7 | Liver Mass 7 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 7 | Liver Mass 7 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 7 | Liver Mass 7 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 7 | Liver Mass 7 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 8 | Liver Mass 8 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 8 | Liver Mass 8 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 8 | Liver Mass 8 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 8 | Liver Mass 8 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 8 | Liver Mass 8 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 8 | Liver Mass 8 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 8 | Liver Mass 8 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 8 | Liver Mass 8 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 8 | Liver Mass 8 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 8 | Liver Mass 8 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 8 | Liver Mass 8 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 8 | Liver Mass 8 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 8 | Liver Mass 8 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 8 | Liver Mass 8 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 8 | Liver Mass 8 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 9 | Liver Mass 9 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 9 | Liver Mass 9 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 9 | Liver Mass 9 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 9 | Liver Mass 9 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 9 | Liver Mass 9 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 9 | Liver Mass 9 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 9 | Liver Mass 9 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 9 | Liver Mass 9 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 9 | Liver Mass 9 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 9 | Liver Mass 9 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 9 | Liver Mass 9 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 9 | Liver Mass 9 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 9 | Liver Mass 9 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 9 | Liver Mass 9 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 9 | Liver Mass 9 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 10 | Liver Mass 10 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 10 | Liver Mass 10 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 10 | Liver Mass 10 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 10 | Liver Mass 10 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 10 | Liver Mass 10 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 10 | Liver Mass 10 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 10 | Liver Mass 10 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 10 | Liver Mass 10 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 10 | Liver Mass 10 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 10 | Liver Mass 10 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 10 | Liver Mass 10 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 10 | Liver Mass 10 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 10 | Liver Mass 10 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 10 | Liver Mass 10 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Mass 10 | Liver Mass 10 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 1 | Liver Site 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 1 | Liver Site 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 1 | Liver Site 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 1 | Liver Site 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 1 | Liver Site 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 1 | Liver Site 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 1 | Liver Site 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 1 | Liver Site 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 1 | Liver Site 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 1 | Liver Site 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 1 | Liver Site 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 1 | Liver Site 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 1 | Liver Site 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 1 | Liver Site 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 1 | Liver Site 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 1 | Liver Site 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 1 | Liver Site 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 1 | Liver Site 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 1 | Liver Site 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 1 | Liver Site 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 1 | Liver Site 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 1 | Liver Site 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 1 | Liver Site 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 1 | Liver Site 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 1 | Liver Site 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 1 | Liver Site 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 1 | Liver Site 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 1 | Liver Site 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 1 | Liver Site 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 1 | Liver Site 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 2 | Liver Site 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 2 | Liver Site 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 2 | Liver Site 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 2 | Liver Site 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 2 | Liver Site 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 2 | Liver Site 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 2 | Liver Site 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 2 | Liver Site 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 2 | Liver Site 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 2 | Liver Site 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 2 | Liver Site 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 2 | Liver Site 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 2 | Liver Site 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 2 | Liver Site 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 2 | Liver Site 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 2 | Liver Site 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 2 | Liver Site 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 2 | Liver Site 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 2 | Liver Site 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 2 | Liver Site 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 2 | Liver Site 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 2 | Liver Site 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 2 | Liver Site 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 2 | Liver Site 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 2 | Liver Site 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 2 | Liver Site 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 2 | Liver Site 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 2 | Liver Site 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 2 | Liver Site 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 2 | Liver Site 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 3 | Liver Site 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 3 | Liver Site 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 3 | Liver Site 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 3 | Liver Site 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 3 | Liver Site 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 3 | Liver Site 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 3 | Liver Site 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 3 | Liver Site 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 3 | Liver Site 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 3 | Liver Site 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 3 | Liver Site 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 3 | Liver Site 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 3 | Liver Site 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 3 | Liver Site 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 3 | Liver Site 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 3 | Liver Site 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 3 | Liver Site 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 3 | Liver Site 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 3 | Liver Site 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 3 | Liver Site 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 3 | Liver Site 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 3 | Liver Site 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 3 | Liver Site 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 3 | Liver Site 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 3 | Liver Site 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 3 | Liver Site 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 3 | Liver Site 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 3 | Liver Site 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 3 | Liver Site 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 3 | Liver Site 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 4 | Liver Site 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 4 | Liver Site 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 4 | Liver Site 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 4 | Liver Site 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 4 | Liver Site 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 4 | Liver Site 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 4 | Liver Site 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 4 | Liver Site 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 4 | Liver Site 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 4 | Liver Site 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 4 | Liver Site 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 4 | Liver Site 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 4 | Liver Site 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 4 | Liver Site 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 4 | Liver Site 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 4 | Liver Site 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 4 | Liver Site 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 4 | Liver Site 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 4 | Liver Site 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 4 | Liver Site 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 4 | Liver Site 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 4 | Liver Site 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 4 | Liver Site 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 4 | Liver Site 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 4 | Liver Site 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 4 | Liver Site 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 4 | Liver Site 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 4 | Liver Site 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 4 | Liver Site 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 4 | Liver Site 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 5 | Liver Site 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 5 | Liver Site 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 5 | Liver Site 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 5 | Liver Site 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 5 | Liver Site 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 5 | Liver Site 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 5 | Liver Site 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 5 | Liver Site 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 5 | Liver Site 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 5 | Liver Site 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 5 | Liver Site 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 5 | Liver Site 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 5 | Liver Site 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 5 | Liver Site 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 5 | Liver Site 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 5 | Liver Site 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 5 | Liver Site 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 5 | Liver Site 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 5 | Liver Site 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 5 | Liver Site 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 5 | Liver Site 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 5 | Liver Site 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 5 | Liver Site 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 5 | Liver Site 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 5 | Liver Site 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 5 | Liver Site 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 5 | Liver Site 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 5 | Liver Site 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 5 | Liver Site 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 5 | Liver Site 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 6 | Liver Site 6 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 6 | Liver Site 6 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 6 | Liver Site 6 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 6 | Liver Site 6 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 6 | Liver Site 6 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 6 | Liver Site 6 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 6 | Liver Site 6 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 6 | Liver Site 6 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 6 | Liver Site 6 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 6 | Liver Site 6 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 6 | Liver Site 6 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 6 | Liver Site 6 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 6 | Liver Site 6 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 6 | Liver Site 6 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 6 | Liver Site 6 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 6 | Liver Site 6 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 6 | Liver Site 6 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 6 | Liver Site 6 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 6 | Liver Site 6 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 6 | Liver Site 6 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 6 | Liver Site 6 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 6 | Liver Site 6 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 6 | Liver Site 6 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 6 | Liver Site 6 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 6 | Liver Site 6 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 6 | Liver Site 6 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 6 | Liver Site 6 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 6 | Liver Site 6 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 6 | Liver Site 6 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 6 | Liver Site 6 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 7 | Liver Site 7 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 7 | Liver Site 7 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 7 | Liver Site 7 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 7 | Liver Site 7 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 7 | Liver Site 7 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 7 | Liver Site 7 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 7 | Liver Site 7 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 7 | Liver Site 7 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 7 | Liver Site 7 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 7 | Liver Site 7 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 7 | Liver Site 7 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 7 | Liver Site 7 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 7 | Liver Site 7 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 7 | Liver Site 7 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 7 | Liver Site 7 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 7 | Liver Site 7 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 7 | Liver Site 7 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 7 | Liver Site 7 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 7 | Liver Site 7 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 7 | Liver Site 7 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 7 | Liver Site 7 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 7 | Liver Site 7 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 7 | Liver Site 7 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 7 | Liver Site 7 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 7 | Liver Site 7 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 7 | Liver Site 7 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 7 | Liver Site 7 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 7 | Liver Site 7 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 7 | Liver Site 7 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 7 | Liver Site 7 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 8 | Liver Site 8 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 8 | Liver Site 8 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 8 | Liver Site 8 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 8 | Liver Site 8 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 8 | Liver Site 8 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 8 | Liver Site 8 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 8 | Liver Site 8 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 8 | Liver Site 8 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 8 | Liver Site 8 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 8 | Liver Site 8 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 8 | Liver Site 8 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 8 | Liver Site 8 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 8 | Liver Site 8 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 8 | Liver Site 8 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 8 | Liver Site 8 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 8 | Liver Site 8 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 8 | Liver Site 8 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 8 | Liver Site 8 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 8 | Liver Site 8 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 8 | Liver Site 8 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 8 | Liver Site 8 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 8 | Liver Site 8 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 8 | Liver Site 8 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 8 | Liver Site 8 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 8 | Liver Site 8 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 8 | Liver Site 8 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 8 | Liver Site 8 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 8 | Liver Site 8 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 8 | Liver Site 8 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 8 | Liver Site 8 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 9 | Liver Site 9 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 9 | Liver Site 9 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 9 | Liver Site 9 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 9 | Liver Site 9 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 9 | Liver Site 9 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 9 | Liver Site 9 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 9 | Liver Site 9 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 9 | Liver Site 9 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 9 | Liver Site 9 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 9 | Liver Site 9 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 9 | Liver Site 9 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 9 | Liver Site 9 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 9 | Liver Site 9 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 9 | Liver Site 9 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 9 | Liver Site 9 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 9 | Liver Site 9 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 9 | Liver Site 9 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 9 | Liver Site 9 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 9 | Liver Site 9 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 9 | Liver Site 9 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 9 | Liver Site 9 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 9 | Liver Site 9 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 9 | Liver Site 9 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 9 | Liver Site 9 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 9 | Liver Site 9 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 9 | Liver Site 9 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 9 | Liver Site 9 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 9 | Liver Site 9 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 9 | Liver Site 9 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 9 | Liver Site 9 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 10 | Liver Site 10 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 10 | Liver Site 10 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 10 | Liver Site 10 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 10 | Liver Site 10 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 10 | Liver Site 10 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 10 | Liver Site 10 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 10 | Liver Site 10 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 10 | Liver Site 10 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 10 | Liver Site 10 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 10 | Liver Site 10 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Site 10 | Liver Site 10 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 10 | Liver Site 10 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 10 | Liver Site 10 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 10 | Liver Site 10 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 10 | Liver Site 10 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 10 | Liver Site 10 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 10 | Liver Site 10 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 10 | Liver Site 10 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 10 | Liver Site 10 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 10 | Liver Site 10 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Site 10 | Liver Site 10 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 10 | Liver Site 10 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 10 | Liver Site 10 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 10 | Liver Site 10 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 10 | Liver Site 10 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 10 | Liver Site 10 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 10 | Liver Site 10 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 10 | Liver Site 10 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 10 | Liver Site 10 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Site 10 | Liver Site 10 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 1 | Liver Segment 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 1 | Liver Segment 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 1 | Liver Segment 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 1 | Liver Segment 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 1 | Liver Segment 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 1 | Liver Segment 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 1 | Liver Segment 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 1 | Liver Segment 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 1 | Liver Segment 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 1 | Liver Segment 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 1 | Liver Segment 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 1 | Liver Segment 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 1 | Liver Segment 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 1 | Liver Segment 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 1 | Liver Segment 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 2 | Liver Segment 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 2 | Liver Segment 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 2 | Liver Segment 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 2 | Liver Segment 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 2 | Liver Segment 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 2 | Liver Segment 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 2 | Liver Segment 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 2 | Liver Segment 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 2 | Liver Segment 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 2 | Liver Segment 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 2 | Liver Segment 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 2 | Liver Segment 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 2 | Liver Segment 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 2 | Liver Segment 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 2 | Liver Segment 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 3 | Liver Segment 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 3 | Liver Segment 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 3 | Liver Segment 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 3 | Liver Segment 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 3 | Liver Segment 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 3 | Liver Segment 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 3 | Liver Segment 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 3 | Liver Segment 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 3 | Liver Segment 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 3 | Liver Segment 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 3 | Liver Segment 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 3 | Liver Segment 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 3 | Liver Segment 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 3 | Liver Segment 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 3 | Liver Segment 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4a | Liver Segment 4a Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4a | Liver Segment 4a Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4a | Liver Segment 4a Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4a | Liver Segment 4a Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4a | Liver Segment 4a Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4a | Liver Segment 4a Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4a | Liver Segment 4a Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4a | Liver Segment 4a Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4a | Liver Segment 4a Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4a | Liver Segment 4a Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4a | Liver Segment 4a Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4a | Liver Segment 4a Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4a | Liver Segment 4a Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4a | Liver Segment 4a Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4a | Liver Segment 4a Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4b | Liver Segment 4b Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4b | Liver Segment 4b Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4b | Liver Segment 4b Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4b | Liver Segment 4b Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4b | Liver Segment 4b Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4b | Liver Segment 4b Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4b | Liver Segment 4b Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4b | Liver Segment 4b Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4b | Liver Segment 4b Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4b | Liver Segment 4b Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4b | Liver Segment 4b Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4b | Liver Segment 4b Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4b | Liver Segment 4b Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4b | Liver Segment 4b Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 4b | Liver Segment 4b Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 5 | Liver Segment 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 5 | Liver Segment 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 5 | Liver Segment 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 5 | Liver Segment 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 5 | Liver Segment 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 5 | Liver Segment 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 5 | Liver Segment 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 5 | Liver Segment 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 5 | Liver Segment 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 5 | Liver Segment 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 5 | Liver Segment 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 5 | Liver Segment 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 5 | Liver Segment 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 5 | Liver Segment 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 5 | Liver Segment 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 6 | Liver Segment 6 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 6 | Liver Segment 6 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 6 | Liver Segment 6 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 6 | Liver Segment 6 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 6 | Liver Segment 6 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 6 | Liver Segment 6 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 6 | Liver Segment 6 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 6 | Liver Segment 6 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 6 | Liver Segment 6 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 6 | Liver Segment 6 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 6 | Liver Segment 6 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 6 | Liver Segment 6 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 6 | Liver Segment 6 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 6 | Liver Segment 6 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 6 | Liver Segment 6 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 7 | Liver Segment 7 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 7 | Liver Segment 7 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 7 | Liver Segment 7 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 7 | Liver Segment 7 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 7 | Liver Segment 7 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 7 | Liver Segment 7 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 7 | Liver Segment 7 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 7 | Liver Segment 7 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 7 | Liver Segment 7 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 7 | Liver Segment 7 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 7 | Liver Segment 7 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 7 | Liver Segment 7 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 7 | Liver Segment 7 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 7 | Liver Segment 7 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 7 | Liver Segment 7 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 8 | Liver Segment 8 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 8 | Liver Segment 8 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 8 | Liver Segment 8 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 8 | Liver Segment 8 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 8 | Liver Segment 8 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 8 | Liver Segment 8 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 8 | Liver Segment 8 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 8 | Liver Segment 8 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 8 | Liver Segment 8 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 8 | Liver Segment 8 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 8 | Liver Segment 8 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 8 | Liver Segment 8 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 8 | Liver Segment 8 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 8 | Liver Segment 8 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Liver Segment 8 | Liver Segment 8 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 1 | Thyroid Left Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 2 | Thyroid Left Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 3 | Thyroid Left Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 4 | Thyroid Left Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Mass 5 | Thyroid Left Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 1 | Thyroid Right Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 2 | Thyroid Right Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 3 | Thyroid Right Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 4 | Thyroid Right Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Mass 5 | Thyroid Right Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 1 | Thyroid Left Site 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 1 | Thyroid Left Site 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 1 | Thyroid Left Site 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 1 | Thyroid Left Site 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 1 | Thyroid Left Site 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 1 | Thyroid Left Site 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 1 | Thyroid Left Site 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 1 | Thyroid Left Site 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 1 | Thyroid Left Site 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 1 | Thyroid Left Site 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 1 | Thyroid Left Site 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 1 | Thyroid Left Site 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 1 | Thyroid Left Site 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 1 | Thyroid Left Site 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 1 | Thyroid Left Site 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 1 | Thyroid Left Site 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 1 | Thyroid Left Site 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 1 | Thyroid Left Site 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 1 | Thyroid Left Site 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 1 | Thyroid Left Site 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 1 | Thyroid Left Site 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 1 | Thyroid Left Site 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 1 | Thyroid Left Site 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 1 | Thyroid Left Site 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 1 | Thyroid Left Site 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 1 | Thyroid Left Site 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 1 | Thyroid Left Site 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 1 | Thyroid Left Site 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 1 | Thyroid Left Site 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 1 | Thyroid Left Site 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 2 | Thyroid Left Site 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 2 | Thyroid Left Site 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 2 | Thyroid Left Site 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 2 | Thyroid Left Site 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 2 | Thyroid Left Site 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 2 | Thyroid Left Site 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 2 | Thyroid Left Site 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 2 | Thyroid Left Site 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 2 | Thyroid Left Site 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 2 | Thyroid Left Site 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 2 | Thyroid Left Site 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 2 | Thyroid Left Site 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 2 | Thyroid Left Site 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 2 | Thyroid Left Site 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 2 | Thyroid Left Site 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 2 | Thyroid Left Site 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 2 | Thyroid Left Site 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 2 | Thyroid Left Site 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 2 | Thyroid Left Site 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 2 | Thyroid Left Site 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 2 | Thyroid Left Site 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 2 | Thyroid Left Site 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 2 | Thyroid Left Site 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 2 | Thyroid Left Site 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 2 | Thyroid Left Site 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 2 | Thyroid Left Site 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 2 | Thyroid Left Site 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 2 | Thyroid Left Site 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 2 | Thyroid Left Site 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 2 | Thyroid Left Site 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 3 | Thyroid Left Site 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 3 | Thyroid Left Site 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 3 | Thyroid Left Site 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 3 | Thyroid Left Site 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 3 | Thyroid Left Site 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 3 | Thyroid Left Site 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 3 | Thyroid Left Site 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 3 | Thyroid Left Site 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 3 | Thyroid Left Site 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 3 | Thyroid Left Site 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 3 | Thyroid Left Site 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 3 | Thyroid Left Site 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 3 | Thyroid Left Site 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 3 | Thyroid Left Site 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 3 | Thyroid Left Site 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 3 | Thyroid Left Site 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 3 | Thyroid Left Site 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 3 | Thyroid Left Site 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 3 | Thyroid Left Site 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 3 | Thyroid Left Site 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 3 | Thyroid Left Site 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 3 | Thyroid Left Site 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 3 | Thyroid Left Site 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 3 | Thyroid Left Site 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 3 | Thyroid Left Site 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 3 | Thyroid Left Site 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 3 | Thyroid Left Site 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 3 | Thyroid Left Site 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 3 | Thyroid Left Site 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 3 | Thyroid Left Site 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 4 | Thyroid Left Site 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 4 | Thyroid Left Site 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 4 | Thyroid Left Site 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 4 | Thyroid Left Site 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 4 | Thyroid Left Site 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 4 | Thyroid Left Site 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 4 | Thyroid Left Site 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 4 | Thyroid Left Site 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 4 | Thyroid Left Site 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 4 | Thyroid Left Site 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 4 | Thyroid Left Site 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 4 | Thyroid Left Site 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 4 | Thyroid Left Site 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 4 | Thyroid Left Site 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 4 | Thyroid Left Site 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 4 | Thyroid Left Site 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 4 | Thyroid Left Site 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 4 | Thyroid Left Site 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 4 | Thyroid Left Site 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 4 | Thyroid Left Site 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 4 | Thyroid Left Site 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 4 | Thyroid Left Site 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 4 | Thyroid Left Site 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 4 | Thyroid Left Site 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 4 | Thyroid Left Site 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 4 | Thyroid Left Site 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 4 | Thyroid Left Site 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 4 | Thyroid Left Site 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 4 | Thyroid Left Site 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 4 | Thyroid Left Site 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 5 | Thyroid Left Site 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 5 | Thyroid Left Site 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 5 | Thyroid Left Site 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 5 | Thyroid Left Site 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 5 | Thyroid Left Site 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 5 | Thyroid Left Site 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 5 | Thyroid Left Site 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 5 | Thyroid Left Site 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 5 | Thyroid Left Site 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 5 | Thyroid Left Site 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Thyroid Site 5 | Thyroid Left Site 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 5 | Thyroid Left Site 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 5 | Thyroid Left Site 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 5 | Thyroid Left Site 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 5 | Thyroid Left Site 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 5 | Thyroid Left Site 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 5 | Thyroid Left Site 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 5 | Thyroid Left Site 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 5 | Thyroid Left Site 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 5 | Thyroid Left Site 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Thyroid Site 5 | Thyroid Left Site 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 5 | Thyroid Left Site 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 5 | Thyroid Left Site 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 5 | Thyroid Left Site 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 5 | Thyroid Left Site 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 5 | Thyroid Left Site 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 5 | Thyroid Left Site 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 5 | Thyroid Left Site 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 5 | Thyroid Left Site 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Thyroid Site 5 | Thyroid Left Site 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 1 | Thyroid Right Site 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 1 | Thyroid Right Site 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 1 | Thyroid Right Site 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 1 | Thyroid Right Site 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 1 | Thyroid Right Site 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 1 | Thyroid Right Site 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 1 | Thyroid Right Site 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 1 | Thyroid Right Site 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 1 | Thyroid Right Site 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 1 | Thyroid Right Site 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 1 | Thyroid Right Site 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 1 | Thyroid Right Site 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 1 | Thyroid Right Site 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 1 | Thyroid Right Site 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 1 | Thyroid Right Site 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 1 | Thyroid Right Site 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 1 | Thyroid Right Site 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 1 | Thyroid Right Site 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 1 | Thyroid Right Site 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 1 | Thyroid Right Site 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 1 | Thyroid Right Site 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 1 | Thyroid Right Site 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 1 | Thyroid Right Site 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 1 | Thyroid Right Site 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 1 | Thyroid Right Site 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 1 | Thyroid Right Site 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 1 | Thyroid Right Site 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 1 | Thyroid Right Site 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 1 | Thyroid Right Site 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 1 | Thyroid Right Site 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 2 | Thyroid Right Site 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 2 | Thyroid Right Site 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 2 | Thyroid Right Site 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 2 | Thyroid Right Site 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 2 | Thyroid Right Site 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 2 | Thyroid Right Site 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 2 | Thyroid Right Site 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 2 | Thyroid Right Site 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 2 | Thyroid Right Site 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 2 | Thyroid Right Site 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 2 | Thyroid Right Site 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 2 | Thyroid Right Site 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 2 | Thyroid Right Site 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 2 | Thyroid Right Site 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 2 | Thyroid Right Site 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 2 | Thyroid Right Site 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 2 | Thyroid Right Site 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 2 | Thyroid Right Site 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 2 | Thyroid Right Site 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 2 | Thyroid Right Site 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 2 | Thyroid Right Site 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 2 | Thyroid Right Site 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 2 | Thyroid Right Site 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 2 | Thyroid Right Site 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 2 | Thyroid Right Site 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 2 | Thyroid Right Site 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 2 | Thyroid Right Site 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 2 | Thyroid Right Site 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 2 | Thyroid Right Site 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 2 | Thyroid Right Site 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 3 | Thyroid Right Site 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 3 | Thyroid Right Site 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 3 | Thyroid Right Site 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 3 | Thyroid Right Site 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 3 | Thyroid Right Site 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 3 | Thyroid Right Site 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 3 | Thyroid Right Site 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 3 | Thyroid Right Site 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 3 | Thyroid Right Site 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 3 | Thyroid Right Site 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 3 | Thyroid Right Site 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 3 | Thyroid Right Site 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 3 | Thyroid Right Site 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 3 | Thyroid Right Site 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 3 | Thyroid Right Site 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 3 | Thyroid Right Site 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 3 | Thyroid Right Site 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 3 | Thyroid Right Site 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 3 | Thyroid Right Site 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 3 | Thyroid Right Site 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 3 | Thyroid Right Site 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 3 | Thyroid Right Site 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 3 | Thyroid Right Site 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 3 | Thyroid Right Site 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 3 | Thyroid Right Site 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 3 | Thyroid Right Site 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 3 | Thyroid Right Site 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 3 | Thyroid Right Site 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 3 | Thyroid Right Site 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 3 | Thyroid Right Site 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 4 | Thyroid Right Site 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 4 | Thyroid Right Site 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 4 | Thyroid Right Site 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 4 | Thyroid Right Site 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 4 | Thyroid Right Site 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 4 | Thyroid Right Site 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 4 | Thyroid Right Site 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 4 | Thyroid Right Site 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 4 | Thyroid Right Site 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 4 | Thyroid Right Site 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 4 | Thyroid Right Site 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 4 | Thyroid Right Site 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 4 | Thyroid Right Site 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 4 | Thyroid Right Site 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 4 | Thyroid Right Site 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 4 | Thyroid Right Site 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 4 | Thyroid Right Site 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 4 | Thyroid Right Site 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 4 | Thyroid Right Site 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 4 | Thyroid Right Site 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 4 | Thyroid Right Site 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 4 | Thyroid Right Site 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 4 | Thyroid Right Site 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 4 | Thyroid Right Site 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 4 | Thyroid Right Site 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 4 | Thyroid Right Site 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 4 | Thyroid Right Site 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 4 | Thyroid Right Site 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 4 | Thyroid Right Site 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 4 | Thyroid Right Site 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 5 | Thyroid Right Site 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 5 | Thyroid Right Site 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 5 | Thyroid Right Site 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 5 | Thyroid Right Site 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 5 | Thyroid Right Site 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 5 | Thyroid Right Site 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 5 | Thyroid Right Site 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 5 | Thyroid Right Site 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 5 | Thyroid Right Site 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 5 | Thyroid Right Site 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Thyroid Site 5 | Thyroid Right Site 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 5 | Thyroid Right Site 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 5 | Thyroid Right Site 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 5 | Thyroid Right Site 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 5 | Thyroid Right Site 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 5 | Thyroid Right Site 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 5 | Thyroid Right Site 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 5 | Thyroid Right Site 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 5 | Thyroid Right Site 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 5 | Thyroid Right Site 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Thyroid Site 5 | Thyroid Right Site 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 5 | Thyroid Right Site 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 5 | Thyroid Right Site 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 5 | Thyroid Right Site 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 5 | Thyroid Right Site 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 5 | Thyroid Right Site 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 5 | Thyroid Right Site 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 5 | Thyroid Right Site 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 5 | Thyroid Right Site 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Thyroid Site 5 | Thyroid Right Site 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 1 | Prostate Left Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 1 | Prostate Left Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 1 | Prostate Left Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 1 | Prostate Left Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 1 | Prostate Left Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 1 | Prostate Left Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 1 | Prostate Left Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 1 | Prostate Left Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 1 | Prostate Left Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 1 | Prostate Left Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 1 | Prostate Left Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 1 | Prostate Left Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 1 | Prostate Left Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 1 | Prostate Left Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 1 | Prostate Left Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 1 | Prostate Left Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 1 | Prostate Left Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 1 | Prostate Left Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 1 | Prostate Left Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 1 | Prostate Left Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 1 | Prostate Left Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 1 | Prostate Left Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 1 | Prostate Left Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 1 | Prostate Left Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 1 | Prostate Left Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 1 | Prostate Left Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 1 | Prostate Left Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 1 | Prostate Left Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 1 | Prostate Left Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 1 | Prostate Left Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 2 | Prostate Left Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 2 | Prostate Left Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 2 | Prostate Left Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 2 | Prostate Left Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 2 | Prostate Left Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 2 | Prostate Left Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 2 | Prostate Left Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 2 | Prostate Left Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 2 | Prostate Left Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 2 | Prostate Left Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 2 | Prostate Left Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 2 | Prostate Left Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 2 | Prostate Left Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 2 | Prostate Left Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 2 | Prostate Left Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 2 | Prostate Left Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 2 | Prostate Left Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 2 | Prostate Left Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 2 | Prostate Left Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 2 | Prostate Left Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 2 | Prostate Left Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 2 | Prostate Left Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 2 | Prostate Left Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 2 | Prostate Left Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 2 | Prostate Left Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 2 | Prostate Left Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 2 | Prostate Left Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 2 | Prostate Left Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 2 | Prostate Left Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 2 | Prostate Left Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 3 | Prostate Left Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 3 | Prostate Left Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 3 | Prostate Left Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 3 | Prostate Left Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 3 | Prostate Left Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 3 | Prostate Left Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 3 | Prostate Left Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 3 | Prostate Left Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 3 | Prostate Left Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 3 | Prostate Left Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 3 | Prostate Left Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 3 | Prostate Left Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 3 | Prostate Left Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 3 | Prostate Left Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 3 | Prostate Left Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 3 | Prostate Left Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 3 | Prostate Left Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 3 | Prostate Left Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 3 | Prostate Left Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 3 | Prostate Left Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 3 | Prostate Left Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 3 | Prostate Left Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 3 | Prostate Left Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 3 | Prostate Left Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 3 | Prostate Left Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 3 | Prostate Left Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 3 | Prostate Left Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 3 | Prostate Left Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 3 | Prostate Left Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 3 | Prostate Left Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 4 | Prostate Left Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 4 | Prostate Left Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 4 | Prostate Left Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 4 | Prostate Left Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 4 | Prostate Left Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 4 | Prostate Left Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 4 | Prostate Left Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 4 | Prostate Left Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 4 | Prostate Left Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 4 | Prostate Left Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 4 | Prostate Left Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 4 | Prostate Left Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 4 | Prostate Left Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 4 | Prostate Left Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 4 | Prostate Left Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 4 | Prostate Left Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 4 | Prostate Left Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 4 | Prostate Left Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 4 | Prostate Left Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 4 | Prostate Left Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 4 | Prostate Left Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 4 | Prostate Left Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 4 | Prostate Left Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 4 | Prostate Left Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 4 | Prostate Left Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 4 | Prostate Left Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 4 | Prostate Left Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 4 | Prostate Left Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 4 | Prostate Left Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 4 | Prostate Left Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 5 | Prostate Left Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 5 | Prostate Left Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 5 | Prostate Left Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 5 | Prostate Left Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 5 | Prostate Left Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 5 | Prostate Left Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 5 | Prostate Left Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 5 | Prostate Left Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 5 | Prostate Left Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 5 | Prostate Left Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Lt Mass 5 | Prostate Left Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 5 | Prostate Left Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 5 | Prostate Left Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 5 | Prostate Left Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 5 | Prostate Left Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 5 | Prostate Left Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 5 | Prostate Left Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 5 | Prostate Left Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 5 | Prostate Left Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 5 | Prostate Left Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Lt Mass 5 | Prostate Left Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 5 | Prostate Left Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 5 | Prostate Left Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 5 | Prostate Left Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 5 | Prostate Left Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 5 | Prostate Left Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 5 | Prostate Left Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 5 | Prostate Left Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 5 | Prostate Left Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Lt Mass 5 | Prostate Left Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 1 | Prostate Right Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 1 | Prostate Right Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 1 | Prostate Right Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 1 | Prostate Right Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 1 | Prostate Right Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 1 | Prostate Right Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 1 | Prostate Right Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 1 | Prostate Right Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 1 | Prostate Right Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 1 | Prostate Right Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 1 | Prostate Right Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 1 | Prostate Right Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 1 | Prostate Right Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 1 | Prostate Right Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 1 | Prostate Right Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 1 | Prostate Right Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 1 | Prostate Right Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 1 | Prostate Right Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 1 | Prostate Right Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 1 | Prostate Right Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 1 | Prostate Right Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 1 | Prostate Right Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 1 | Prostate Right Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 1 | Prostate Right Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 1 | Prostate Right Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 1 | Prostate Right Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 1 | Prostate Right Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 1 | Prostate Right Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 1 | Prostate Right Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 1 | Prostate Right Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 2 | Prostate Right Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 2 | Prostate Right Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 2 | Prostate Right Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 2 | Prostate Right Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 2 | Prostate Right Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 2 | Prostate Right Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 2 | Prostate Right Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 2 | Prostate Right Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 2 | Prostate Right Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 2 | Prostate Right Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 2 | Prostate Right Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 2 | Prostate Right Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 2 | Prostate Right Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 2 | Prostate Right Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 2 | Prostate Right Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 2 | Prostate Right Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 2 | Prostate Right Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 2 | Prostate Right Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 2 | Prostate Right Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 2 | Prostate Right Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 2 | Prostate Right Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 2 | Prostate Right Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 2 | Prostate Right Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 2 | Prostate Right Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 2 | Prostate Right Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 2 | Prostate Right Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 2 | Prostate Right Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 2 | Prostate Right Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 2 | Prostate Right Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 2 | Prostate Right Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 3 | Prostate Right Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 3 | Prostate Right Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 3 | Prostate Right Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 3 | Prostate Right Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 3 | Prostate Right Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 3 | Prostate Right Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 3 | Prostate Right Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 3 | Prostate Right Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 3 | Prostate Right Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 3 | Prostate Right Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 3 | Prostate Right Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 3 | Prostate Right Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 3 | Prostate Right Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 3 | Prostate Right Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 3 | Prostate Right Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 3 | Prostate Right Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 3 | Prostate Right Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 3 | Prostate Right Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 3 | Prostate Right Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 3 | Prostate Right Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 3 | Prostate Right Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 3 | Prostate Right Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 3 | Prostate Right Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 3 | Prostate Right Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 3 | Prostate Right Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 3 | Prostate Right Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 3 | Prostate Right Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 3 | Prostate Right Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 3 | Prostate Right Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 3 | Prostate Right Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 4 | Prostate Right Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 4 | Prostate Right Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 4 | Prostate Right Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 4 | Prostate Right Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 4 | Prostate Right Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 4 | Prostate Right Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 4 | Prostate Right Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 4 | Prostate Right Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 4 | Prostate Right Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 4 | Prostate Right Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 4 | Prostate Right Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 4 | Prostate Right Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 4 | Prostate Right Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 4 | Prostate Right Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 4 | Prostate Right Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 4 | Prostate Right Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 4 | Prostate Right Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 4 | Prostate Right Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 4 | Prostate Right Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 4 | Prostate Right Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 4 | Prostate Right Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 4 | Prostate Right Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 4 | Prostate Right Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 4 | Prostate Right Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 4 | Prostate Right Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 4 | Prostate Right Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 4 | Prostate Right Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 4 | Prostate Right Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 4 | Prostate Right Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 4 | Prostate Right Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 5 | Prostate Right Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 5 | Prostate Right Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 5 | Prostate Right Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 5 | Prostate Right Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 5 | Prostate Right Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 5 | Prostate Right Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 5 | Prostate Right Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 5 | Prostate Right Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 5 | Prostate Right Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 5 | Prostate Right Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Prostate Rt Mass 5 | Prostate Right Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 5 | Prostate Right Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 5 | Prostate Right Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 5 | Prostate Right Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 5 | Prostate Right Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 5 | Prostate Right Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 5 | Prostate Right Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 5 | Prostate Right Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 5 | Prostate Right Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 5 | Prostate Right Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Prostate Rt Mass 5 | Prostate Right Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 5 | Prostate Right Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 5 | Prostate Right Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 5 | Prostate Right Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 5 | Prostate Right Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 5 | Prostate Right Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 5 | Prostate Right Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 5 | Prostate Right Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 5 | Prostate Right Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Prostate Rt Mass 5 | Prostate Right Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Overall Statistics: Liver | Overall Liver Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Overall Interquartile Range of Elasticity Medians;99SIEMENS;OIQREMedians | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Liver | Overall Liver Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Overall IQR/Median Ratio of Elasticity Medians;99SIEMENS;OIQRMdEMedians | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Liver | Overall Liver Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Overall Standard Deviation of Elasticity Medians;99SIEMENS;OStdDevEMedians | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Liver | Overall Liver Elasticity Mean | Mean | kPa | 1 | 1 | Overall Mean of Elasticity Medians;99SIEMENS;OMeanEMedians | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Liver | Overall Liver Elasticity Median | Median | kPa | 1 | 1 | Overall Median of Elasticity Medians;99SIEMENS;OMedianEMedians | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Liver | Overall Liver Interquartile Range of Shear Wave Velocity Medians | IQR | m/s | 2 | 1 | Overall Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;OIQRVsMedians | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Liver | Overall Liver IQR/Median of Shear Wave Velocity Medians | IQR/Median | unitless | 2 | 1 | Overall IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;OIQRMdVsMedians | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Liver | Overall Liver Standard Deviation of Shear Wave Velocity Medians | Std Dev | m/s | 2 | 1 | Overall Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;OStdDevVsMedians | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Liver | Overall Liver Mean of Shear Wave Velocity Medians | Mean | m/s | 2 | 1 | Overall Mean of Shear Wave Velocity Medians;99SIEMENS;OMeanVsMedians | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Liver | Overall Liver Median of Shear Wave Velocity Medians | Median | m/s | 2 | 1 | Overall Median of Shear Wave Velocity Medians;99SIEMENS;OMedianVsMedians | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Prostate | Overall Prostate Interquartile Range of Shear Wave Velocity Medians | IQR | m/s | 2 | 1 | Overall Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;OIQRVsMedians | NUM | Prostate;SRT;T-92000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Prostate | Overall Prostate IQR/Median of Shear Wave Velocity Medians | IQR/Median | unitless | 2 | 1 | Overall IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;OIQRMdVsMedians | NUM | Prostate;SRT;T-92000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Prostate | Overall Prostate Standard Deviation of Shear Wave Velocity Medians | Std Dev | m/s | 2 | 1 | Overall Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;OStdDevVsMedians | NUM | Prostate;SRT;T-92000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Prostate | Overall Prostate Mean of Shear Wave Velocity Medians | Mean | m/s | 2 | 1 | Overall Mean of Shear Wave Velocity Medians;99SIEMENS;OMeanVsMedians | NUM | Prostate;SRT;T-92000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Prostate | Overall Prostate Median of Shear Wave Velocity Medians | Median | m/s | 2 | 1 | Overall Median of Shear Wave Velocity Medians;99SIEMENS;OMedianVsMedians | NUM | Prostate;SRT;T-92000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Prostate | Overall Prostate Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Overall Interquartile Range of Elasticity Medians;99SIEMENS;OIQREMedians | NUM | Prostate;SRT;T-92000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Prostate | Overall Prostate Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Overall IQR/Median Ratio of Elasticity Medians;99SIEMENS;OIQRMdEMedians | NUM | Prostate;SRT;T-92000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Prostate | Overall Prostate Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Overall Standard Deviation of Elasticity Medians;99SIEMENS;OStdDevEMedians | NUM | Prostate;SRT;T-92000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Prostate | Overall Prostate Elasticity Mean | Mean | kPa | 1 | 1 | Overall Mean of Elasticity Medians;99SIEMENS;OMeanEMedians | NUM | Prostate;SRT;T-92000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Prostate | Overall Prostate Elasticity Median | Median | kPa | 1 | 1 | Overall Median of Elasticity Medians;99SIEMENS;OMedianEMedians | NUM | Prostate;SRT;T-92000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Thyroid | Overall Thyroid Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Overall Interquartile Range of Elasticity Medians;99SIEMENS;OIQREMedians | NUM | Thyroid;SRT;T-B6000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Thyroid | Overall Thyroid Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Overall IQR/Median Ratio of Elasticity Medians;99SIEMENS;OIQRMdEMedians | NUM | Thyroid;SRT;T-B6000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Thyroid | Overall Thyroid Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Overall Standard Deviation of Elasticity Medians;99SIEMENS;OStdDevEMedians | NUM | Thyroid;SRT;T-B6000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Thyroid | Overall Thyroid Elasticity Mean | Mean | kPa | 1 | 1 | Overall Mean of Elasticity Medians;99SIEMENS;OMeanEMedians | NUM | Thyroid;SRT;T-B6000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Thyroid | Overall Thyroid Elasticity Median | Median | kPa | 1 | 1 | Overall Median of Elasticity Medians;99SIEMENS;OMedianEMedians | NUM | Thyroid;SRT;T-B6000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Thyroid | Overall Thyroid Interquartile Range of Shear Wave Velocity Medians | IQR | m/s | 2 | 1 | Overall Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;OIQRVsMedians | NUM | Thyroid;SRT;T-B6000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Thyroid | Overall Thyroid IQR/Median of Shear Wave Velocity Medians | IQR/Median | unitless | 2 | 1 | Overall IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;OIQRMdVsMedians | NUM | Thyroid;SRT;T-B6000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Thyroid | Overall Thyroid Standard Deviation of Shear Wave Velocity Medians | Std Dev | m/s | 2 | 1 | Overall Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;OStdDevVsMedians | NUM | Thyroid;SRT;T-B6000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Thyroid | Overall Thyroid Mean of Shear Wave Velocity Medians | Mean | m/s | 2 | 1 | Overall Mean of Shear Wave Velocity Medians;99SIEMENS;OMeanVsMedians | NUM | Thyroid;SRT;T-B6000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Thyroid | Overall Thyroid Median of Shear Wave Velocity Medians | Median | m/s | 2 | 1 | Overall Median of Shear Wave Velocity Medians;99SIEMENS;OMedianVsMedians | NUM | Thyroid;SRT;T-B6000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Liver Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Liver Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Liver Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Liver Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Liver Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Liver Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Liver Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Liver Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Liver Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Liver Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Liver Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Liver Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Liver Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Liver Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Liver Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Liver Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Liver Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Liver Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Liver Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Liver Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Liver Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Liver Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Liver Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Liver Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Liver Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Liver Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Liver Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Liver Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Liver Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Liver Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Liver Mass 6 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Liver Mass 6 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Liver Mass 6 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Liver Mass 6 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Liver Mass 6 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Liver Mass 6 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Liver Mass 7 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Liver Mass 7 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Liver Mass 7 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Liver Mass 7 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Liver Mass 7 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Liver Mass 7 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Liver Mass 8 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Liver Mass 8 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Liver Mass 8 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Liver Mass 8 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Liver Mass 8 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Liver Mass 8 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Liver Mass 9 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Liver Mass 9 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Liver Mass 9 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Liver Mass 9 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Liver Mass 9 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Liver Mass 9 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Liver Mass 10 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Liver Mass 10 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Liver Mass 10 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Liver Mass 10 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Liver Mass 10 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Liver Mass 10 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Liver Site 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Liver Site 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Liver Site 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Liver Site 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Liver Site 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Liver Site 1 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Liver Site 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Liver Site 1 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Liver Site 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Liver Site 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Liver Site 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Liver Site 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Liver Site 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Liver Site 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Liver Site 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Liver Site 2 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Liver Site 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Liver Site 2 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Liver Site 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Liver Site 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Liver Site 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Liver Site 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Liver Site 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Liver Site 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Liver Site 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Liver Site 3 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Liver Site 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Liver Site 3 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Liver Site 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Liver Site 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Liver Site 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Liver Site 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Liver Site 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Liver Site 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Liver Site 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Liver Site 4 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Liver Site 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Liver Site 4 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Liver Site 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Liver Site 4 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Liver Site 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Liver Site 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Liver Site 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Liver Site 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Liver Site 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Liver Site 5 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Liver Site 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Liver Site 5 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Liver Site 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Liver Site 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Liver Site 6 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Liver Site 6 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Liver Site 6 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Liver Site 6 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Liver Site 6 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Liver Site 6 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Liver Site 6 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Liver Site 6 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Liver Site 6 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Liver Site 6 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Liver Site 7 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Liver Site 7 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Liver Site 7 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Liver Site 7 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Liver Site 7 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Liver Site 7 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Liver Site 7 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Liver Site 7 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Liver Site 7 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Liver Site 7 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Liver Site 8 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Liver Site 8 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Liver Site 8 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Liver Site 8 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Liver Site 8 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Liver Site 8 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Liver Site 8 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Liver Site 8 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Liver Site 8 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Liver Site 8 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Liver Site 9 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Liver Site 9 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Liver Site 9 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Liver Site 9 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Liver Site 9 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Liver Site 9 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Liver Site 9 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Liver Site 9 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Liver Site 9 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Liver Site 9 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Liver Site 10 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Liver Site 10 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Liver Site 10 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Liver Site 10 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Liver Site 10 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Liver Site 10 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Liver Site 10 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Liver Site 10 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Liver Site 10 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Liver Site 10 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Liver Segment 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Liver Segment 1 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Liver Segment 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Liver Segment 1 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Liver Segment 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Liver Segment 1 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Liver Segment 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Liver Segment 2 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Liver Segment 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Liver Segment 2 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Liver Segment 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Liver Segment 2 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Liver Segment 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Liver Segment 3 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Liver Segment 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Liver Segment 3 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Liver Segment 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Liver Segment 3 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Liver Segment 4a Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Liver Segment 4a Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Liver Segment 4a Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Liver Segment 4a Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Liver Segment 4a Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Liver Segment 4a Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Liver Segment 4b Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Liver Segment 4b Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Liver Segment 4b Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Liver Segment 4b Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Liver Segment 4b Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Liver Segment 4b Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Liver Segment 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Liver Segment 5 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Liver Segment 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Liver Segment 5 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Liver Segment 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Liver Segment 5 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Liver Segment 6 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Liver Segment 6 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Liver Segment 6 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Liver Segment 6 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Liver Segment 6 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Liver Segment 6 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Liver Segment 7 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Liver Segment 7 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Liver Segment 7 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Liver Segment 7 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Liver Segment 7 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Liver Segment 7 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Shear Wave Velocity Mean;99SIEMENS;VsMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity;99SIEMENS;VsStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Shear Wave Velocity Interquartile Range;99SIEMENS;VsIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Liver Segment 8 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | Shear Wave Velocity IQR/Median Ratio;99SIEMENS;VsIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Liver Segment 8 Elasticity Mean | Mean | kPa | 1 | 1 | Elasticity Mean;99SIEMENS;EMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Liver Segment 8 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity;99SIEMENS;EStdDev | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Liver Segment 8 Elasticity Median | Median | kPa | 1 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Liver Segment 8 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Elasticity Interquartile Range;99SIEMENS;EIQR | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Liver Segment 8 Elasticity IQR/Median | IQR/Median | unitless | 2 | 1 | Elasticity IQR/Median Ratio;99SIEMENS;EIQRMRatio | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Overall Statistics: Liver | Overall Statistics: Liver Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Overall Shear Velocity Mean;99SIEMENS;OVsMean | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Overall Statistics: Liver | Overall Statistics: Liver Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Overall Shear Velocity Standard Deviation;99SIEMENS;OVsStdDev | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Overall Statistics: Liver | Overall Statistics: Liver Shear Wave Velocity Median | Median | m/s | 2 | 1 | Overall Shear Velocity Median;99SIEMENS;OVsMedian | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Overall Statistics: Liver | Overall Statistics: Liver Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Overall Shear Velocity Interquartile Range;99SIEMENS;OVsIQR | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Overall Statistics: Liver | Overall Statistics: Liver Shear Wave Velocity IQR/Median Ratio | IQR/Median | unitless | 2 | 1 | Overall Shear Velocity IQR/Median Ratio;99SIEMENS;OVsIQRMednRatio | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Overall Statistics: Liver | Overall Statistics: Liver Elasticity Mean | Mean | kPa | 1 | 1 | Overall Elasticity Mean;99SIEMENS;OElastictyMean | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Overall Statistics: Liver | Overall Statistics: Liver Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Overall Elasticity Standard Deviation;99SIEMENS;OElastictyStdDev | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Overall Statistics: Liver | Overall Statistics: Liver Elasticity Median | Median | kPa | 1 | 1 | Overall Elasticity Median;99SIEMENS;OElastictyMedian | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Overall Statistics: Liver | Overall Statistics: Liver Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Overall Elasticity Interquartile Range;99SIEMENS;OElastictyIQR | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Overall Statistics: Liver | Overall Statistics: Liver Elasticity IQR/Median Ratio | IQR/Median | unitless | 2 | 1 | Overall Elasticity IQR/Median Ratio;99SIEMENS;OEIQRMednRatio | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Liver Mass 1 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Liver Mass 1 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Liver Mass 1 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Liver Mass 1 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Liver Mass 1 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Liver Mass 2 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Liver Mass 2 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Liver Mass 2 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Liver Mass 2 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Liver Mass 2 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Liver Mass 3 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Liver Mass 3 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Liver Mass 3 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Liver Mass 3 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Liver Mass 3 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Liver Mass 4 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Liver Mass 4 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Liver Mass 4 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Liver Mass 4 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Liver Mass 4 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Liver Mass 5 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Liver Mass 5 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Liver Mass 5 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Liver Mass 5 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Liver Mass 5 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Liver Mass 6 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Liver Mass 6 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Liver Mass 6 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Liver Mass 6 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Liver Mass 6 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Liver Mass 7 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Liver Mass 7 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Liver Mass 7 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Liver Mass 7 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Liver Mass 7 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Liver Mass 8 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Liver Mass 8 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Liver Mass 8 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Liver Mass 8 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Liver Mass 8 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Liver Mass 9 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Liver Mass 9 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Liver Mass 9 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Liver Mass 9 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Liver Mass 9 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Liver Mass 10 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Liver Mass 10 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Liver Mass 10 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Liver Mass 10 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Liver Mass 10 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Liver Site 1 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Liver Site 1 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Liver Site 1 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Liver Site 1 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Liver Site 1 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Liver Site 2 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Liver Site 2 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Liver Site 2 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Liver Site 2 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Liver Site 2 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Liver Site 3 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Liver Site 3 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Liver Site 3 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Liver Site 3 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Liver Site 3 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Liver Site 4 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Liver Site 4 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Liver Site 4 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Liver Site 4 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Liver Site 4 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Liver Site 5 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Liver Site 5 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Liver Site 5 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Liver Site 5 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Liver Site 5 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Liver Site 6 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Liver Site 6 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Liver Site 6 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Liver Site 6 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Liver Site 6 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Liver Site 7 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Liver Site 7 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Liver Site 7 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Liver Site 7 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Liver Site 7 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Liver Site 8 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Liver Site 8 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Liver Site 8 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Liver Site 8 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Liver Site 8 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Liver Site 9 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Liver Site 9 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Liver Site 9 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Liver Site 9 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Liver Site 9 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Liver Site 10 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Liver Site 10 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Liver Site 10 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Liver Site 10 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Liver Site 10 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Liver Segment 1 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Liver Segment 1 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Liver Segment 1 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Liver Segment 1 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Liver Segment 1 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Liver Segment 2 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Liver Segment 2 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Liver Segment 2 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Liver Segment 2 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Liver Segment 2 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Liver Segment 3 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Liver Segment 3 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Liver Segment 3 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Liver Segment 3 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Liver Segment 3 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Liver Segment 4a UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Liver Segment 4a UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Liver Segment 4a UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Liver Segment 4a UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Liver Segment 4a UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Liver Segment 4b UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Liver Segment 4b UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Liver Segment 4b UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Liver Segment 4b UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Liver Segment 4b UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Liver Segment 5 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Liver Segment 5 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Liver Segment 5 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Liver Segment 5 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Liver Segment 5 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Liver Segment 6 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Liver Segment 6 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Liver Segment 6 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Liver Segment 6 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Liver Segment 6 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Liver Segment 7 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Liver Segment 7 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Liver Segment 7 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Liver Segment 7 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Liver Segment 7 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Liver Segment 8 UDFF Interquartile Range | IQR | % | 1 | 1 | Interquartile Range of UDFF;99SIEMENS;IQRUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Liver Segment 8 UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of UDFF;99SIEMENS;IQRMdUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Liver Segment 8 UDFF Standard Deviation | Std Dev | % | 1 | 1 | Standard Deviation of UDFF;99SIEMENS;StdDevUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Liver Segment 8 UDFF Mean | Mean | % | 1 | 1 | Mean of UDFF;99SIEMENS;MeanUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Liver Segment 8 UDFF Median | Median | % | 1 | 1 | Median of UDFF;99SIEMENS;MedianUDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Overall Statistics: Liver | Overall Liver UDFFInterquartile Range | IQR | % | 1 | 1 | Overall UDFF Interquartile Range of UDFF;99SIEMENS;OVIQRUDFF | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Overall Statistics: Liver | Overall Liver UDFF IQR/Median | IQR/Median | unitless | 2 | 1 | Overall UDFF IQR/Median Ratio of UDFF;99SIEMENS;OVIQRMdUDFF | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Overall Statistics: Liver | Overall Liver UDFF Standard Deviation | Std Dev | % | 1 | 1 | Overall UDFF Standard Deviation of UDFF;99SIEMENS;OVStdDevUDFF | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Overall Statistics: Liver | Overall Liver UDFF Mean | Mean | % | 1 | 1 | Overall UDFF Mean of UDFF;99SIEMENS;OVMeanUDFF | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Overall Statistics: Liver | Overall Liver UDFF Median | Median | % | 1 | 1 | Overall UDFF Median of UDFF;99SIEMENS;OVMedianUDFF | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE

## OB Gyn Measurements
EnglishLabel | Description | DetailEnglishLabel | PossibleUnits | MaxInstances | DefaultPrecision | ConceptCoding | ValueType | ParentContainer | ParentLaterality | SRSection | SRSectionLaterality | AdditionalModifiers | AdditionalModifiers2
AC | Abdominal Circumference |  | cm mm | 5 | 2 | Abdominal Circumference;LN;11979-2 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Amniotic Fluid Volume | Fetal Biophysical Profile: Amniotic Fluid Volume |  | unitless | 1 | 0 | Amniotic Fluid Volume;LN;11630-1 |  | Biophysical Profile;DCM;125006 |  | Biophysical Profile;DCM;125006
LV-A | Anterior Horn Lateral ventricular width |  | cm mm | 5 | 2 | Anterior Horn Lateral ventricular width;LN;33197-5 | NUM | Biometry Group;DCM;125005 |  | Fetal Cranium;DCM;125004
APAD | Anterior-Posterior Abdominal Diameter |  | cm mm | 5 | 2 | Anterior-Posterior Abdominal Diameter;LN;11818-2 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
APTD | Anterior-Posterior Trunk Diameter |  | cm mm | 5 | 2 | Anterior-Posterior Trunk Diameter;LN;11819-0 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Lt Axillary LN 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Rt Axillary LN 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Lt Axillary LN 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Rt Axillary LN 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Lt Axillary LN 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Rt Axillary LN 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Lt Axillary LN 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Rt Axillary LN 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Lt Axillary LN 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Rt Axillary LN 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Lt Axillary LN 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Rt Axillary LN 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Lt Axillary LN 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Rt Axillary LN 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Lt Axillary LN 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Rt Axillary LN 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Lt Axillary LN 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Rt Axillary LN 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Lt Axillary LN 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Rt Axillary LN 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Lt Axillary LN 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Rt Axillary LN 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Lt Axillary LN 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Rt Axillary LN 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Lt Axillary LN 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Rt Axillary LN 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Lt Axillary LN 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Rt Axillary LN 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Lt Axillary LN 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Rt Axillary LN 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Binoc D | Binocular Distance |  | cm mm | 5 | 2 | Outer Orbital Diameter;LN;11629-3 | NUM | Biometry Group;DCM;125005 |  | Fetal Cranium;DCM;125004
BPD | Biparietal Diameter |  | cm mm | 5 | 2 | Biparietal Diameter;LN;11820-8 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Bladder neck-Symphysis Distance | Bladder neck-Symphysis Distance |  | cm mm | 5 | 2 | Bladder neck-Symphysis Distance;99SIEMENS;BldrNckSymphDist | NUM | Pelvic Floor;SRT;T-D6236 |  | Finding Site;SRT;G-C0E3=Pelvic Floor;SRT;T-D6236
Lt Breast Cyst 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Rt Breast Cyst 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Lt Breast Cyst 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Rt Breast Cyst 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Lt Breast Cyst 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Rt Breast Cyst 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Lt Breast Cyst 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Rt Breast Cyst 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Lt Breast Cyst 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Rt Breast Cyst 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Lt Breast Cyst 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Rt Breast Cyst 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Lt Breast Cyst 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Rt Breast Cyst 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Lt Breast Cyst 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Rt Breast Cyst 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Lt Breast Cyst 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Rt Breast Cyst 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Lt Breast Cyst 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Rt Breast Cyst 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Lt Breast Cyst 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Rt Breast Cyst 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Lt Breast Cyst 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Rt Breast Cyst 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Lt Breast Cyst 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Rt Breast Cyst 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Lt Breast Cyst 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Rt Breast Cyst 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Lt Breast Cyst 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Rt Breast Cyst 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
CSP | Cavum Septum Pellucidum |  | cm mm | 5 | 2 | Cavum septi pellucidi;SRT;T-A1630 | NUM | Biometry Group;DCM;125005 |  | Fetal Cranium;DCM;125004
Lt MCA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Embryonic Vascular Structure;SRT;T-F6800
Rt MCA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Embryonic Vascular Structure;SRT;T-F6800
Cervix (TA) | Cervix Length (Transabdominal) |  | cm mm | 5 | 2 | Cervix Length;LN;11961-0 | NUM | Pelvis and Uterus;DCM;125011 |  | Finding Site;SRT;G-C0E3=Pelvis and Uterus;DCM;125011
Cervix (EV) | Cervix Length (Endovaginal) |  | cm mm | 5 | 2 | Cervix Length;LN;11961-0 | NUM | Pelvis and Uterus;DCM;125011 |  | Finding Site;SRT;G-C0E3=Pelvis and Uterus;DCM;125011 |  | Imaging procedure;SRT;P0-0099A=Transvaginal echography;SRT;P5-B8012
Cist Magna | Cisterna Magna length |  | cm mm | 5 | 2 | Cisterna Magna length;LN;11860-4 | NUM | Biometry Group;DCM;125005 |  | Fetal Cranium;DCM;125004
Clavicle | Clavicle Length |  | cm mm | 5 | 2 | Clavicle length;LN;11962-8 | NUM | Biometry Group;DCM;125005 |  | Fetal Long Bones;DCM;125003
CRL | Crown Rump Length |  | cm mm | 5 | 2 | Crown Rump Length;LN;11957-8 | NUM | Biometry Group;DCM;125005 |  | Early Gestation;DCM;125009
Lt Cyst 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Cyst 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Cyst 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Cyst 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Cyst 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Cyst 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Cyst 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Cyst 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Cyst 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Cyst 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Cyst 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Cyst 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Cyst 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Cyst 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Cyst 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Cyst 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Cyst 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Cyst 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Cyst 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Cyst 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Cyst 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Cyst 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Cyst 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Cyst 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Cyst 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Cyst 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Cyst 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Cyst 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Cyst 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Cyst 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Ductus V A | Atrial Systole Vmax | Vmax | cm/s m/s | 5 | 2 | A-Wave Peak Velocity;LN;59081-0 | NUM | Ductus Venosus;SRT;T-F6806 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Ductus V A | Doppler Angle Property of Atrial Systole Vmax | Angle | ° | 5 | 0 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Ductus V D | Diastole Vmax | Vmax | cm/s m/s | 5 | 2 | Diastolic blood velocity, peak;DCM;122202 | NUM | Ductus Venosus;SRT;T-F6806 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Ductus V D | Doppler Angle Property of Diastole Vmax | Angle | ° | 5 | 0 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Ductus V S | Systole Vmax | Vmax | cm/s m/s | 5 | 2 | Systolic blood velocity, peak;DCM;122204 | NUM | Ductus Venosus;SRT;T-F6806 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Ductus V S | Doppler Angle Property of Systole Vmax | Angle | ° | 5 | 0 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Endo Thick (TA) | Endometrium Thickness (Transabdominal) |  | cm mm | 5 | 2 | Endometrium Thickness;LN;12145-9 | NUM | Pelvis and Uterus;DCM;125011 |  | Finding Site;SRT;G-C0E3=Pelvis and Uterus;DCM;125011
Endo Thick (EV) | Endometrium Thickness (Endovaginal) |  | cm mm | 5 | 2 | Endometrium Thickness;LN;12145-9 | NUM | Pelvis and Uterus;DCM;125011 |  | Finding Site;SRT;G-C0E3=Pelvis and Uterus;DCM;125011 |  | Imaging procedure;SRT;P0-0099A=Transvaginal echography;SRT;P5-B8012
FL | Femur Length |  | cm mm | 5 | 2 | Femur Length;LN;11963-6 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Aorta | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Aorta | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Aorta | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Aorta | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Aorta | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Aorta | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Aorta | Doppler Angle Property of PSV, EDV, TAMx | Angle | ° | 5 | 0 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Aorta | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Aorta | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Fetal Respiration | Fetal Biophysical Profile: Fetal Breathing |  | unitless | 1 | 0 | Fetal Breathing;LN;11632-7 |  | Biophysical Profile;DCM;125006 |  | Biophysical Profile;DCM;125006
Facial Angle | Facial Angle |  | ° | 5 | 0 | Facial Angle;99SIEMENS;FacialAngle | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Fetal HR | Fetal Heart Rate |  | bpm | 5 | 0 | Fetal Heart Rate;LN;11948-7 | NUM | Fetus Summary;DCM;125008 |  | Summary;DCM;121111
Lt Fetal Kidney | Fetal Left Kidney Thickness | AP | cm mm | 5 | 2 | Left Kidney thickness;LN;11853-9 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Rt Fetal Kidney | Fetal Right Kidney Thickness | AP | cm mm | 5 | 2 | Right Kidney thickness;LN;11855-4 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Lt Fetal Kidney | Fetal Left Kidney Length | Length | cm mm | 5 | 2 | Left Kidney length;LN;11834-9 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Rt Fetal Kidney | Fetal Right Kidney Length | Length | cm mm | 5 | 2 | Right Kidney length;LN;11836-4 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Lt Fetal Kidney | Fetal Left Kidney Width | Width | cm mm | 5 | 2 | Left Kidney width;LN;11825-7 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Rt Fetal Kidney | Fetal Right Kidney Width | Width | cm mm | 5 | 2 | Right Kidney width;LN;11827-3 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Fetal Movements | Fetal Biophysical Profile: Gross Body Movement |  | unitless | 1 | 0 | Gross Body Movement;LN;11631-9 |  | Biophysical Profile;DCM;125006 |  | Biophysical Profile;DCM;125006
Fetal Tone | Fetal Biophysical Profile: Fetal Tone |  | unitless | 1 | 0 | Fetal Tone;LN;11635-0 |  | Biophysical Profile;DCM;125006 |  | Biophysical Profile;DCM;125006
FTA | Fetal Trunk Area |  | cm² mm² | 5 | 2 | Fetal Trunk Area;99SIEMENS;FetalTrunkArea | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Fibroid 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid
Fibroid 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid
Fibroid 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid
Fibroid 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid
Fibroid 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid
Fibroid 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid
Fibroid 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid
Fibroid 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid
Fibroid 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid
Fibroid 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid
Fibroid 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid
Fibroid 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid
Fibroid 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid
Fibroid 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid
Fibroid 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid
Fibula | Fibula Length |  | cm mm | 5 | 2 | Fibula length;LN;11964-4 | NUM | Biometry Group;DCM;125005 |  | Fetal Long Bones;DCM;125003
Lt Follicle 1 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 1 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 1 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 1 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 1 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 1 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 1 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 1 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 1 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 1 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 10 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 10 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 10 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 10 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 10 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 10 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 10 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 10 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 10 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 10 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 11 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 11 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 11 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 11 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 11 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 11 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 11 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 11 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 11 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 11 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 12 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 12 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 12 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 12 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 12 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 12 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 12 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 12 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 12 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 12 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 13 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 13 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 13 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 13 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 13 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 13 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 13 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 13 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 13 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 13 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 14 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 14 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 14 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 14 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 14 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 14 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 14 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 14 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 14 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 14 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 15 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 15 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 15 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 15 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 15 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 15 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 15 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 15 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 15 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 15 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 16 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 16 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 16 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 16 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 16 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 16 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 16 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 16 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 16 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 16 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 17 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 17 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 17 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 17 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 17 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 17 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 17 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 17 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 17 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 17 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 18 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 18 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 18 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 18 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 18 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 18 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 18 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 18 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 18 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 18 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 19 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 19 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 19 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 19 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 19 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 19 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 19 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 19 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 19 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 19 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 2 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 2 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 2 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 2 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 2 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 2 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 2 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 2 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 2 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 2 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 20 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 20 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 20 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 20 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 20 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 20 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 20 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 20 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 20 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 20 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 3 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 3 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 3 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 3 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 3 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 3 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 3 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 3 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 3 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 3 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 4 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 4 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 4 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 4 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 4 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 4 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 4 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 4 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 4 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 4 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 5 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 5 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 5 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 5 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 5 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 5 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 5 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 5 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 5 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 5 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 6 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 6 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 6 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 6 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 6 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 6 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 6 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 6 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 6 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 6 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 7 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 7 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 7 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 7 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 7 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 7 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 7 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 7 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 7 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 7 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 8 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 8 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 8 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 8 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 8 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 8 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 8 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 8 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 8 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 8 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Follicle 9 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 9 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 9 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 9 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Follicle 9 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Follicle 9 Diam | Diameter |  | cm mm | 5 | 2 | Follicle diameter;LN;11793-7 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 9 | Volume |  | cm³ mm³ | 5 | 2 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 9 | Volume D1 | D1 | cm mm | 5 | 2 | Distance 1 of Volume;99SIEMENS;VolumeD1 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 9 | Volume D2 | D2 | cm mm | 5 | 2 | Distance 2 of Volume;99SIEMENS;VolumeD2 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Follicle 9 | Volume D3 | D3 | cm mm | 5 | 2 | Distance 3 of Volume;99SIEMENS;VolumeD3 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Ovarian Follicle;SRT;T-87600 | Laterality;SRT;G-C171=Right;SRT;G-A100
Foot | Foot Length |  | cm mm | 5 | 2 | Foot length;LN;11965-1 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
GSD | Gestational Sac Diameter |  | cm mm | 5 | 2 | Gestational Sac Diameter;LN;11850-5 | NUM | Biometry Group;DCM;125005 |  | Early Gestation;DCM;125009
HC | Head Circumference |  | cm mm | 5 | 2 | Head Circumference;LN;11984-2 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
HW | Hemispheric Width |  | cm mm | 5 | 2 | Width of Hemisphere;LN;12170-7 | NUM | Biometry Group;DCM;125005 |  | Fetal Cranium;DCM;125004
HL | Humerus Length |  | cm mm | 5 | 2 | Humerus length;LN;11966-9 | NUM | Biometry Group;DCM;125005 |  | Fetal Long Bones;DCM;125003
Lt Int Mammary A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Int Mammary A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Int Mammary A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Int Mammary A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Int Mammary A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Int Mammary A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Int Mammary A | Doppler Angle Property of PSV, EDV | Angle | ° | 5 | 0 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Lt Int Mammary A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Int Mammary A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Int Mammary A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Int Mammary A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Int Mammary A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Int Mammary A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Int Mammary A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Int Mammary A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Int Mammary A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Int Mammary A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Int Mammary A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Int Mammary A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Int Mammary A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Int Mammary A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Int Mammary A | Doppler Angle Property of PSV, EDV | Angle | ° | 5 | 0 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Rt Int Mammary A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Int Mammary A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Int Mammary A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Int Mammary A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Int Mammary A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Int Mammary A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Int Mammary A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Int Mammary A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
IOD | Interocular Distance |  | cm mm | 5 | 2 | Inner Orbital Diameter;LN;33070-4 | NUM | Biometry Group;DCM;125005 |  | Fetal Cranium;DCM;125004
Lt Kidney | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Kidney;SRT;T-71000 |  | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Kidney | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Kidney;SRT;T-71000 |  | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Kidney | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Kidney;SRT;T-71000 |  | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Kidney | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Kidney;SRT;T-71000 |  | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Kidney | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Kidney;SRT;T-71000 |  | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Kidney | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Kidney;SRT;T-71000 |  | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lat Vent | Lateral Ventricle width |  | cm mm | 5 | 2 | Lateral Ventricle width;LN;12171-5 | NUM | Biometry Group;DCM;125005 |  | Fetal Cranium;DCM;125004
Levator Hiatus AP Distance | Levator Hiatus AP Distance |  | cm | 5 | 2 | Levator Hiatus AP Distance;99SIEMENS;LevatrHiatusAP | NUM | Pelvic Floor;SRT;T-D6236 |  | Finding Site;SRT;G-C0E3=Pelvic Floor;SRT;T-D6236
Levator Hiatal Area | Levator Hiatal Area |  | cm² mm² | 5 | 2 | Levator Hiatal Area;99SIEMENS;LevatrHiatalArea | NUM | Pelvic Floor;SRT;T-D6236 |  | Finding Site;SRT;G-C0E3=Pelvic Floor;SRT;T-D6236
Levator Hiatus Lateral Distance | Levator Hiatus Lateral Distance |  | cm | 5 | 2 | Levator Hiatus Lateral Distance;99SIEMENS;LevatrHiatusLatD | NUM | Pelvic Floor;SRT;T-D6236 |  | Finding Site;SRT;G-C0E3=Pelvic Floor;SRT;T-D6236
Lt Breast Mass 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Rt Breast Mass 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Lt Mass 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Mass 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Mass 1 | Elasticity Imaging/B Mode Ratio | EI/B Area Ratio | unitless | 5 | 2 | EI/B Area Ratio;99SIEMENS;EBARatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Lt Mass 1 | E-mode | A (E) | cm² mm² | 5 | 2 | E-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioEArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Lt Mass 1 | B-mode | A (B) | cm² mm² | 5 | 2 | B-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioBArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Rt Mass 1 | Elasticity Imaging/B Mode Ratio | EI/B Area Ratio | unitless | 5 | 2 | EI/B Area Ratio;99SIEMENS;EBARatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Rt Mass 1 | E-mode | A (E) | cm² mm² | 5 | 2 | E-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioEArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Rt Mass 1 | B-mode | A (B) | cm² mm² | 5 | 2 | B-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioBArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Lt Mass 1 | Circumference | C | cm mm | 5 | 2 | Circumference;SRT;M-02560 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Lt Mass 1 | A | A | cm² mm² | 5 | 2 | Area;SRT;G-A166 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Lt Mass 1 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Lt Mass 1 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Lt Mass 1 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Rt Mass 1 | Circumference | C | cm mm | 5 | 2 | Circumference;SRT;M-02560 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Rt Mass 1 | A | A | cm² mm² | 5 | 2 | Area;SRT;G-A166 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Rt Mass 1 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Rt Mass 1 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Rt Mass 1 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Lt Breast Mass 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Rt Breast Mass 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Lt Mass 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Mass 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Mass 1 | Elasticity Imaging/B Mode Ratio | EI/B Distance Ratio | unitless | 5 | 2 | EI/B Distance Ratio;99SIEMENS;EBDRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Lt Mass 1 | E-mode | D (E) | cm mm | 5 | 2 | E-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioEDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Lt Mass 1 | B-mode | D (B) | cm mm | 5 | 2 | B-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioBDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Rt Mass 1 | Elasticity Imaging/B Mode Ratio | EI/B Distance Ratio | unitless | 5 | 2 | EI/B Distance Ratio;99SIEMENS;EBDRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Rt Mass 1 | E-mode | D (E) | cm mm | 5 | 2 | E-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioEDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Rt Mass 1 | B-mode | D (B) | cm mm | 5 | 2 | B-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioBDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Lt Breast Mass 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1
Rt Breast Mass 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1
Lt Mass 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Mass 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Breast Mass 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Rt Breast Mass 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Lt Mass 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Mass 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Mass 2 | Elasticity Imaging/B Mode Ratio | EI/B Area Ratio | unitless | 5 | 2 | EI/B Area Ratio;99SIEMENS;EBARatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Lt Mass 2 | E-mode | A (E) | cm² mm² | 5 | 2 | E-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioEArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Lt Mass 2 | B-mode | A (B) | cm² mm² | 5 | 2 | B-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioBArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Rt Mass 2 | Elasticity Imaging/B Mode Ratio | EI/B Area Ratio | unitless | 5 | 2 | EI/B Area Ratio;99SIEMENS;EBARatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Rt Mass 2 | E-mode | A (E) | cm² mm² | 5 | 2 | E-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioEArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Rt Mass 2 | B-mode | A (B) | cm² mm² | 5 | 2 | B-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioBArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Lt Mass 2 | Circumference | C | cm mm | 5 | 2 | Circumference;SRT;M-02560 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Lt Mass 2 | A | A | cm² mm² | 5 | 2 | Area;SRT;G-A166 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Lt Mass 2 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Lt Mass 2 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Lt Mass 2 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Rt Mass 2 | Circumference | C | cm mm | 5 | 2 | Circumference;SRT;M-02560 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Rt Mass 2 | A | A | cm² mm² | 5 | 2 | Area;SRT;G-A166 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Rt Mass 2 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Rt Mass 2 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Rt Mass 2 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Lt Breast Mass 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Rt Breast Mass 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Lt Mass 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Mass 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Mass 2 | Elasticity Imaging/B Mode Ratio | EI/B Distance Ratio | unitless | 5 | 2 | EI/B Distance Ratio;99SIEMENS;EBDRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Lt Mass 2 | E-mode | D (E) | cm mm | 5 | 2 | E-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioEDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Lt Mass 2 | B-mode | D (B) | cm mm | 5 | 2 | B-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioBDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Rt Mass 2 | Elasticity Imaging/B Mode Ratio | EI/B Distance Ratio | unitless | 5 | 2 | EI/B Distance Ratio;99SIEMENS;EBDRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Rt Mass 2 | E-mode | D (E) | cm mm | 5 | 2 | E-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioEDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Rt Mass 2 | B-mode | D (B) | cm mm | 5 | 2 | B-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioBDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Lt Breast Mass 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2
Rt Breast Mass 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2
Lt Mass 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Mass 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Breast Mass 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Rt Breast Mass 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Lt Mass 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Mass 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Mass 3 | Elasticity Imaging/B Mode Ratio | EI/B Area Ratio | unitless | 5 | 2 | EI/B Area Ratio;99SIEMENS;EBARatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Lt Mass 3 | E-mode | A (E) | cm² mm² | 5 | 2 | E-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioEArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Lt Mass 3 | B-mode | A (B) | cm² mm² | 5 | 2 | B-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioBArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Rt Mass 3 | Elasticity Imaging/B Mode Ratio | EI/B Area Ratio | unitless | 5 | 2 | EI/B Area Ratio;99SIEMENS;EBARatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Rt Mass 3 | E-mode | A (E) | cm² mm² | 5 | 2 | E-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioEArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Rt Mass 3 | B-mode | A (B) | cm² mm² | 5 | 2 | B-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioBArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Lt Mass 3 | Circumference | C | cm mm | 5 | 2 | Circumference;SRT;M-02560 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Lt Mass 3 | A | A | cm² mm² | 5 | 2 | Area;SRT;G-A166 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Lt Mass 3 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Lt Mass 3 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Lt Mass 3 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Rt Mass 3 | Circumference | C | cm mm | 5 | 2 | Circumference;SRT;M-02560 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Rt Mass 3 | A | A | cm² mm² | 5 | 2 | Area;SRT;G-A166 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Rt Mass 3 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Rt Mass 3 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Rt Mass 3 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Lt Breast Mass 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Rt Breast Mass 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Lt Mass 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Mass 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Mass 3 | Elasticity Imaging/B Mode Ratio | EI/B Distance Ratio | unitless | 5 | 2 | EI/B Distance Ratio;99SIEMENS;EBDRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Lt Mass 3 | E-mode | D (E) | cm mm | 5 | 2 | E-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioEDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Lt Mass 3 | B-mode | D (B) | cm mm | 5 | 2 | B-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioBDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Rt Mass 3 | Elasticity Imaging/B Mode Ratio | EI/B Distance Ratio | unitless | 5 | 2 | EI/B Distance Ratio;99SIEMENS;EBDRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Rt Mass 3 | E-mode | D (E) | cm mm | 5 | 2 | E-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioEDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Rt Mass 3 | B-mode | D (B) | cm mm | 5 | 2 | B-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioBDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Lt Breast Mass 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3
Rt Breast Mass 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3
Lt Mass 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Mass 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Breast Mass 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Rt Breast Mass 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Lt Mass 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Mass 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Mass 4 | Elasticity Imaging/B Mode Ratio | EI/B Area Ratio | unitless | 5 | 2 | EI/B Area Ratio;99SIEMENS;EBARatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Lt Mass 4 | E-mode | A (E) | cm² mm² | 5 | 2 | E-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioEArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Lt Mass 4 | B-mode | A (B) | cm² mm² | 5 | 2 | B-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioBArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Rt Mass 4 | Elasticity Imaging/B Mode Ratio | EI/B Area Ratio | unitless | 5 | 2 | EI/B Area Ratio;99SIEMENS;EBARatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Rt Mass 4 | E-mode | A (E) | cm² mm² | 5 | 2 | E-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioEArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Rt Mass 4 | B-mode | A (B) | cm² mm² | 5 | 2 | B-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioBArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Lt Mass 4 | Circumference | C | cm mm | 5 | 2 | Circumference;SRT;M-02560 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Lt Mass 4 | A | A | cm² mm² | 5 | 2 | Area;SRT;G-A166 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Rt Mass 4 | Circumference | C | cm mm | 5 | 2 | Circumference;SRT;M-02560 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Rt Mass 4 | A | A | cm² mm² | 5 | 2 | Area;SRT;G-A166 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Lt Breast Mass 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Rt Breast Mass 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Lt Mass 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Mass 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Mass 4 | Elasticity Imaging/B Mode Ratio | EI/B Distance Ratio | unitless | 5 | 2 | EI/B Distance Ratio;99SIEMENS;EBDRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Lt Mass 4 | E-mode | D (E) | cm mm | 5 | 2 | E-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioEDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Lt Mass 4 | B-mode | D (B) | cm mm | 5 | 2 | B-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioBDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Rt Mass 4 | Elasticity Imaging/B Mode Ratio | EI/B Distance Ratio | unitless | 5 | 2 | EI/B Distance Ratio;99SIEMENS;EBDRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Rt Mass 4 | E-mode | D (E) | cm mm | 5 | 2 | E-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioEDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Rt Mass 4 | B-mode | D (B) | cm mm | 5 | 2 | B-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioBDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Lt Mass 4 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Lt Mass 4 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Lt Mass 4 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Rt Mass 4 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Rt Mass 4 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Rt Mass 4 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Lt Breast Mass 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4
Rt Breast Mass 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4
Lt Mass 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Mass 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Breast Mass 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Rt Breast Mass 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Lt Mass 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Mass 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Mass 5 | Elasticity Imaging/B Mode Ratio | EI/B Area Ratio | unitless | 5 | 2 | EI/B Area Ratio;99SIEMENS;EBARatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Lt Mass 5 | E-mode | A (E) | cm² mm² | 5 | 2 | E-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioEArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Lt Mass 5 | B-mode | A (B) | cm² mm² | 5 | 2 | B-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioBArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Rt Mass 5 | Elasticity Imaging/B Mode Ratio | EI/B Area Ratio | unitless | 5 | 2 | EI/B Area Ratio;99SIEMENS;EBARatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Rt Mass 5 | E-mode | A (E) | cm² mm² | 5 | 2 | E-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioEArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Rt Mass 5 | B-mode | A (B) | cm² mm² | 5 | 2 | B-mode Area of EI/B Area Ratio;99SIEMENS;EBARatioBArea | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Lt Mass 5 | Circumference | C | cm mm | 5 | 2 | Circumference;SRT;M-02560 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Lt Mass 5 | A | A | cm² mm² | 5 | 2 | Area;SRT;G-A166 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Rt Mass 5 | Circumference | C | cm mm | 5 | 2 | Circumference;SRT;M-02560 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Rt Mass 5 | A | A | cm² mm² | 5 | 2 | Area;SRT;G-A166 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Lt Breast Mass 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Rt Breast Mass 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Lt Mass 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Mass 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Mass 5 | Elasticity Imaging/B Mode Ratio | EI/B Distance Ratio | unitless | 5 | 2 | EI/B Distance Ratio;99SIEMENS;EBDRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Lt Mass 5 | E-mode | D (E) | cm mm | 5 | 2 | E-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioEDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Lt Mass 5 | B-mode | D (B) | cm mm | 5 | 2 | B-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioBDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Rt Mass 5 | Elasticity Imaging/B Mode Ratio | EI/B Distance Ratio | unitless | 5 | 2 | EI/B Distance Ratio;99SIEMENS;EBDRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Rt Mass 5 | E-mode | D (E) | cm mm | 5 | 2 | E-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioEDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Rt Mass 5 | B-mode | D (B) | cm mm | 5 | 2 | B-mode Distance of EI/B Distance Ratio;99SIEMENS;EBDRatioBDist | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Lt Mass 5 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Lt Mass 5 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Lt Mass 5 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Rt Mass 5 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Rt Mass 5 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Rt Mass 5 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Lt Breast Mass 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5
Rt Breast Mass 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5
Lt Mass 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Mass 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt MCA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Lt MCA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Lt MCA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Lt MCA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Lt MCA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Lt MCA | Doppler Angle Property of PSV, EDV, TAMx | Angle | ° | 5 | 0 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Lt MCA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Lt MCA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Rt MCA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Rt MCA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Rt MCA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Rt MCA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Rt MCA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Rt MCA | Doppler Angle Property of PSV, EDV, TAMx | Angle | ° | 5 | 0 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Rt MCA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Rt MCA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Nasal Bone | Nasal Bone Length |  | cm mm | 5 | 2 | Fetal Nasal bone length;LN;53697-9 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Nonstress Test | Fetal Biophysical Profile: Fetal Heart Reactivity |  | unitless | 1 | 0 | Fetal Heart Reactivity;LN;11635-5 |  | Biophysical Profile;DCM;125006 |  | Biophysical Profile;DCM;125006
Nuchal Thickness | Nuchal Fold Thickness |  | cm mm | 5 | 2 | Nuchal Fold thickness;LN;12146-7 | NUM | Biometry Group;DCM;125005 |  | Fetal Cranium;DCM;125004
Nuchal Translucency | Nuchal Translucency |  | cm mm | 5 | 2 | Nuchal Translucency;LN;33069-6 | NUM | Biometry Group;DCM;125005 |  | Early Gestation;DCM;125009
OFD | Occipital-Frontal Diameter |  | cm mm | 5 | 2 | Occipital-Frontal Diameter;LN;11851-3 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Lt Ovarian A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian A | Doppler Angle Property of PSV, EDV, TAMx | Angle | ° | 5 | 0 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Lt Ovarian A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian A | Doppler Angle Property of PSV, EDV, TAMx | Angle | ° | 5 | 0 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Rt Ovarian A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ovarian Vein;SRT;T-48780 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ovarian Vein;SRT;T-48780 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Ovarian Vein;SRT;T-48780 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Rt Ovarian V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Ovarian Vein;SRT;T-48780 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Lt Ovarian V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Ovarian Vein;SRT;T-48780 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Ovarian V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Ovarian V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Ovarian Vein;SRT;T-48780 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Ovary (TA) | Anterior-Posterior Diameter (Transabdominal) | AP | cm mm | 5 | 2 | Left Ovary Height;LN;11857-0 | NUM | Ovary;SRT;T-87000 |  | Finding Site;SRT;G-C0E3=Ovary;SRT;T-87000
Rt Ovary (TA) | Anterior-Posterior Diameter (Transabdominal) | AP | cm mm | 5 | 2 | Right Ovary Height;LN;11858-8 | NUM | Ovary;SRT;T-87000 |  | Finding Site;SRT;G-C0E3=Ovary;SRT;T-87000
Lt Ovary (TA) | Length (Transabdominal) | Length | cm mm | 5 | 2 | Left Ovary Length;LN;11840-6 | NUM | Ovary;SRT;T-87000 |  | Finding Site;SRT;G-C0E3=Ovary;SRT;T-87000
Rt Ovary (TA) | Length (Transabdominal) | Length | cm mm | 5 | 2 | Right Ovary Length;LN;11841-4 | NUM | Ovary;SRT;T-87000 |  | Finding Site;SRT;G-C0E3=Ovary;SRT;T-87000
Lt Ovary (EV) | Anterior-Posterior Diameter (Endovaginal) | AP | cm mm | 5 | 2 | Left Ovary Height;LN;11857-0 | NUM | Ovary;SRT;T-87000 |  | Finding Site;SRT;G-C0E3=Ovary;SRT;T-87000 |  | Imaging procedure;SRT;P0-0099A=Transvaginal echography;SRT;P5-B8012
Rt Ovary (EV) | Anterior-Posterior Diameter (Endovaginal) | AP | cm mm | 5 | 2 | Right Ovary Height;LN;11858-8 | NUM | Ovary;SRT;T-87000 |  | Finding Site;SRT;G-C0E3=Ovary;SRT;T-87000 |  | Imaging procedure;SRT;P0-0099A=Transvaginal echography;SRT;P5-B8012
Lt Ovary (EV) | Length (Endovaginal) | Length | cm mm | 5 | 2 | Left Ovary Length;LN;11840-6 | NUM | Ovary;SRT;T-87000 |  | Finding Site;SRT;G-C0E3=Ovary;SRT;T-87000 |  | Imaging procedure;SRT;P0-0099A=Transvaginal echography;SRT;P5-B8012
Rt Ovary (EV) | Length (Endovaginal) | Length | cm mm | 5 | 2 | Right Ovary Length;LN;11841-4 | NUM | Ovary;SRT;T-87000 |  | Finding Site;SRT;G-C0E3=Ovary;SRT;T-87000 |  | Imaging procedure;SRT;P0-0099A=Transvaginal echography;SRT;P5-B8012
Lt Ovary (EV) | Width (Endovaginal) | Width | cm mm | 5 | 2 | Left Ovary Width;LN;11829-9 | NUM | Ovary;SRT;T-87000 |  | Finding Site;SRT;G-C0E3=Ovary;SRT;T-87000 |  | Imaging procedure;SRT;P0-0099A=Transvaginal echography;SRT;P5-B8012
Rt Ovary (EV) | Width (Endovaginal) | Width | cm mm | 5 | 2 | Right Ovary Width;LN;11830-7 | NUM | Ovary;SRT;T-87000 |  | Finding Site;SRT;G-C0E3=Ovary;SRT;T-87000 |  | Imaging procedure;SRT;P0-0099A=Transvaginal echography;SRT;P5-B8012
Lt Ovary (TA) | Width (Transabdominal) | Width | cm mm | 5 | 2 | Left Ovary Width;LN;11829-9 | NUM | Ovary;SRT;T-87000 |  | Finding Site;SRT;G-C0E3=Ovary;SRT;T-87000
Rt Ovary (TA) | Width (Transabdominal) | Width | cm mm | 5 | 2 | Right Ovary Width;LN;11830-7 | NUM | Ovary;SRT;T-87000 |  | Finding Site;SRT;G-C0E3=Ovary;SRT;T-87000
LV-P | Posterior Horn Lateral ventricular width |  | cm mm | 5 | 2 | Posterior Horn Lateral ventricular width;LN;33196-7 | NUM | Biometry Group;DCM;125005 |  | Fetal Cranium;DCM;125004
Q1 | Q1 Diameter |  | cm mm | 1 | 2 | First Quadrant Diameter;LN;11624-4 | NUM | Amniotic Sac;SRT;T-F1300 |  | Finding Site;SRT;G-C0E3=Amniotic Sac;SRT;T-F1300
Q1 | Q1 Diameter |  | cm mm | 1 | 2 | First Quadrant Diameter;LN;11624-4 | NUM | Amniotic Sac;SRT;T-F1300 |  | Finding Site;SRT;G-C0E3=Amniotic Sac;SRT;T-F1300 |  | Identifier;DCM;125010=AFI 2
Q1 | Q1 Diameter |  | cm mm | 1 | 2 | First Quadrant Diameter;LN;11624-4 | NUM | Amniotic Sac;SRT;T-F1300 |  | Finding Site;SRT;G-C0E3=Amniotic Sac;SRT;T-F1300 |  | Identifier;DCM;125010=AFI 3
Q2 | Q2 Diameter |  | cm mm | 1 | 2 | Second Quadrant Diameter;LN;11626-9 | NUM | Amniotic Sac;SRT;T-F1300 |  | Finding Site;SRT;G-C0E3=Amniotic Sac;SRT;T-F1300
Q2 | Q2 Diameter |  | cm mm | 1 | 2 | Second Quadrant Diameter;LN;11626-9 | NUM | Amniotic Sac;SRT;T-F1300 |  | Finding Site;SRT;G-C0E3=Amniotic Sac;SRT;T-F1300 |  | Identifier;DCM;125010=AFI 2
Q2 | Q2 Diameter |  | cm mm | 1 | 2 | Second Quadrant Diameter;LN;11626-9 | NUM | Amniotic Sac;SRT;T-F1300 |  | Finding Site;SRT;G-C0E3=Amniotic Sac;SRT;T-F1300 |  | Identifier;DCM;125010=AFI 3
Q3 | Q3 Diameter |  | cm mm | 1 | 2 | Third Quadrant Diameter;LN;11625-1 | NUM | Amniotic Sac;SRT;T-F1300 |  | Finding Site;SRT;G-C0E3=Amniotic Sac;SRT;T-F1300
Q3 | Q3 Diameter |  | cm mm | 1 | 2 | Third Quadrant Diameter;LN;11625-1 | NUM | Amniotic Sac;SRT;T-F1300 |  | Finding Site;SRT;G-C0E3=Amniotic Sac;SRT;T-F1300 |  | Identifier;DCM;125010=AFI 2
Q3 | Q3 Diameter |  | cm mm | 1 | 2 | Third Quadrant Diameter;LN;11625-1 | NUM | Amniotic Sac;SRT;T-F1300 |  | Finding Site;SRT;G-C0E3=Amniotic Sac;SRT;T-F1300 |  | Identifier;DCM;125010=AFI 3
Q4 | Q4 Diameter |  | cm mm | 1 | 2 | Fourth Quadrant Diameter;LN;11623-6 | NUM | Amniotic Sac;SRT;T-F1300 |  | Finding Site;SRT;G-C0E3=Amniotic Sac;SRT;T-F1300
Q4 | Q4 Diameter |  | cm mm | 1 | 2 | Fourth Quadrant Diameter;LN;11623-6 | NUM | Amniotic Sac;SRT;T-F1300 |  | Finding Site;SRT;G-C0E3=Amniotic Sac;SRT;T-F1300 |  | Identifier;DCM;125010=AFI 2
Q4 | Q4 Diameter |  | cm mm | 1 | 2 | Fourth Quadrant Diameter;LN;11623-6 | NUM | Amniotic Sac;SRT;T-F1300 |  | Finding Site;SRT;G-C0E3=Amniotic Sac;SRT;T-F1300 |  | Identifier;DCM;125010=AFI 3
Radius | Radius Length |  | cm mm | 5 | 2 | Radius length;LN;11967-7 | NUM | Biometry Group;DCM;125005 |  | Fetal Long Bones;DCM;125003
Rectal Ampulla descent | Rectal Ampulla descent distance |  | cm | 5 | 2 | Rectal Ampulla descent;99SIEMENS;RectalAmpDescent | NUM | Pelvic Floor;SRT;T-D6236 |  | Finding Site;SRT;G-C0E3=Pelvic Floor;SRT;T-D6236
Rectovesical Angle at Rest | Rectovesical Angle at Rest |  | ° | 5 | 0 | Rectovesical Angle at Rest;99SIEMENS;RectovAngleRest | NUM | Pelvic Floor;SRT;T-D6236 |  | Finding Site;SRT;G-C0E3=Pelvic Floor;SRT;T-D6236
Rectovesical Angle at Valsalvas | Rectovesical Angle at Valsalvas |  | ° | 5 | 0 | Rectovesical Angle at Valsalvas;99SIEMENS;RectovAngleValsa | NUM | Pelvic Floor;SRT;T-D6236 |  | Finding Site;SRT;G-C0E3=Pelvic Floor;SRT;T-D6236
TA | Thoracic Area |  | cm² mm² | 5 | 2 | Thoracic Area;LN;33068-8 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
TC | Thoracic Circumference |  | cm mm | 5 | 2 | Thoracic Circumference;LN;11988-3 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Tibia | Tibia Length |  | cm mm | 5 | 2 | Tibia length;LN;11968-5 | NUM | Biometry Group;DCM;125005 |  | Fetal Long Bones;DCM;125003
TAD | Transverse Abdominal Diameter |  | cm mm | 5 | 2 | Transverse Abdominal Diameter;LN;11862-0 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Cerebellum | Cerebellum |  | cm mm | 5 | 2 | Trans Cerebellar Diameter;LN;11863-8 | NUM | Biometry Group;DCM;125005 |  | Fetal Cranium;DCM;125004
TTD | Transverse Thoracic Diameter |  | cm mm | 5 | 2 | Transverse Thoracic Diameter;LN;11864-6 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Ulna | Ulna Length |  | cm mm | 5 | 2 | Ulna length;LN;11969-3 | NUM | Biometry Group;DCM;125005 |  | Fetal Long Bones;DCM;125003
Umb A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Umbilical Artery;SRT;T-F1810 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Umb A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Umbilical Artery;SRT;T-F1810 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Umb A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Umbilical Artery;SRT;T-F1810 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Umb A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Umbilical Artery;SRT;T-F1810 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Umb A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Umbilical Artery;SRT;T-F1810 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Umb A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Umbilical Artery;SRT;T-F1810 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Umb A | Doppler Angle Property of PSV, EDV, TAMx | Angle | ° | 5 | 0 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Umb A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Umbilical Artery;SRT;T-F1810 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Umb A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Umbilical Artery;SRT;T-F1810 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Lt Ureter | Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Ureter;SRT;T-73000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
Rt Ureter | Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Ureter;SRT;T-73000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
Lt Uterine A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Uterine A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Uterine A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Uterine A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Uterine A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Uterine A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Uterine A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Uterine A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Uterine A | Doppler Angle Property of PSV, EDV, TAMx | Angle | ° | 5 | 0 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Lt Uterine A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Uterine A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Uterine A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Uterine A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Uterine A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Uterine A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | Doppler Angle Property of PSV, EDV, TAMx | Angle | ° | 5 | 0 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Rt Uterine A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Uterus (TA) | Anterior-Posterior Diameter (Transabdominal) | AP | cm mm | 5 | 2 | Uterus Height;LN;11859-6 | NUM | Uterus;SRT;T-83000 |  | Finding Site;SRT;G-C0E3=Pelvis and Uterus;DCM;125011
Uterus (TA) | Length (Transabdominal) | Length | cm mm | 5 | 2 | Uterus Length;LN;11842-2 | NUM | Uterus;SRT;T-83000 |  | Finding Site;SRT;G-C0E3=Pelvis and Uterus;DCM;125011
Uterus (EV) | Anterior-Posterior Diameter (Endovaginal) | AP | cm mm | 5 | 2 | Uterus Height;LN;11859-6 | NUM | Uterus;SRT;T-83000 |  | Finding Site;SRT;G-C0E3=Pelvis and Uterus;DCM;125011 |  | Imaging procedure;SRT;P0-0099A=Transvaginal echography;SRT;P5-B8012
Uterus (EV) | Length (Endovaginal) | Length | cm mm | 5 | 2 | Uterus Length;LN;11842-2 | NUM | Uterus;SRT;T-83000 |  | Finding Site;SRT;G-C0E3=Pelvis and Uterus;DCM;125011 |  | Imaging procedure;SRT;P0-0099A=Transvaginal echography;SRT;P5-B8012
Uterus (EV) | Width (Endovaginal) | Width | cm mm | 5 | 2 | Uterus Width;LN;11865-3 | NUM | Uterus;SRT;T-83000 |  | Finding Site;SRT;G-C0E3=Pelvis and Uterus;DCM;125011 |  | Imaging procedure;SRT;P0-0099A=Transvaginal echography;SRT;P5-B8012
Uterus (TA) | Width (Transabdominal) | Width | cm mm | 5 | 2 | Uterus Width;LN;11865-3 | NUM | Uterus;SRT;T-83000 |  | Finding Site;SRT;G-C0E3=Pelvis and Uterus;DCM;125011
Yolk Sac | Yolk Sac Length |  | cm mm | 5 | 2 | Yolk Sac length;LN;11816-6 | NUM | Biometry Group;DCM;125005 |  | Early Gestation;DCM;125009
Aborta | Aborta |  | unitless | 1 | 0 | Aborta;LN;11612-9 | NUM | Patient Characteristics;DCM;121118 |  | Patient Characteristics;DCM;121118
Ectopics | Ectopics |  | unitless | 1 | 0 | Ectopic Pregnancies;LN;33065-4 | NUM | Patient Characteristics;DCM;121118 |  | Patient Characteristics;DCM;121118
Gravida | Gravida |  | unitless | 1 | 0 | Gravida;LN;11996-6 | NUM | Patient Characteristics;DCM;121118 |  | Patient Characteristics;DCM;121118
Para | Para |  | unitless | 1 | 0 | Para;LN;11977-6 | NUM | Patient Characteristics;DCM;121118 |  | Patient Characteristics;DCM;121118
Lt Breast Mass 1 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 1 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 1 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 1 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 1 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 1 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 2 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 2 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 2 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 2 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 2 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 2 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 3 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 3 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 3 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 3 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 3 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 3 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 4 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 4 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 4 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 4 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 4 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 4 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 5 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 5 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 5 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 5 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 5 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 5 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 1 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 1 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 1 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 1 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 1 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 1 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 2 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 2 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 2 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 2 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 2 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 2 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 3 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 3 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 3 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 3 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 3 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 3 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 4 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 4 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 4 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 4 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 4 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 4 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 5 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 5 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 5 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 5 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 5 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 5 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 1 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 1 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 1 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 1 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 1 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 1 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 2 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 2 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 2 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 2 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 2 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 2 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 3 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 3 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 3 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 3 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 3 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 3 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 4 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 4 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 4 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 4 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 4 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 4 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 5 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 5 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 5 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 5 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 5 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 5 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 1 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 1 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 1 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 1 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 1 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 1 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 2 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 2 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 2 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 2 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 2 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 2 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 3 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 3 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 3 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 3 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 3 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 3 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 4 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 4 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 4 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 4 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 4 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 4 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 5 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 5 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 5 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 5 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 5 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 5 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
PostVoid Bladder | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Post voiding;DCM;109135 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
PostVoid Bladder | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Post voiding;DCM;109135 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
PostVoid Bladder | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Post voiding;DCM;109135 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
PreVoid Bladder | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Prior to voiding;DCM;109134 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
PreVoid Bladder | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Prior to voiding;DCM;109134 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
PreVoid Bladder | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Prior to voiding;DCM;109134 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
LVOT Diam s | Left Ventricular Outflow Tract Diameter systole |  | cm mm | 5 | 2 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | Left Ventricle;SRT;T-32600 |  | Finding Site;SRT;G-C0E3=Fetal Heart;99SIEMENS;FetalHeart |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650¦Cardiac Cycle Point;SRT;R-4089A=End Systole;SRT;R-FAB5B
RVOT Diam s | Right Ventricular Outflow Tract Diameter systole |  | cm mm | 5 | 2 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | Right Ventricle;SRT;T-32500 |  | Finding Site;SRT;G-C0E3=Fetal Heart;99SIEMENS;FetalHeart |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550¦Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020
HA | Heart Area |  | cm² mm² | 5 | 2 | Cardiac Cross-sectional Area, transverse by US;LN;59075-2 | NUM | Fetal Heart;99SIEMENS;FetalHeart |  | Finding Site;SRT;G-C0E3=Fetal Heart;99SIEMENS;FetalHeart
Heart Circum | Heart Circumference |  | cm mm | 5 | 2 | Cardiac Circumference, transverse by US;LN;59073-7 | NUM | Fetal Heart;99SIEMENS;FetalHeart |  | Finding Site;SRT;G-C0E3=Fetal Heart;99SIEMENS;FetalHeart
MSD | Mean Gestational Sac Diameter |  | cm mm | 1 | 2 | Mean Sac Diameter;99SIEMENS;MeanSacDiam | NUM | Biometry Group;DCM;125005 |  | Early Gestation;DCM;125009
MSD | Distance 1 of Mean Gestational Sac Diameter | D1 | cm mm | 1 | 2 | Gestational Sac Diameter;LN;11850-5 | NUM | Biometry Group;DCM;125005 |  | Early Gestation;DCM;125009
MSD | Distance 2 of Mean Gestational Sac Diameter | D2 | cm mm | 1 | 2 | Gestational Sac Diameter;LN;11850-5 | NUM | Biometry Group;DCM;125005 |  | Early Gestation;DCM;125009
MSD | Distance 3 of Mean Gestational Sac Diameter | D3 | cm mm | 1 | 2 | Gestational Sac Diameter;LN;11850-5 | NUM | Biometry Group;DCM;125005 |  | Early Gestation;DCM;125009
Lt Int Mammary A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Int Mammary A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Ovarian A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Uterine A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Int Mammary A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Int Mammary A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Ovarian A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Ovarian A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Uterine A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Mass 1 | Distance 1 of Breast Area Mass 1 | D1 | cm mm | 5 | 2
Rt Mass 1 | Distance 2 of Breast Area Mass 1 | D2 | cm mm | 5 | 2
Rt Mass 2 | Distance 1 of Breast Area Mass 2 | D1 | cm mm | 5 | 2
Rt Mass 2 | Distance 2 of Breast Area Mass 2 | D2 | cm mm | 5 | 2
Rt Mass 3 | Distance 1 of Breast Area Mass 3 | D1 | cm mm | 5 | 2
Rt Mass 3 | Distance 2 of Breast Area Mass 3 | D2 | cm mm | 5 | 2
Rt Mass 4 | Distance 1 of Breast Area Mass 4 | D1 | cm mm | 5 | 2
Rt Mass 4 | Distance 2 of Breast Area Mass 4 | D2 | cm mm | 5 | 2
Rt Mass 5 | Distance 1 of Breast Area Mass 5 | D1 | cm mm | 5 | 2
Rt Mass 5 | Distance 2 of Breast Area Mass 5 | D2 | cm mm | 5 | 2
Lt Mass 1 | Distance 1 of Breast Area Mass 1 | D1 | cm mm | 5 | 2
Lt Mass 1 | Distance 2 of Breast Area Mass 1 | D2 | cm mm | 5 | 2
Lt Mass 2 | Distance 1 of Breast Area Mass 2 | D1 | cm mm | 5 | 2
Lt Mass 2 | Distance 2 of Breast Area Mass 2 | D2 | cm mm | 5 | 2
Lt Mass 3 | Distance 1 of Breast Area Mass 3 | D1 | cm mm | 5 | 2
Lt Mass 3 | Distance 2 of Breast Area Mass 3 | D2 | cm mm | 5 | 2
Lt Mass 4 | Distance 1 of Breast Area Mass 4 | D1 | cm mm | 5 | 2
Lt Mass 4 | Distance 2 of Breast Area Mass 4 | D2 | cm mm | 5 | 2
Lt Mass 5 | Distance 1 of Breast Area Mass 5 | D1 | cm mm | 5 | 2
Lt Mass 5 | Distance 2 of Breast Area Mass 5 | D2 | cm mm | 5 | 2
HC | Head Circumference |  | cm mm | 5 | 2
HC | Head Circumference |  | cm mm | 5 | 2
HC | Head Circumference |  | cm mm | 5 | 2
HC | Head Circumference | A | cm² mm² | 5 | 2
AC | Head Circumference |  | cm mm | 5 | 2
AC | Head Circumference |  | cm mm | 5 | 2
AC | Head Circumference |  | cm mm | 5 | 2
AC | Abdominal Circumference | A | cm² mm² | 5 | 2
Lt Breast Mass 1 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 1 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 2 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 2 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 3 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 3 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 4 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 4 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 5 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Mass 5 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 1 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 1 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 2 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 2 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 3 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 3 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 4 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 4 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 5 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Breast Site 5 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 1 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 1 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 2 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 2 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 3 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 3 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 4 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 4 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 5 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Mass 5 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 1 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 1 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 2 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 2 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 3 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 3 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 4 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 4 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 5 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Breast Site 5 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Aorta | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Lt Int Mammary A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt MCA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Lt Ovarian A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Lt Uterine A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Int Mammary A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Mammary Artery;SRT;T-46200 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt MCA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Rt Ovarian A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Ovarian Artery;SRT;T-46980 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Rt Uterine A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Pelvic Vascular Structure;SRT;T-D6007
Umb A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Umbilical Artery;SRT;T-F1810 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800

## OB Gyn Calculations
EnglishLabel | Description | DetailEnglishLabel | DefaultUnits | DefaultPrecision | MaxInstances | ConceptCoding | ValueType | ParentContainer | ParentLaterality | SRSection | SRSectionLaterality | AdditionalModifiers | AdditionalModifiers2 | AdditionalModifiers3
Rt Breast Mass 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Breast Mass 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Breast Mass 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Breast Mass 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Breast Mass 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Breast Mass 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Breast Mass 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Breast Mass 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Breast Mass 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Breast Mass 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Axillary Lymph Node 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Axillary Lymph Node 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Axillary Lymph Node 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Axillary Lymph Node 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Axillary Lymph Node 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Axillary Lymph Node 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Axillary Lymph Node 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Axillary Lymph Node 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Axillary Lymph Node 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Axillary Lymph Node 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Axillary Lymph Node;SRT;T-C4710 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Breast Cyst 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Breast Cyst 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Breast Cyst 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Breast Cyst 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Breast Cyst 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Breast Cyst 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Breast Cyst 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Breast Cyst 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Breast Cyst 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Breast Cyst 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Cyst;SRT;M-3340A |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Mass 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Mass 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Mass 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Mass 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Mass 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Mass 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Mass 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Mass 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Mass 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Mass 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Cyst 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Cyst 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Cyst 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Cyst 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Cyst 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Cyst 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Cyst 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Cyst 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Cyst 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Cyst 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Cyst;SRT;M-3340A | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Ovary (TA) | Volume | Vol | cm³ | 2 | 1 | Right Ovary Volume;LN;12165-7 | NUM | Ovary;SRT;T-87000 |  | Finding Site;SRT;G-C0E3=Ovary;SRT;T-87000
Lt Ovary (TA) | Volume | Vol | cm³ | 2 | 1 | Left Ovary Volume;LN;12164-0 | NUM | Ovary;SRT;T-87000 |  | Finding Site;SRT;G-C0E3=Ovary;SRT;T-87000
Rt Ovary (EV) | Volume | Vol | cm³ | 2 | 1 | Right Ovary Volume;LN;12165-7 | NUM | Ovary;SRT;T-87000 |  | Finding Site;SRT;G-C0E3=Ovary;SRT;T-87000 |  | Imaging procedure;SRT;P0-0099A=Transvaginal echography;SRT;P5-B8012
Lt Ovary (EV) | Volume | Vol | cm³ | 2 | 1 | Left Ovary Volume;LN;12164-0 | NUM | Ovary;SRT;T-87000 |  | Finding Site;SRT;G-C0E3=Ovary;SRT;T-87000 |  | Imaging procedure;SRT;P0-0099A=Transvaginal echography;SRT;P5-B8012
Uterus (TA) | Volume | Vol | cm³ | 2 | 1 | Uterus Volume;LN;33192-6 | NUM | Uterus;SRT;T-83000 |  | Finding Site;SRT;G-C0E3=Pelvis and Uterus;DCM;125011
Uterus (EV) | Volume | Vol | cm³ | 2 | 1 | Uterus Volume;LN;33192-6 | NUM | Uterus;SRT;T-83000 |  | Finding Site;SRT;G-C0E3=Pelvis and Uterus;DCM;125011 |  | Imaging procedure;SRT;P0-0099A=Transvaginal echography;SRT;P5-B8012
Rt Kidney | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Kidney;SRT;T-71000 |  | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Kidney | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Kidney;SRT;T-71000 |  | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Left;SRT;G-A101
BPDa | Corrected Biparietal Diameter |  | cm | 1 | 1 | BPD area corrected;LN;11824-0 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
APTDxTTD | Fetal Thoracic Cross-sectional Area |  | cm² | 1 | 1 | APTD x TTD;99SIEMENS;APTDxTTD | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
APADxTAD | Abdominal Cross-sectional Area |  | cm² | 1 | 1 | APAD * TAD;LN;33191-8 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
AFI | Amniotic Fluid Index |  | cm | 1 | 1 | Amniotic Fluid Index;LN;11627-7 | NUM | Amniotic Sac;SRT;T-F1300 |  | Finding Site;SRT;G-C0E3=Amniotic Sac;SRT;T-F1300
AFI 2 | Amniotic Fluid Index 2 |  | cm | 1 | 1 | Amniotic Fluid Index;LN;11627-7 | NUM | Amniotic Sac;SRT;T-F1300 |  | Finding Site;SRT;G-C0E3=Amniotic Sac;SRT;T-F1300
AFI 3 | Amniotic Fluid Index 3 |  | cm | 1 | 1 | Amniotic Fluid Index;LN;11627-7 | NUM | Amniotic Sac;SRT;T-F1300 |  | Finding Site;SRT;G-C0E3=Amniotic Sac;SRT;T-F1300
Rt Fetal Kidney | Volume | Vol | cm³ | 2 | 1 | Right Fetal Kidney volume;99SIEMENS;RtFetalKidneyVol | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
Lt Fetal Kidney | Volume | Vol | cm³ | 2 | 1 | Left Fetal Kidney volume;99SIEMENS;LtFetalKidneyVol | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
HC/AC | HC/AC |  | unitless | 2 | 1 | HC/AC;LN;11947-9 | NUM | Fetal Biometry Ratios;DCM;125001 |  | Fetal Biometry Ratios;DCM;125001
FL/BPD | FL/BPD |  | unitless | 2 | 1 | FL/BPD;LN;11872-9 | NUM | Fetal Biometry Ratios;DCM;125001 |  | Fetal Biometry Ratios;DCM;125001
FL/HC | FL/HC |  | unitless | 2 | 1 | FL/HC;LN;11873-7 | NUM | Fetal Biometry Ratios;DCM;125001 |  | Fetal Biometry Ratios;DCM;125001
FL/AC | FL/AC |  | unitless | 2 | 1 | FL/AC;LN;11871-1 | NUM | Fetal Biometry Ratios;DCM;125001 |  | Fetal Biometry Ratios;DCM;125001
TCD/AC | TCD/AC |  | unitless | 2 | 1 | TCD/AC;99SIEMENS;TCDOverAC | NUM | Fetal Biometry Ratios;DCM;125001 |  | Fetal Biometry Ratios;DCM;125001
LVW/HW | LVW/HW |  | unitless | 2 | 1 | LVW/HW;99SIEMENS;LVWOverHW | NUM | Fetal Biometry Ratios;DCM;125001 |  | Fetal Biometry Ratios;DCM;125001
MAD | Mean Abdominal Diameter |  | cm | 2 | 1 | Mean Abdominal Diameter;99SIEMENS;MeanAbdDiam | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002
CI | Cephalic Index |  | unitless | 2 | 1 | Cephalic Index;LN;11823-2 | NUM | Fetal Biometry Ratios;DCM;125001 |  | Fetal Biometry Ratios;DCM;125001
Clinical Age | Fetus age based on LMP |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM | Summary;DCM;121111 |  | Summary;DCM;121111
Total Score | Biophysical Profile Sum Score |  | unitless | 0 | 1 | Biophysical Profile Sum Score;LN;11634-3 | NUM | Biophysical Profile;DCM;125006 |  | Biophysical Profile;DCM;125006
Fibroid 1 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid |  | Identifier;DCM;125010=1
Fibroid 2 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid |  | Identifier;DCM;125010=2
Fibroid 3 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid |  | Identifier;DCM;125010=3
Fibroid 4 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid |  | Identifier;DCM;125010=4
Fibroid 5 | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Measurement Group;DCM;125007 |  | Finding Site;SRT;G-C0E3=Fibroid;99SIEMENS;Fibroid |  | Identifier;DCM;125010=5
LMP | Last Menstrual Period Date |  | MCR_DATE |  | 1 | LMP;LN;11955-2 | DATE | Summary;DCM;121111 |  | Summary;DCM;121111
EDD | Estimated Delivery Date |  | MCR_DATE |  | 1 | EDD;LN;11778-8 | DATE | Summary;DCM;121111 |  | Summary;DCM;121111
DOC | In-vitro Fertilization Date |  | MCR_DATE |  | 1 | Conception date;LN;33067-0 | DATE | Summary;DCM;121111 |  | Summary;DCM;121111
Num of fetuses | Number of fetuses |  | MCR_COUNT | 0 | 1 | Number of Fetuses;LN;11878-6 | NUM | Summary;DCM;121111 |  | Summary;DCM;121111
Rt CP Ratio | Rt MCA PI / Umb A PI |  | unitless | 2 | 1 | Cerebroplacental ratio;99SIEMENS;CPRatio | NUM | Fetal Vascular Ratios;99SIEMENS;FetalVascRatios | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Lt CP Ratio | Lt MCA PI / Umb A PI |  | unitless | 2 | 1 | Cerebroplacental ratio;99SIEMENS;CPRatio | NUM | Fetal Vascular Ratios;99SIEMENS;FetalVascRatios | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Hill | TCD (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Cranium;DCM;125004 |  | Equation;DCM;121420=TCD, Hill 1990;LN;33134-8 | 2 Sigma deviation of population ;DCM;121417
Hadlock | BPD (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=BPD, Hadlock 1984;LN;11902-4 | 2 Sigma deviation of population ;DCM;121417
Verburg (outer-outer) | BPD (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD (outer-outer), Verburg 2008;99SIEMENS;GA107-22-01 | 2 Sigma deviation of population ;DCM;121417
Jeanty | BPD (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD, Jeanty 1984;LN;11905-7 | 2 Sigma deviation of population ;DCM;121417
Lasser (outer-inner) | BPD (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD (outer-inner), Lasser 1993;99SIEMENS;GA107-14-01 | 2 Sigma deviation of population ;DCM;121417
Rempen | BPD (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD, Rempen 2001;99SIEMENS;GA107-25-01 | 2 Sigma deviation of population ;DCM;121417
Merz | BPD (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD, Merz 1991;99SIEMENS;GA107-16-01 | 2 Sigma deviation of population ;DCM;121417
Hansmann | BPD (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD, Hansmann 1986;LN;33538-0 | 2 Sigma deviation of population ;DCM;121417
JSUM 2003 | BPD (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD, JSUM 2003;99SIEMENS;GA107-12-01 | 2 Sigma deviation of population ;DCM;121417
Chitty (outer-inner) | BPD (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD-oi, Chitty 1997;LN;33086-0 | 2 Sigma deviation of population ;DCM;121417
Chitty (outer-outer) | BPD (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD-oo, Chitty 1997;LN;33087-8 | 2 Sigma deviation of population ;DCM;121417
ASUM 2003 | BPD (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=BPD, ASUM 2003;99SIEMENS;GA107-01-01
Osaka* | BPD (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD, Osaka 1983;99SIEMENS;GA107-18-01
Hadlock | HC (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=HC, Hadlock 1984;LN;11932-1 | 2 Sigma deviation of population ;DCM;121417
Merz | HC (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC, Merz 1991;99SIEMENS;GA116-16-01 | 2 Sigma deviation of population ;DCM;121417
Lasser | HC (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC, Lasser 1993;99SIEMENS;GA116-14-01 | 2 Sigma deviation of population ;DCM;121417
Hansmann | HC (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC, Hansmann 1986;LN;33543-0 | 2 Sigma deviation of population ;DCM;121417
ASUM 2003 | HC (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=HC, ASUM 2003;99SIEMENS;GA116-01-01 | 2 Sigma deviation of population ;DCM;121417
Chitty | HC (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC measured, Chitty 1997;LN;33110-8 | 2 Sigma deviation of population ;DCM;121417
Jeanty | HC (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC, Jeanty 2015;99SIEMENS;GA102-19-01 | 2 Sigma deviation of population ;DCM;121417
ASUM 2003 | HC (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=HC, ASUM 2003;99SIEMENS;GA116-01-01
Hadlock | HC (c) (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=HC, Hadlock 1984;LN;11932-1 | 2 Sigma deviation of population ;DCM;121417
Merz | HC (c) (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC, Merz 1991;99SIEMENS;GA116-16-01 | 2 Sigma deviation of population ;DCM;121417
Hansmann | HC (c) (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC, Hansmann 1986;LN;33543-0 | 2 Sigma deviation of population ;DCM;121417
Chitty(Derived) | HC (c) (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC derived, Chitty 1997;LN;33111-6 | 2 Sigma deviation of population ;DCM;121417
Jeanty | HC (c) (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC, Jeanty 2015;99SIEMENS;GA102-19-01 | 2 Sigma deviation of population ;DCM;121417
Hadlock | AC (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=AC, Hadlock 1984;LN;11892-7 | 2 Sigma deviation of population ;DCM;121417
Lasser | AC (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC, Lasser 1993;99SIEMENS;GA102-14-01 | 2 Sigma deviation of population ;DCM;121417
Merz | AC (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC, Merz 1991;99SIEMENS;GA102-16-01 | 2 Sigma deviation of population ;DCM;121417
JSUM 2003 | AC (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC, JSUM 2003;99SIEMENS;GA102-12-01 | 2 Sigma deviation of population ;DCM;121417
ASUM 2003 | AC (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=AC, ASUM 2003;99SIEMENS;GA102-01-01
Chitty Plotted | AC (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC, Chitty Plotted 1994;99SIEMENS;GA102-17-01 | 2 Sigma deviation of population ;DCM;121417
Hadlock | AC (c) GA |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=AC, Hadlock 1984;LN;11892-7 | 2 Sigma deviation of population ;DCM;121417
Merz | AC (c) GA |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC, Merz 1991;99SIEMENS;GA102-16-01 | 2 Sigma deviation of population ;DCM;121417
JSUM 2003 | AC (c) GA |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC, JSUM 2003;99SIEMENS;GA102-12-01 | 2 Sigma deviation of population ;DCM;121417
ASUM 2003 | AC (c) GA |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=AC, ASUM 2003;99SIEMENS;GA102-01-01
Chitty(Derived) | AC (c) GA |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC, Chitty Derived 1994;99SIEMENS;GA102-18-01
Hadlock | FL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=FL, Hadlock 1984;LN;11920-6 | 2 Sigma deviation of population ;DCM;121417
Jeanty | FL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=FL, Jeanty 1984;LN;11923-0 | 2 Sigma deviation of population ;DCM;121417
Merz | FL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FL, Merz 1991;99SIEMENS;GA112-16-01 | 2 Sigma deviation of population ;DCM;121417
Warda | FL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=FL, Warda;99SIEMENS;GA112-23-01 | 2 Sigma deviation of population ;DCM;121417
JSUM 2003 | FL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FL, JSUM 2003;99SIEMENS;GA112-12-01 | 2 Sigma deviation of population ;DCM;121417
Tokyo* | FL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FL, Tokyo 1986;LN;33103-3
Osaka* | FL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FL, Osaka 1983;99SIEMENS;GA112-18-01
ASUM 2003 | FL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=FL, ASUM 2003;99SIEMENS;GA112-01-01
Chitty | FL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FL, Chitty 1997;LN;33098-5 | 2 Sigma deviation of population ;DCM;121417
Hadlock | CRL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Early Gestation;DCM;125009 |  | Equation;DCM;121420=CRL, Hadlock 1992;LN;11910-7 | 2 Sigma deviation of population ;DCM;121417
Robinson | CRL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Early Gestation;DCM;125009 |  | Equation;DCM;121420=CRL, Robinson 1975;LN;11914-9
Lasser | CRL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=CRL, Lasser 1993;99SIEMENS;GA111-14-01 | 2 Sigma deviation of population ;DCM;121417
Hansmann | CRL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=CRL, Hansmann 1986;LN;33540-6 | 2 Sigma deviation of population ;DCM;121417
Rempen | CRL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=CRL, Rempen 2001;99SIEMENS;GA111-25-01
Verburg | CRL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=CRL, Verburg 2008;99SIEMENS;GA111-22-01
JSUM 2003 | CRL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=CRL, JSUM 2003;99SIEMENS;GA111-12-01 | 2 Sigma deviation of population ;DCM;121417
Tokyo* | CRL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=CRL, Tokyo 1986;LN;33096-9
ASUM 2003 | CRL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Early Gestation;DCM;125009 |  | Equation;DCM;121420=CRL, ASUM 2003;99SIEMENS;GA111-01-01
Osaka* | CRL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=CRL, Osaka 1983;99SIEMENS;GA111-18-01
Tongsong | Binoc D (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Cranium;DCM;125004 |  | Equation;DCM;121420=Binocular Distance, Tongsong 1992;99SIEMENS;GA106-21-01 | 2 Sigma deviation of population ;DCM;121417
Jeanty | Binoc D (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Cranium;DCM;125004 |  | Table of Values;DCM;121424=Binocular Distance, Jeanty 1984;LN;11896-8 | 2 Sigma deviation of population ;DCM;121417
Rempen | GSD(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=GS, Rempen 1991;LN;11929-7 | 2 Sigma deviation of population ;DCM;121417
Tokyo* | GSD(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=GS, Tokyo 1986;LN;33108-2
Jeanty | HL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Equation;DCM;121420=Humerus, Jeanty 1984;LN;11936-2 | 2 Sigma deviation of population ;DCM;121417
Osaka* | HL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Table of Values;DCM;121424=HL, Osaka 1983;99SIEMENS;GA117-18-01
ASUM 2003 | HL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Equation;DCM;121420=HL, ASUM 2003;99SIEMENS;GA117-01-01
Jeanty | Ulna (GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Equation;DCM;121420=Ulna, Jeanty 1984;LN;11944-6 | 2 Sigma deviation of population ;DCM;121417
Jeanty | Tibia(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Equation;DCM;121420=Tibia, Jeanty 1984;LN;11941-2 | 2 Sigma deviation of population ;DCM;121417
Mercer | Foot Length(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=Foot Length, Mercer 1987;LN;11926-3 | 2 Sigma deviation of population ;DCM;121417
ASUM 2003 | OFD(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=OFD, ASUM 2003;99SIEMENS;GA119-01-01
Osaka* | FTA(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FTA, Osaka 1983;99SIEMENS;GA114-18-01
Tokyo* | EFW(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Summary;DCM;121111 |  | Table of Values;DCM;121424=Gestational Age by EFW, Tokyo 1987;99SIEMENS;GA101-20-01
Osaka* | EFW(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Summary;DCM;121111 |  | Table of Values;DCM;121424=Gestational Age by EFW, Osaka 1983;99SIEMENS;GA101-18-01
JSUM 2003* | EFW(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Summary;DCM;121111 |  | Table of Values;DCM;121424=Gestational Age by EFW, JSUM 2003;99SIEMENS;GA101-12-01
Tokyo* | APTDxTTD(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=APTDxTTD, Tokyo;99SIEMENS;GA105-20-01
Jeanty | Ulna(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Table of Values;DCM;121424=Ulna by GA, Jeanty 1984;99SIEMENS;GR124-11-01
Hansmann | Ulna(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Table of Values;DCM;121424=Ulna by GA, Hansmann 1986;99SIEMENS;GR124-09-01
Chitty | Ulna(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Table of Values;DCM;121424=Ulna by GA, Chitty 2002;99SIEMENS;GR124-05-01
Merz | Ulna(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Table of Values;DCM;121424=Ulna by GA, Merz 1996;99SIEMENS;GR124-16-01
Chitkara | TC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=TC by GA, Chitkara 1987;99SIEMENS;GR122-04-01
Kurmanavicius | MAD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=MAD by GA, Kurmanavicius 1997;99SIEMENS;GR118-13-01
JSUM 2003 RI | Rt MCA RI(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800 |  | Equation;DCM;121420=Right MCA RI by GA JSUM 2003;99SIEMENS;GR132-12-01
JSUM 2003 PI | Rt MCA PI(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800 |  | Equation;DCM;121420=Right MCA PI by GA JSUM 2003;99SIEMENS;GR131-12-01
Ebbing | Rt MCA PI(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800 |  | Table of Values;DCM;121424=Right MCA PI by GA Ebbing;99SIEMENS;RightMCAPIEbbing
JSUM 2003 RI | Lt MCA RI(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800 |  | Equation;DCM;121420=Left MCA RI by GA JSUM 2003;99SIEMENS;GR130-12-01
JSUM 2003 PI | Lt MCA PI(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800 |  | Equation;DCM;121420=Left MCA PI by GA JSUM 2003;99SIEMENS;GR129-12-01
Ebbing | Lt MCA PI(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800 |  | Table of Values;DCM;121424=Left MCA PI by GA Ebbing;99SIEMENS;LeftMCAPIEbbing
JSUM 2003 RI | Umb A RI(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800 |  | Equation;DCM;121420=Umbilical Artery RI by GA JSUM 2003;99SIEMENS;GR134-12-01
JSUM 2003 PI | Umb A PI(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800 |  | Equation;DCM;121420=Umbilical Artery PI by GA JSUM 2003;99SIEMENS;GR133-12-01
Hadlock | CRL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Early Gestation;DCM;125009 |  | Equation;DCM;121420=CRL by GA, Hadlock 1982;99SIEMENS;GR111-08-01
Rempen | CRL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=CRL by GA, Rempen 1991;LN;33160-3
Robinson | CRL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Early Gestation;DCM;125009 |  | Equation;DCM;121420=CRL by GA, Robinson 1975;99SIEMENS;GR111-19-01
Hansmann | CRL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=CRL by GA, Hansmann 1986;99SIEMENS;GR111-09-01
JSUM 2003 | CRL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=CRL by GA, JSUM 2003;99SIEMENS;GR111-12-01
Tokyo* | CRL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=CRL by GA, Tokyo;99SIEMENS;GR111-20-01
Osaka | CRL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=CRL by GA, Osaka 1983;99SIEMENS;GR111-18-01
ASUM 2003 | CRL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=CRL by GA, ASUM 2003;99SIEMENS;GR111-01-01
Pexsters | CRL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=CRL by GA, Pexsters 2010;99SIEMENS;GR111-28-01
Hadlock | BPD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=BPD by GA, Hadlock 1984;LN;33198-3
Verburg | BPD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD by GA, Verburg 2008;99SIEMENS;GR107-22-01
Hansmann | BPD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD by GA, Hansmann 1986;99SIEMENS;GR107-09-01
Lasser | BPD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=BPD by GA, Lasser 1993;99SIEMENS;GR107-14-01
Merz | BPD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD by GA, Merz 1996;99SIEMENS;GR107-16-01
Rempen | BPD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD by GA, Rempen 1991;LN;33155-3
Chitty (outer-inner) | BPD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=BPD outer-inner by GA, Chitty 1994;LN;33556-2
Chitty (outer-outer) | BPD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=BPD outer-outer by GA, Chitty 1994;LN;33152-0
Grumbach (twins) | BPD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD by GA, Grumbach 1986 (twin);99SIEMENS;GR107-07-01
JSUM 2003 | BPD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD by GA, JSUM 2003;99SIEMENS;GR107-12-01
Osaka | BPD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD by GA, Osaka 1983;99SIEMENS;GR107-18-01
ASUM 2003 | BPD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD by GA, ASUM 2003;99SIEMENS;GR107-01-01
CFEF | BPD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD by GA, CFEF 2000;99SIEMENS;GR107-03-01
Nicolaides | BPD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD by GA, Nicolaides 1994;99SIEMENS;GR107-17-01
Hadlock | HC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=HC by GA, Hadlock 1984;LN;33173-6
Hansmann | HC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC by GA, Hansmann 1986;99SIEMENS;GR116-09-01
Merz | HC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC by GA, Merz 1996;99SIEMENS;GR116-16-01
Chitty | HC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=HC by GA, Chitty 1994;99SIEMENS;GR116-05-01
Deter | HC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=HC by GA, Deter 1982;99SIEMENS;GR116-06-01
ASUM 2003 | HC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC by GA, ASUM 2003;99SIEMENS;GR116-01-01
CFEF | HC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC by GA, CFEF 2000;99SIEMENS;GR116-03-01
Verburg | HC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC by GA, Verburg 2008;99SIEMENS;GR116-22-01
Nicolaides | HC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC by GA, Nicolaides 1994;99SIEMENS;GR116-17-01
Hadlock | HC(c)(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=HC by GA, Hadlock 1984;LN;33173-6
Hansmann | HC(c)(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC by GA, Hansmann 1986;99SIEMENS;GR116-09-01
Merz | HC(c)(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC by GA, Merz 1996;99SIEMENS;GR116-16-01
Chitty(Derived) | HC(c)(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=HC derived by GA, Chitty 1994;LN;33174-4
Deter | HC(c)(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=HC by GA, Deter 1982;99SIEMENS;GR116-06-01
ASUM 2003 | HC(c)(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC by GA, ASUM 2003;99SIEMENS;GR116-01-01
CFEF | HC(c)(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC by GA, CFEF 2000;99SIEMENS;GR116-03-01
Verburg | HC(c)(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC by GA, Verburg 2008;99SIEMENS;GR116-22-01
Nicolaides | HC(c)(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC by GA, Nicolaides 1994;99SIEMENS;GR116-17-01
Hadlock | AC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=AC by GA, Hadlock 1984;LN;33146-2
Nicolaides | AC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, Nicolaides 1994;99SIEMENS;GR102-17-01
Jeanty | AC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, Jeanty 1984;99SIEMENS;GR102-11-01
Merz | AC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, Merz 1996;99SIEMENS;GR102-16-01
Grumbach (twins) | AC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, Grumbach 1986 (twin);99SIEMENS;GR102-07-01
Verburg | AC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, Verburg 2008;99SIEMENS;GR102-22-01
Chitty | AC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=AC (measured) by GA, Chitty 1994;LN;33147-0
Deter | AC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=AC by GA, Deter 1982;99SIEMENS;GR102-06-01
JSUM 2003 | AC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, JSUM 2003;99SIEMENS;GR102-12-01
ASUM 2003 | AC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, ASUM 2003;99SIEMENS;GR102-01-01
CFEF | AC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, CFEF 2000;99SIEMENS;GR102-03-01
Hadlock | AC(c)GC |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=AC by GA, Hadlock 1984;LN;33146-2
Nicolaides | AC(c)GC |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, Nicolaides 1994;99SIEMENS;GR102-17-01
Jeanty | AC(c)GC |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, Jeanty 1984;99SIEMENS;GR102-11-01
Merz | AC(c)GC |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, Merz 1996;99SIEMENS;GR102-16-01
Grumbach (twins) | AC(c)GC |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, Grumbach 1986 (twin);99SIEMENS;GR102-07-01
Verburg | AC(c)GC |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, Verburg 2008;99SIEMENS;GR102-22-01
Chitty(Derived) | AC(c)GC |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=AC (derived) by GA, Chitty 1994;LN;33546-3
Deter | AC(c)GC |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=AC by GA, Deter 1982;99SIEMENS;GR102-06-01
JSUM 2003 | AC(c)GC |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, JSUM 2003;99SIEMENS;GR102-12-01
ASUM 2003 | AC(c)GC |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, ASUM 2003;99SIEMENS;GR102-01-01
CFEF | AC(c)GC |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, CFEF 2000;99SIEMENS;GR102-03-01
Moore | AFI(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Amniotic Sac;DCM;125002 |  | Table of Values;DCM;121424=AFI by GA, Moore 1990;99SIEMENS;GR135-27-01
Moore | AFI(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Amniotic Sac;DCM;125002 |  | Table of Values;DCM;121424=AFI by GA, Moore 1990;99SIEMENS;GR135-27-01
Moore | AFI(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Amniotic Sac;DCM;125002 |  | Table of Values;DCM;121424=AFI by GA, Moore 1990;99SIEMENS;GR135-27-01
Hadlock | FL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=FL by GA, Hadlock 1984;LN;33166-0
Jeanty | FL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FL by GA, Jeanty 1984;99SIEMENS;GR112-11-01
Merz | FL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FL by GA, Merz 1991;99SIEMENS;GR112-16-01
Chitty | FL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=FL by GA, Chitty 1994;LN;33167-8
Warda | FL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=FL by GA, Warda 1985;99SIEMENS;GR112-23-01
JSUM 2003 | FL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FL by GA, JSUM 2003;99SIEMENS;GR112-12-01
Tokyo | FL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FL by GA, Tokyo;99SIEMENS;GR112-20-01
Osaka | FL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Equation;DCM;121420=FL by GA, Osaka 1983;99SIEMENS;GR112-18-01
ASUM 2003 | FL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FL by GA, ASUM 2003;99SIEMENS;GR112-01-01
CFEF | FL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FL by GA, CFEF 2000;99SIEMENS;GR112-03-01
Hansmann | FL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FL by GA, Hansmann 1986;99SIEMENS;GR112-09-01
Nicolaides | FL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FL by GA, Nicolaides 1994;99SIEMENS;GR112-17-01
Verburg | FL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FL by GA, Verburg 2008;99SIEMENS;GR112-22-01
Mercer | Foot Length(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=Foot by GA, Mercer 1987;99SIEMENS;GR113-15-01
Chitty | Foot Length(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=Foot by GA, Chitty 2002;99SIEMENS;GR113-05-01
Hellman | GSD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Early Gestation;DCM;125009 |  | Equation;DCM;121420=GS by GA, Hellman 1969;99SIEMENS;GR115-10-01
Tokyo* | GSD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=GS by GA, Tokyo;99SIEMENS;GR115-20-01
Rempen | GSD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=GS by GA, Rempen 1991;LN;33171-0
Jeanty | HL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Table of Values;DCM;121424=HL by GA, Jeanty 1984;99SIEMENS;GR117-11-01
Chitty | HL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Table of Values;DCM;121424=HL by GA, Chitty 2002;99SIEMENS;GR117-05-01
Merz | HL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Table of Values;DCM;121424=HL by GA, Merz 1996;99SIEMENS;GR117-16-01
Osaka | HL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Table of Values;DCM;121424=HL by GA, Osaka 1983;99SIEMENS;GR117-18-01
ASUM 2003 | HL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Table of Values;DCM;121424=HL by GA, ASUM 2003;99SIEMENS;GR117-01-01
Hansmann | HL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Table of Values;DCM;121424=HL by GA, Hansmann 1986;99SIEMENS;GR117-09-01
Hadlock | EFW(GC) |  | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Summary;DCM;121111 |  | Equation;DCM;121420=EFW by GA, Hadlock 1991;LN;33162-9
Jeanty | EFW(GC) |  | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Summary;DCM;121111 |  | Table of Values;DCM;121424=FWP by GA, Jeanty 1994;99SIEMENS;GR101-11-01
Yarkoni (twins) | EFW(GC) |  | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Summary;DCM;121111 |  | Table of Values;DCM;121424=FWP by GA, Yarkoni 1987;99SIEMENS;GR101-24-01
JSUM 2003 | EFW(GC) |  | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Summary;DCM;121111 |  | Table of Values;DCM;121424=FWP by GA, JSUM 2003;99SIEMENS;GR101-12-01
Tokyo* | EFW(GC) |  | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Summary;DCM;121111 |  | Table of Values;DCM;121424=FWP by GA, Tokyo;99SIEMENS;GR101-20-01
Osaka | EFW(GC) |  | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Summary;DCM;121111 |  | Table of Values;DCM;121424=FWP by GA, Osaka 1983;99SIEMENS;GR101-18-01
Williams Female | EFW(GC) |  | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Summary;DCM;121111 |  | Table of Values;DCM;121424=FWP by GA, Williams 1982 (female);LN;33184-3
Brenner | EFW(GC) |  | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Summary;DCM;121111 |  | Table of Values;DCM;121424=FWP by GA, Brenner 1976;LN;33189-2
Doubilet | EFW(GC) |  | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Summary;DCM;121111 |  | Table of Values;DCM;121424=FWP by GA, Doubilet 1996;99SIEMENS;GR101-26-01
Campbell | HC/AC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry Ratios;DCM;125001 |  | Table of Values;DCM;121424=HC/AC by GA, Campbell 1977;LN;33182-7
Osaka | FTA(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FTA by GA, Osaka 1983;99SIEMENS;GR114-18-01
Tokyo | APTDXTTD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=APTDxTTD by GA, Tokyo;99SIEMENS;GR105-20-01
CFEF | TAD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=TAD by GA, CFEF 2000;99SIEMENS;GR121-03-01
Nicolaides | LV -A(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Cranium;DCM;125004 |  | Table of Values;DCM;121424=Ant CV by GA, Nicolaides 1994;99SIEMENS;GR103-17-01
Nicolaides | LV -P(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Cranium;DCM;125004 |  | Table of Values;DCM;121424=Post CV by GA, Nicolaides 1994;99SIEMENS;GR120-17-01
Merz | Tibia Length(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Table of Values;DCM;121424=Tibia by GA, Merz 1996;99SIEMENS;GR123-16-01
Jeanty | Tibia Length(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Table of Values;DCM;121424=Tibia by GA, Jeanty 1984;99SIEMENS;GR123-11-01
Hansmann | Tibia Length(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Table of Values;DCM;121424=Tibia by GA, Hansmann 1986;99SIEMENS;GR123-09-01
Chitty | Tibia Length(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Table of Values;DCM;121424=Tibia by GA, Chitty 2002;99SIEMENS;GR123-05-01
ASUM 2003 | OFD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=OFD by GA, ASUM 2003;99SIEMENS;GR119-01-01
Hansmann | OFD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=OFD by GA, Hansmann 1986;99SIEMENS;GR119-09-01
Nicolaides | OFD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=OFD by GA, Nicolaides 1994;99SIEMENS;GR119-17-01
Chitty | OFD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=OFD by GA, Chitty 1994;LN;33179-3
Sherer | Clavicle(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Long Bones;DCM;125003 |  | Equation;DCM;121420=Clavicle by GA, Sherer 2006;99SIEMENS;GR110-30-01
Nicolaides | TCD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Cranium;DCM;125004 |  | Table of Values;DCM;121424=TCD by GA, Nicolaides 1994;99SIEMENS;GR108-17-01
Verburg | TCD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Cranium;DCM;125004 |  | Table of Values;DCM;121424=TCD by GA, Verburg 2008;99SIEMENS;GR108-22-01
Nicolaides | Cist Magna Length(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Cranium;DCM;125004 |  | Table of Values;DCM;121424=Cisterna Magna by GA, Nicolaides 1994;99SIEMENS;GR109-17-01
Bertagnoli | Lt Kidney AP(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=Left Kidney AP by GA, Bertagnoli 1983;99SIEMENS;GR125-02-01
Hansmann | Lt Kidney AP(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=Left Kidney AP by GA, Hansmann 1986;99SIEMENS;GR125-09-01
Bertagnoli | Rt Kidney AP(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=Right Kidney AP by GA, Bertagnoli 1983;99SIEMENS;GR127-02-01
Hansmann | Rt Kidney AP(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=Right Kidney AP by GA, Hansmann 1986;99SIEMENS;GR127-09-01
Bertagnoli | Lt Kidney Length(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=Left Kidney Length by GA, Bertagnoli 1983;99SIEMENS;GR126-02-01
Hansmann | Lt Kidney Length(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=Left Kidney Length by GA, Hansmann 1986;99SIEMENS;GR126-09-01
Bertagnoli | Rt Kidney Length(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=Right Kidney Length by GA, Bertagnoli 1983;99SIEMENS;GR128-02-01
Hansmann | Rt Kidney Length(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=Right Kidney Length by GA, Hansmann 1986;99SIEMENS;GR128-09-01
Average | US Age |  | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Summary;DCM;121111 |  | Equation;DCM;121420=Average Ultrasound Age;LN;11884-4
Hadlock (BPD, HC, AC, FL) | US Age |  | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Summary;DCM;121111 |  | Equation;DCM;121420=Composite US Age, Hadlock (BPD, HC, AC, FL);99SIEMENS;CA100-08-07
Hadlock (BPD, HC, AC) | US Age |  | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Summary;DCM;121111 |  | Equation;DCM;121420=Composite US Age, Hadlock (BPD, HC, AC);99SIEMENS;CA100-08-06
Hadlock (HC, AC, FL) | US Age |  | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Summary;DCM;121111 |  | Equation;DCM;121420=Composite US Age, Hadlock (HC, AC, FL);99SIEMENS;CA100-08-10
Hadlock (BPD, HC, FL) | US Age |  | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Summary;DCM;121111 |  | Equation;DCM;121420=Composite US Age, Hadlock (BPD, HC, FL);99SIEMENS;CA100-08-08
Hadlock (BPD, AC, FL) | US Age |  | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Summary;DCM;121111 |  | Equation;DCM;121420=Composite US Age, Hadlock (BPD, AC, FL);99SIEMENS;CA100-08-03
Hadlock (AC, FL) | US Age |  | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Summary;DCM;121111 |  | Equation;DCM;121420=Composite US Age, Hadlock (AC, FL);99SIEMENS;CA100-08-01
Hadlock (HC, FL) | US Age |  | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Summary;DCM;121111 |  | Equation;DCM;121420=Composite US Age, Hadlock (HC, FL);99SIEMENS;CA100-08-11
Hadlock (HC, AC) | US Age |  | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Summary;DCM;121111 |  | Equation;DCM;121420=Composite US Age, Hadlock (HC, AC);99SIEMENS;CA100-08-09
Hadlock (BPD, FL) | US Age |  | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Summary;DCM;121111 |  | Equation;DCM;121420=Composite US Age, Hadlock (BPD, FL);99SIEMENS;CA100-08-04
Hadlock (BPD, HC) | US Age |  | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Summary;DCM;121111 |  | Equation;DCM;121420=Composite US Age, Hadlock (BPD, HC);99SIEMENS;CA100-08-05
Hadlock (BPD, AC) | US Age |  | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Summary;DCM;121111 |  | Equation;DCM;121420=Composite US Age, Hadlock (BPD, AC);99SIEMENS;CA100-08-02
Hadlock (BPD, HC, AC, FL) | EFW |  | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Summary;DCM;121112 |  | Equation;DCM;121420=EFW by AC, BPD, FL, HC, Hadlock 1985;LN;11732-5
Hadlock (BPD, AC, FL) | EFW |  | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Summary;DCM;121113 |  | Equation;DCM;121420=EFW by AC, BPD, FL, Hadlock 1985;LN;11735-8
Hadlock (HC, AC, FL) | EFW |  | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Summary;DCM;121114 |  | Equation;DCM;121420=EFW by AC, FL, HC, Hadlock 1985;LN;11746-5
Hadlock (AC, FL) | EFW |  | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Summary;DCM;121115 |  | Equation;DCM;121420=EFW by AC, FL, Hadlock 1985;LN;11751-5
Shepard (BPD, AC) | EFW |  | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Summary;DCM;121116 |  | Equation;DCM;121420=EFW by AC and BPD, Shepard 1982;LN;11739-0
Merz (BPD, AC) | EFW |  | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Summary;DCM;121117 |  | Equation;DCM;121420=EFW by BPD, AC, Merz 2005;99SIEMENS;EFW101-16-01
Tokyo (BPD, APTD, TTD, FL) | EFW |  | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Summary;DCM;121118 |  | Equation;DCM;121420=EFW by BPD, APTD, TTD, FL, Tokyo 1987;99SIEMENS;EFW101-20-01
Osaka (BPD, FTA, FL) | EFW |  | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Summary;DCM;121119 |  | Equation;DCM;121420=EFW by BPD, FTA, FL, Osaka 1990;LN;33140-5
JSUM (BPD, AC, FL) | EFW |  | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Summary;DCM;121120 |  | Equation;DCM;121420=EFW2 by Shinozuka 1996;LN;33142-1
JSUM (BPD, APTD, TTD, FL) | EFW |  | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Summary;DCM;121121 |  | Equation;DCM;121420=EFW by BPD, APTD, TTD, FL, JSUM 2003;99SIEMENS;EFW101-12-01
Hansmann (BPD, TTD) | EFW |  | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Summary;DCM;121122 |  | Equation;DCM;121420=EFW by BPD, TTD, Hansmann 1986;LN;33139-7
INTERGROWTH-21st | CRL(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Early Gestation;DCM;125009 |  | Equation;DCM;121420=CRL, INTERGROWTH-21st 2014;99SIEMENS;GA111-29-01 | 2 Sigma deviation of population ;DCM;121417
INTERGROWTH-21st | CRL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=CRL by GA, INTERGROWTH-21st 2014;99SIEMENS;GR111-29-01
INTERGROWTH-21st | BPD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=BPD by GA, INTERGROWTH-21st 2014;99SIEMENS;GR107-29-01
INTERGROWTH-21st | HC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC by GA, INTERGROWTH-21st 2014;99SIEMENS;GR116-29-01
INTERGROWTH-21st | HC(c )(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=HC by GA, INTERGROWTH-21st 2014;99SIEMENS;GR116-29-01
INTERGROWTH-21st | AC(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, INTERGROWTH-21st 2014;99SIEMENS;GR102-29-01
INTERGROWTH-21st | AC(c)(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=AC by GA, INTERGROWTH-21st 2014;99SIEMENS;GR102-29-01
INTERGROWTH-21st | FL(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=FL by GA, INTERGROWTH-21st 2014;99SIEMENS;GR112-29-01
INTERGROWTH-21st | OFD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetal Biometry;DCM;125002 |  | Table of Values;DCM;121424=OFD by GA, INTERGROWTH-21st 2014;99SIEMENS;GR119-29-01
PreVoid Bladder | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Prior to voiding;DCM;109134 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
PostVoid Bladder | Volume | Vol | cm³ | 2 | 1 | Volume;SRT;G-D705 | NUM | Post voiding;DCM;109135 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
Micturated Vol | Micturated Volume |  | cm³ | 2 | 1 | Micturated Volume;99SIEMENS;MicturatedVol | NUM | Bladder;SRT;T-74000 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
US EDD | Estimated Delivery Date from average ultrasound age |  | MCR_DATE |  | 1 | EDD from average ultrasound age;LN;11781-2 | DATE | Summary;DCM;121111 |  | Summary;DCM;121111
Williams Male | EFW% |  | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Summary;DCM;121111 |  | Table of Values;DCM;121424=FWP by GA, Williams 1982 (male);LN;33184-3
LVOT Area s | Left Ventricular Outflow Tract Area s (derived) |  | cm² | 1 | 1 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Left Ventricle;SRT;T-32600 |  | Finding Site;SRT;G-C0E3=Fetal Heart;99SIEMENS;FetalHeart |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 | Cardiac Cycle Point;SRT;R-4089A=End Systole;SRT;R-FAB5B | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVOT Area s | Right Ventricular Outflow Tract Area systole (derived) |  | cm² | 1 | 1 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Right Ventricle;SRT;T-32500 |  | Finding Site;SRT;G-C0E3=Fetal Heart;99SIEMENS;FetalHeart |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
CTR Area | Cardiothoracic Area Ratio |  | unitless | 1 | 1 | Cardiothoracic Area Ratio;LN;59076-0 | NUM | Fetal Heart;99SIEMENS;FetalHeart |  | Finding Site;SRT;G-C0E3=Fetal Heart;99SIEMENS;FetalHeart
CTR Circum | Cardiothoracic Circumference Ratio |  | unitless | 1 | 1 | Cardiothoracic Circumference Ratio;LN;59074-5 | NUM | Fetal Heart;99SIEMENS;FetalHeart |  | Finding Site;SRT;G-C0E3=Fetal Heart;99SIEMENS;FetalHeart
Rempen | MSD(GA) |  | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=MSD, Rempen 1991;99SIEMENS;GA136-25-01 | 2 Sigma deviation of population ;DCM;121417
Hellman | MSD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Early Gestation;DCM;125009 |  | Equation;DCM;121420=MSD by GA, Hellman 1969;99SIEMENS;GR136-10-01
Rempen | MSD(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Early Gestation;DCM;125009 |  | Table of Values;DCM;121424=MSD by GA, Rempen 1991;99SIEMENS;GR136-25-01
Ductus V S/D Vmax | Ductus Venosus S Vmax/D Vmax |  | unitless | 1 | 1 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Ductus venosus;SRT;T-F6806 |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800
Lt Breast Mass 1 | Left Breast Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 1 | Left Breast Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 1 | Left Breast Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 1 | Left Breast Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 1 | Left Breast Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 1 | Left Breast Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 1 | Left Breast Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 1 | Left Breast Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 1 | Left Breast Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 1 | Left Breast Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 1 | Left Breast Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 1 | Left Breast Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 1 | Left Breast Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 1 | Left Breast Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 1 | Left Breast Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 1 | Left Breast Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 1 | Left Breast Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 1 | Left Breast Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 1 | Left Breast Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 1 | Left Breast Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 1 | Left Breast Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 1 | Left Breast Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 1 | Left Breast Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 1 | Left Breast Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 1 | Left Breast Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 1 | Left Breast Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 1 | Left Breast Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 1 | Left Breast Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 1 | Left Breast Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 1 | Left Breast Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 2 | Left Breast Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 2 | Left Breast Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 2 | Left Breast Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 2 | Left Breast Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 2 | Left Breast Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 2 | Left Breast Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 2 | Left Breast Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 2 | Left Breast Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 2 | Left Breast Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 2 | Left Breast Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 2 | Left Breast Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 2 | Left Breast Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 2 | Left Breast Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 2 | Left Breast Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 2 | Left Breast Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 2 | Left Breast Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 2 | Left Breast Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 2 | Left Breast Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 2 | Left Breast Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 2 | Left Breast Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 2 | Left Breast Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 2 | Left Breast Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 2 | Left Breast Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 2 | Left Breast Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 2 | Left Breast Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 2 | Left Breast Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 2 | Left Breast Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 2 | Left Breast Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 2 | Left Breast Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 2 | Left Breast Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 3 | Left Breast Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 3 | Left Breast Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 3 | Left Breast Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 3 | Left Breast Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 3 | Left Breast Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 3 | Left Breast Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 3 | Left Breast Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 3 | Left Breast Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 3 | Left Breast Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 3 | Left Breast Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 3 | Left Breast Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 3 | Left Breast Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 3 | Left Breast Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 3 | Left Breast Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 3 | Left Breast Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 3 | Left Breast Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 3 | Left Breast Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 3 | Left Breast Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 3 | Left Breast Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 3 | Left Breast Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 3 | Left Breast Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 3 | Left Breast Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 3 | Left Breast Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 3 | Left Breast Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 3 | Left Breast Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 3 | Left Breast Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 3 | Left Breast Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 3 | Left Breast Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 3 | Left Breast Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 3 | Left Breast Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 4 | Left Breast Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 4 | Left Breast Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 4 | Left Breast Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 4 | Left Breast Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 4 | Left Breast Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 4 | Left Breast Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 4 | Left Breast Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 4 | Left Breast Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 4 | Left Breast Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 4 | Left Breast Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 4 | Left Breast Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 4 | Left Breast Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 4 | Left Breast Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 4 | Left Breast Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 4 | Left Breast Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 4 | Left Breast Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 4 | Left Breast Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 4 | Left Breast Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 4 | Left Breast Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 4 | Left Breast Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 4 | Left Breast Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 4 | Left Breast Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 4 | Left Breast Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 4 | Left Breast Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 4 | Left Breast Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 4 | Left Breast Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 4 | Left Breast Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 4 | Left Breast Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 4 | Left Breast Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 4 | Left Breast Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 5 | Left Breast Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 5 | Left Breast Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 5 | Left Breast Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 5 | Left Breast Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 5 | Left Breast Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 5 | Left Breast Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 5 | Left Breast Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 5 | Left Breast Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 5 | Left Breast Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 5 | Left Breast Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Mass 5 | Left Breast Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 5 | Left Breast Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 5 | Left Breast Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 5 | Left Breast Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 5 | Left Breast Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 5 | Left Breast Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 5 | Left Breast Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 5 | Left Breast Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 5 | Left Breast Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 5 | Left Breast Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Mass 5 | Left Breast Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 5 | Left Breast Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 5 | Left Breast Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 5 | Left Breast Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 5 | Left Breast Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 5 | Left Breast Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 5 | Left Breast Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 5 | Left Breast Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 5 | Left Breast Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Mass 5 | Left Breast Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 1 | Right Breast Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 1 | Right Breast Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 1 | Right Breast Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 1 | Right Breast Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 1 | Right Breast Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 1 | Right Breast Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 1 | Right Breast Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 1 | Right Breast Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 1 | Right Breast Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 1 | Right Breast Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 1 | Right Breast Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 1 | Right Breast Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 1 | Right Breast Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 1 | Right Breast Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 1 | Right Breast Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 1 | Right Breast Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 1 | Right Breast Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 1 | Right Breast Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 1 | Right Breast Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 1 | Right Breast Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 1 | Right Breast Mass 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 1 | Right Breast Mass 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 1 | Right Breast Mass 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 1 | Right Breast Mass 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 1 | Right Breast Mass 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 1 | Right Breast Mass 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 1 | Right Breast Mass 1 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 1 | Right Breast Mass 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 1 | Right Breast Mass 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 1 | Right Breast Mass 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 2 | Right Breast Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 2 | Right Breast Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 2 | Right Breast Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 2 | Right Breast Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 2 | Right Breast Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 2 | Right Breast Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 2 | Right Breast Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 2 | Right Breast Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 2 | Right Breast Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 2 | Right Breast Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 2 | Right Breast Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 2 | Right Breast Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 2 | Right Breast Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 2 | Right Breast Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 2 | Right Breast Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 2 | Right Breast Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 2 | Right Breast Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 2 | Right Breast Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 2 | Right Breast Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 2 | Right Breast Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 2 | Right Breast Mass 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 2 | Right Breast Mass 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 2 | Right Breast Mass 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 2 | Right Breast Mass 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 2 | Right Breast Mass 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 2 | Right Breast Mass 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 2 | Right Breast Mass 2 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 2 | Right Breast Mass 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 2 | Right Breast Mass 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 2 | Right Breast Mass 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 3 | Right Breast Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 3 | Right Breast Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 3 | Right Breast Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 3 | Right Breast Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 3 | Right Breast Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 3 | Right Breast Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 3 | Right Breast Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 3 | Right Breast Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 3 | Right Breast Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 3 | Right Breast Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 3 | Right Breast Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 3 | Right Breast Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 3 | Right Breast Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 3 | Right Breast Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 3 | Right Breast Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 3 | Right Breast Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 3 | Right Breast Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 3 | Right Breast Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 3 | Right Breast Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 3 | Right Breast Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 3 | Right Breast Mass 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 3 | Right Breast Mass 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 3 | Right Breast Mass 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 3 | Right Breast Mass 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 3 | Right Breast Mass 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 3 | Right Breast Mass 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 3 | Right Breast Mass 3 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 3 | Right Breast Mass 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 3 | Right Breast Mass 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 3 | Right Breast Mass 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 4 | Right Breast Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 4 | Right Breast Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 4 | Right Breast Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 4 | Right Breast Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 4 | Right Breast Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 4 | Right Breast Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 4 | Right Breast Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 4 | Right Breast Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 4 | Right Breast Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 4 | Right Breast Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 4 | Right Breast Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 4 | Right Breast Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 4 | Right Breast Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 4 | Right Breast Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 4 | Right Breast Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 4 | Right Breast Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 4 | Right Breast Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 4 | Right Breast Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 4 | Right Breast Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 4 | Right Breast Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 4 | Right Breast Mass 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 4 | Right Breast Mass 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 4 | Right Breast Mass 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 4 | Right Breast Mass 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 4 | Right Breast Mass 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 4 | Right Breast Mass 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 4 | Right Breast Mass 4 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 4 | Right Breast Mass 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 4 | Right Breast Mass 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 4 | Right Breast Mass 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 5 | Right Breast Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 5 | Right Breast Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 5 | Right Breast Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 5 | Right Breast Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 5 | Right Breast Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 5 | Right Breast Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 5 | Right Breast Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 5 | Right Breast Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 5 | Right Breast Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 5 | Right Breast Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Mass 5 | Right Breast Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 5 | Right Breast Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 5 | Right Breast Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 5 | Right Breast Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 5 | Right Breast Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 5 | Right Breast Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 5 | Right Breast Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 5 | Right Breast Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 5 | Right Breast Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 5 | Right Breast Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Mass 5 | Right Breast Mass 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 5 | Right Breast Mass 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 5 | Right Breast Mass 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 5 | Right Breast Mass 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 5 | Right Breast Mass 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 5 | Right Breast Mass 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 5 | Right Breast Mass 5 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 5 | Right Breast Mass 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 5 | Right Breast Mass 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Mass 5 | Right Breast Mass 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 1 | Left Breast Site 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 1 | Left Breast Site 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 1 | Left Breast Site 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 1 | Left Breast Site 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 1 | Left Breast Site 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 1 | Left Breast Site 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 1 | Left Breast Site 1 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 1 | Left Breast Site 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 1 | Left Breast Site 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 1 | Left Breast Site 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 1 | Left Breast Site 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 1 | Left Breast Site 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 1 | Left Breast Site 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 1 | Left Breast Site 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 1 | Left Breast Site 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 1 | Left Breast Site 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 1 | Left Breast Site 1 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 1 | Left Breast Site 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 1 | Left Breast Site 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 1 | Left Breast Site 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 1 | Left Breast Site 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 1 | Left Breast Site 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 1 | Left Breast Site 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 1 | Left Breast Site 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 1 | Left Breast Site 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 1 | Left Breast Site 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 1 | Left Breast Site 1 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 1 | Left Breast Site 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 1 | Left Breast Site 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 1 | Left Breast Site 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 2 | Left Breast Site 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 2 | Left Breast Site 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 2 | Left Breast Site 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 2 | Left Breast Site 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 2 | Left Breast Site 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 2 | Left Breast Site 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 2 | Left Breast Site 2 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 2 | Left Breast Site 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 2 | Left Breast Site 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 2 | Left Breast Site 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 2 | Left Breast Site 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 2 | Left Breast Site 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 2 | Left Breast Site 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 2 | Left Breast Site 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 2 | Left Breast Site 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 2 | Left Breast Site 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 2 | Left Breast Site 2 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 2 | Left Breast Site 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 2 | Left Breast Site 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 2 | Left Breast Site 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 2 | Left Breast Site 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 2 | Left Breast Site 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 2 | Left Breast Site 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 2 | Left Breast Site 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 2 | Left Breast Site 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 2 | Left Breast Site 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 2 | Left Breast Site 2 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 2 | Left Breast Site 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 2 | Left Breast Site 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 2 | Left Breast Site 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 3 | Left Breast Site 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 3 | Left Breast Site 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 3 | Left Breast Site 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 3 | Left Breast Site 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 3 | Left Breast Site 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 3 | Left Breast Site 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 3 | Left Breast Site 3 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 3 | Left Breast Site 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 3 | Left Breast Site 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 3 | Left Breast Site 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 3 | Left Breast Site 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 3 | Left Breast Site 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 3 | Left Breast Site 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 3 | Left Breast Site 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 3 | Left Breast Site 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 3 | Left Breast Site 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 3 | Left Breast Site 3 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 3 | Left Breast Site 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 3 | Left Breast Site 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 3 | Left Breast Site 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 3 | Left Breast Site 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 3 | Left Breast Site 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 3 | Left Breast Site 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 3 | Left Breast Site 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 3 | Left Breast Site 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 3 | Left Breast Site 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 3 | Left Breast Site 3 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 3 | Left Breast Site 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 3 | Left Breast Site 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 3 | Left Breast Site 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 4 | Left Breast Site 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 4 | Left Breast Site 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 4 | Left Breast Site 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 4 | Left Breast Site 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 4 | Left Breast Site 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 4 | Left Breast Site 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 4 | Left Breast Site 4 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 4 | Left Breast Site 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 4 | Left Breast Site 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 4 | Left Breast Site 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 4 | Left Breast Site 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 4 | Left Breast Site 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 4 | Left Breast Site 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 4 | Left Breast Site 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 4 | Left Breast Site 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 4 | Left Breast Site 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 4 | Left Breast Site 4 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 4 | Left Breast Site 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 4 | Left Breast Site 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 4 | Left Breast Site 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 4 | Left Breast Site 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 4 | Left Breast Site 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 4 | Left Breast Site 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 4 | Left Breast Site 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 4 | Left Breast Site 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 4 | Left Breast Site 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 4 | Left Breast Site 4 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 4 | Left Breast Site 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 4 | Left Breast Site 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 4 | Left Breast Site 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 5 | Left Breast Site 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 5 | Left Breast Site 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 5 | Left Breast Site 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 5 | Left Breast Site 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 5 | Left Breast Site 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 5 | Left Breast Site 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 5 | Left Breast Site 5 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 5 | Left Breast Site 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 5 | Left Breast Site 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 5 | Left Breast Site 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Lt Breast Site 5 | Left Breast Site 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 5 | Left Breast Site 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 5 | Left Breast Site 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 5 | Left Breast Site 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 5 | Left Breast Site 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 5 | Left Breast Site 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 5 | Left Breast Site 5 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 5 | Left Breast Site 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 5 | Left Breast Site 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 5 | Left Breast Site 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Lt Breast Site 5 | Left Breast Site 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 5 | Left Breast Site 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 5 | Left Breast Site 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 5 | Left Breast Site 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 5 | Left Breast Site 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 5 | Left Breast Site 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 5 | Left Breast Site 5 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 5 | Left Breast Site 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 5 | Left Breast Site 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Lt Breast Site 5 | Left Breast Site 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 1 | Right Breast Site 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 1 | Right Breast Site 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 1 | Right Breast Site 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 1 | Right Breast Site 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 1 | Right Breast Site 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 1 | Right Breast Site 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 1 | Right Breast Site 1 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 1 | Right Breast Site 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 1 | Right Breast Site 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 1 | Right Breast Site 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 1 | Right Breast Site 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 1 | Right Breast Site 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 1 | Right Breast Site 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 1 | Right Breast Site 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 1 | Right Breast Site 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 1 | Right Breast Site 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 1 | Right Breast Site 1 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 1 | Right Breast Site 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 1 | Right Breast Site 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 1 | Right Breast Site 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 1 | Right Breast Site 1 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 1 | Right Breast Site 1 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 1 | Right Breast Site 1 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 1 | Right Breast Site 1 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 1 | Right Breast Site 1 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 1 | Right Breast Site 1 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 1 | Right Breast Site 1 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 1 | Right Breast Site 1 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 1 | Right Breast Site 1 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 1 | Right Breast Site 1 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 2 | Right Breast Site 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 2 | Right Breast Site 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 2 | Right Breast Site 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 2 | Right Breast Site 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 2 | Right Breast Site 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 2 | Right Breast Site 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 2 | Right Breast Site 2 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 2 | Right Breast Site 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 2 | Right Breast Site 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 2 | Right Breast Site 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 2 | Right Breast Site 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 2 | Right Breast Site 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 2 | Right Breast Site 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 2 | Right Breast Site 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 2 | Right Breast Site 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 2 | Right Breast Site 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 2 | Right Breast Site 2 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 2 | Right Breast Site 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 2 | Right Breast Site 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 2 | Right Breast Site 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 2 | Right Breast Site 2 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 2 | Right Breast Site 2 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 2 | Right Breast Site 2 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 2 | Right Breast Site 2 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 2 | Right Breast Site 2 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 2 | Right Breast Site 2 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 2 | Right Breast Site 2 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 2 | Right Breast Site 2 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 2 | Right Breast Site 2 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 2 | Right Breast Site 2 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 3 | Right Breast Site 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 3 | Right Breast Site 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 3 | Right Breast Site 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 3 | Right Breast Site 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 3 | Right Breast Site 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 3 | Right Breast Site 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 3 | Right Breast Site 3 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 3 | Right Breast Site 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 3 | Right Breast Site 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 3 | Right Breast Site 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 3 | Right Breast Site 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 3 | Right Breast Site 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 3 | Right Breast Site 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 3 | Right Breast Site 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 3 | Right Breast Site 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 3 | Right Breast Site 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 3 | Right Breast Site 3 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 3 | Right Breast Site 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 3 | Right Breast Site 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 3 | Right Breast Site 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 3 | Right Breast Site 3 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 3 | Right Breast Site 3 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 3 | Right Breast Site 3 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 3 | Right Breast Site 3 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 3 | Right Breast Site 3 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 3 | Right Breast Site 3 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 3 | Right Breast Site 3 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 3 | Right Breast Site 3 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 3 | Right Breast Site 3 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 3 | Right Breast Site 3 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 4 | Right Breast Site 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 4 | Right Breast Site 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 4 | Right Breast Site 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 4 | Right Breast Site 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 4 | Right Breast Site 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 4 | Right Breast Site 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 4 | Right Breast Site 4 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 4 | Right Breast Site 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 4 | Right Breast Site 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 4 | Right Breast Site 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 4 | Right Breast Site 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 4 | Right Breast Site 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 4 | Right Breast Site 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 4 | Right Breast Site 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 4 | Right Breast Site 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 4 | Right Breast Site 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 4 | Right Breast Site 4 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 4 | Right Breast Site 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 4 | Right Breast Site 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 4 | Right Breast Site 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 4 | Right Breast Site 4 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 4 | Right Breast Site 4 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 4 | Right Breast Site 4 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 4 | Right Breast Site 4 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 4 | Right Breast Site 4 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 4 | Right Breast Site 4 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 4 | Right Breast Site 4 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 4 | Right Breast Site 4 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 4 | Right Breast Site 4 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 4 | Right Breast Site 4 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 5 | Right Breast Site 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;IQRVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 5 | Right Breast Site 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;IQRMdVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 5 | Right Breast Site 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;StdDevVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 5 | Right Breast Site 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Medians;99SIEMENS;MeanVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 5 | Right Breast Site 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Medians;99SIEMENS;MedianVsMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 5 | Right Breast Site 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Medians;99SIEMENS;IQREMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 5 | Right Breast Site 5 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Medians;99SIEMENS;IQRMdEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 5 | Right Breast Site 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Medians;99SIEMENS;StdDevEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 5 | Right Breast Site 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Medians;99SIEMENS;MeanEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 5 | Right Breast Site 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Medians;99SIEMENS;MedianEMedians | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Rt Breast Site 5 | Right Breast Site 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Means;99SIEMENS;IQRVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 5 | Right Breast Site 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Means;99SIEMENS;IQRMdVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 5 | Right Breast Site 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Means;99SIEMENS;StdDevVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 5 | Right Breast Site 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Means;99SIEMENS;MeanVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 5 | Right Breast Site 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Means;99SIEMENS;MedianVsMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 5 | Right Breast Site 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Means;99SIEMENS;IQREMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 5 | Right Breast Site 5 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Means;99SIEMENS;IQRMdEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 5 | Right Breast Site 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Means;99SIEMENS;StdDevEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 5 | Right Breast Site 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Means;99SIEMENS;MeanEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 5 | Right Breast Site 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Means;99SIEMENS;MedianEMeans | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Means;99SIEMENS;SWEMeans
Rt Breast Site 5 | Right Breast Site 5 Shear Wave Velocity Interquartile Range | IQR | m/s | 2 | 1 | Interquartile Range of Shear Wave Velocity Max;99SIEMENS;IQRVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 5 | Right Breast Site 5 Shear Wave Velocity IQR/Median | IQR/Median | unitless | 2 | 1 | IQR/Median Ratio of Shear Wave Velocity Max;99SIEMENS;IQRMdVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 5 | Right Breast Site 5 Shear Wave Velocity Standard Deviation | Std Dev | m/s | 2 | 1 | Standard Deviation of Shear Wave Velocity Max;99SIEMENS;StdDevVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 5 | Right Breast Site 5 Shear Wave Velocity Mean | Mean | m/s | 2 | 1 | Mean of Shear Wave Velocity Max;99SIEMENS;MeanVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 5 | Right Breast Site 5 Shear Wave Velocity Median | Median | m/s | 2 | 1 | Median of Shear Wave Velocity Max;99SIEMENS;MedianVsMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 5 | Right Breast Site 5 Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Interquartile Range of Elasticity Max;99SIEMENS;IQREMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 5 | Right Breast Site 5 Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | IQR/Median Ratio of Elasticity Max;99SIEMENS;IQRMdEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 5 | Right Breast Site 5 Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Standard Deviation of Elasticity Max;99SIEMENS;StdDevEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 5 | Right Breast Site 5 Elasticity Mean | Mean | kPa | 1 | 1 | Mean of Elasticity Max;99SIEMENS;MeanEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Rt Breast Site 5 | Right Breast Site 5 Elasticity Median | Median | kPa | 1 | 1 | Median of Elasticity Max;99SIEMENS;MedianEMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography with Max Velocities;99SIEMENS;SWEMax
Overall Statistics: Breast | Overall Breast Elasticity Interquartile Range | IQR | kPa | 1 | 1 | Overall Interquartile Range of Elasticity Medians;99SIEMENS;OIQREMedians | NUM | Breast;SRT;T-04000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Breast | Overall Breast Elasticity IQR/Median | IQR/Median | unitless | 1 | 1 | Overall IQR/Median Ratio of Elasticity Medians;99SIEMENS;OIQRMdEMedians | NUM | Breast;SRT;T-04000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Breast | Overall Breast Elasticity Standard Deviation | Std Dev | kPa | 1 | 1 | Overall Standard Deviation of Elasticity Medians;99SIEMENS;OStdDevEMedians | NUM | Breast;SRT;T-04000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Breast | Overall Breast Elasticity Mean | Mean | kPa | 1 | 1 | Overall Mean of Elasticity Medians;99SIEMENS;OMeanEMedians | NUM | Breast;SRT;T-04000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Breast | Overall Breast Elasticity Median | Median | kPa | 1 | 1 | Overall Median of Elasticity Medians;99SIEMENS;OMedianEMedians | NUM | Breast;SRT;T-04000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Breast | Overall Breast Interquartile Range of Shear Wave Velocity Medians | IQR | m/s | 2 | 1 | Overall Interquartile Range of Shear Wave Velocity Medians;99SIEMENS;OIQRVsMedians | NUM | Breast;SRT;T-04000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Breast | Overall Breast IQR/Median of Shear Wave Velocity Medians | IQR/Median | unitless | 2 | 1 | Overall IQR/Median Ratio of Shear Wave Velocity Medians;99SIEMENS;OIQRMdVsMedians | NUM | Breast;SRT;T-04000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Breast | Overall Breast Standard Deviation of Shear Wave Velocity Medians | Std Dev | m/s | 2 | 1 | Overall Standard Deviation of Shear Wave Velocity Medians;99SIEMENS;OStdDevVsMedians | NUM | Breast;SRT;T-04000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Breast | Overall Breast Mean of Shear Wave Velocity Medians | Mean | m/s | 2 | 1 | Overall Mean of Shear Wave Velocity Medians;99SIEMENS;OMeanVsMedians | NUM | Breast;SRT;T-04000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Overall Statistics: Breast | Overall Breast Median of Shear Wave Velocity Medians | Median | m/s | 2 | 1 | Overall Median of Shear Wave Velocity Medians;99SIEMENS;OMedianVsMedians | NUM | Breast;SRT;T-04000 |  | Finding Site;SRT;G-C0E3=Breast;SRT;T-04000 |  |  | Measurement Method;SRT;G-C036=Shear Wave Elastography with Medians;99SIEMENS;SWEMedians
Williams | EFW |  | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Summary;DCM;121111 |  | Table of Values;DCM;121424=FWP by GA, Williams 1982 (average);LN;33184-3
Ebbing | Umb A PI(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800 |  | Table of Values;DCM;121424=Umbilical Artery PI by GA Ebbing;99SIEMENS;UmbArtPIEbbing
Ebbing | Rt CP Ratio(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM | Fetal Vascular Ratios;99SIEMENS;FetalVascRatios | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800 |  | Table of Values;DCM;121424=Right Cerebroplacental Ratio by GA Ebbing;99SIEMENS;RtCPRatioEbbing
Ebbing | Lt CP Ratio(GC) |  | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM | Fetal Vascular Ratios;99SIEMENS;FetalVascRatios | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Embryonic Vascular Structure;SRT;T-F6800 |  | Table of Values;DCM;121424=Left Cerebroplacental Ratio by GA Ebbing;99SIEMENS;LtCPRatioEbbing
HC(c) | Head Circumference from BPD and OFD |  | mm | 3 | 1 | Head Circumference;LN;11984-2 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002 |  | Derivation;DCM;121401=Calculated;DCM;121428
AC(c) | Abdominal Circumference from APAD and TAD |  | mm | 3 | 1 | Abdominal Circumference;LN;11979-2 | NUM | Biometry Group;DCM;125005 |  | Fetal Biometry;DCM;125002 |  | Derivation;DCM;121401=Calculated;DCM;121428