## UCUM Units
Coding Scheme Designator | Meaning | Value
UCUM | percent | %
UCUM | beats per minute | bpm
UCUM | cubic centimeter | cc
UCUM | centimeter | cm
UCUM | centimeters per second | cm/s
UCUM | centimeters per square second | cm/s2
UCUM | square centimeter | cm2
UCUM | day | d
UCUM | degree of angle | deg
UCUM | degrees per second | deg/s
UCUM | gram | g
UCUM | grams per cubic centimeter | g/cm3
UCUM | kilogram | kg
UCUM | kilopascal | kPa
UCUM | liters per minute | l/min
UCUM | pound | lb
UCUM | meter | m
UCUM | meters per second | m/s
UCUM | meters per square second | m/s2
UCUM | square meter | m2
UCUM | milliliter | ml
UCUM | milliliters per second | ml/s
UCUM | millimeter | mm
UCUM | millimeters per second | mm/s
UCUM | square millimeter | mm2
UCUM | cubic millimeter | mm3
UCUM | millimeter of mercury column | mm[Hg]
UCUM | millimeters of mercury per second | mm[Hg]/s
UCUM | month | mo
UCUM | millisecond | ms
UCUM | ounce | oz
UCUM | second | s
UCUM | no units | unitless
UCUM | milliliters per minute | ml/min
UCUM | beats per minute | bpm
UCUM | centimeters per second | cm/s
UCUM | centimeter | cm
UCUM | meters per second | m/s
UCUM | meter | m
UCUM | millimeter | mm
UCUM | Wood U | [wood'U]
UCUM | milliliter per centimeters | ml/cm
UCUM | square centimeters per square meter | cm2/m2
UCUM | grams per centimeter | g/cm
UCUM | grams per square centimeter | g/cm2
UCUM | grams per square meter | g/m2
UCUM | liters per minute per square meter | l/min/m2
UCUM | milliliters per square meter | ml/m2
UCUM | liter | l
UCUM | minute | min

## Cardiac SR Mappings
EnglishLabel | Possible Units | Precision | Parent Container | Concept Coding | Value Type | Image View | Finding Site | Measurement Method | Cardiac Cycle Point | Topographical Modifier | Image Mode | Respiratory Cycle Point | Flow Direction | Blood Vessel Finding | Index | AdditionalModifier
Ao Diam d | cm mm | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Aortic Root Diameter;LN;18015-8 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=M mode;SRT;G-0394
AoV Cusp Sep | cm mm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Aortic Valve Cusp Separation;LN;17996-0 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LA Diam s | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium Antero-posterior Systolic Dimension;LN;29469-4 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
MAPSE | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Annular Plane Systolic Excursion;99SIEMENS;MAPSE | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
MV DE Excursion | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | D-E Excursion;LN;59091-9 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
MV DE Slope | cm/s m/s | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | D-E Slope;LN;59127-1 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
MV EF Slope | cm/s m/s | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-F Slope by M-Mode;LN;18040-6 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
MV EPSS | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve EPSS, E wave;LN;18036-4 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
CA Amp | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Amplitude of C point to A point;99SIEMENS;MVCAAmp | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
CE Amp | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Amplitude of C point to E point;99SIEMENS;MVCEAmp | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
RVAW d | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Anterior Wall Diastolic Thickness;LN;18153-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
RVAW s | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Anterior Wall Systolic Thickness;LN;18157-8 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
RVID d | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Internal Diastolic Dimension;LN;20304-2 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
RVID s | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Internal Systolic Dimension;LN;20305-9 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
RV Inflow Vp | cm/s m/s | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Velocity of Flow Propagation;LN;59115-6 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
RV PEP (M-mode) | ms | 0 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Pre-Ejection Period;LN;59085-1 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
RV ET (M-mode) | ms | 0 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Ejection Time;DCM;122213 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
IVS s | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Interventricular Septum Systolic Thickness;LN;18158-6 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LVID s | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Internal Systolic Dimension;LN;29438-9 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LVPW s | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Posterior Wall Systolic Thickness;LN;18156-0 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
IVS d | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Interventricular Septum Diastolic Thickness;LN;18154-5 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LVID d | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Internal End Diastolic Dimension;LN;29436-3 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LVPW d | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Posterior Wall Diastolic Thickness;LN;18152-9 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LVOT Diam s (M-mode) | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV Inflow Vp | cm/s m/s | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Velocity of Flow Propagation;LN;59115-6 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV PEP (M-mode) | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Pre-Ejection Period;LN;59085-1 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV ET (M-mode) | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Ejection Time;DCM;122211 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV Major s | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Systolic Major Axis;LN;18076-0 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
PE d | cm mm | 2 | Finding Site;SRT;G-C0E3=Pericardial Cavity;SRT;T-39050 | Diameter;SRT;M-02550 | NUM |  | Finding Site;SRT;G-C0E3=Pericardial effusion;SRT;D3-90008 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=M mode;SRT;G-0394
IVC Diam (M-mode) | cm mm | 2 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | Inferior Vena Cava Diameter;LN;18006-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
IVC Collapsed Diam (M-mode) | cm mm | 2 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | Inferior Vena Cava Diameter;LN;18006-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394 |  |  | Blood vessel finding;SRT;F-3A000=Collapsed;SRT;G-A282
RA Pressure | mmHg | 0 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Right Atrium Systolic Pressure;LN;18070-3 | NUM
TAPSE | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Annular Plane Systolic Excursion;99SIEMENS;TAPSE | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
TV DE Excursion | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | D-E Excursion;LN;59091-9 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
TV DE Slope | cm/s m/s | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | D-E Slope;LN;59127-1 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
TV EF Slope | cm/s m/s | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | E-F Slope;LN;59128-9 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
Ao Root d | cm mm | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Aortic Root Diameter;LN;18015-8 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Sinus d | cm mm | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Aortic Root Diameter;LN;18015-8 | NUM |  | Finding Site;SRT;G-C0E3=Structure Sinus of Valsalva;SRT;T-42200 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao STJ d | cm mm | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Aortic Root Diameter;LN;18015-8 | NUM |  | Finding Site;SRT;G-C0E3=Aortic Sinotubular Junction;SRT;T-42102 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Asc d | cm mm | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Ascending Aortic Diameter;LN;18012-5 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Asc Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Peak Systolic Velocity;LN;11726-7 | NUM |  | Finding Site;SRT;G-C0E3=Ascending Aorta;SRT;T-42100
Ao Asc VTI | cm | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Ascending Aorta;SRT;T-42100
Ao Asc Mean Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Mean Gradient;LN;20256-4 | NUM |  | Finding Site;SRT;G-C0E3=Ascending Aorta;SRT;T-42100
Ao Arch d | cm mm | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Aortic Arch Diameter;LN;18011-7 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Arch Trans Prox | cm mm | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Diameter;SRT;M-02550 | NUM |  | Finding Site;SRT;G-C0E3=Transverse Aortic Arch;SRT;T-42304 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Arch Distal | cm mm | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Aortic Arch Diameter;LN;18011-7 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Arch Trans Distal | cm mm | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Diameter;SRT;M-02550 | NUM |  | Finding Site;SRT;G-C0E3=Transverse Aortic Arch;SRT;T-42304 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Isthmus | cm mm | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Aortic Isthmus Diameter;LN;18014-1 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Desc | cm mm | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Descending Aortic Diameter;LN;18013-3 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Desc Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Peak Systolic Velocity;LN;11726-7 | NUM |  | Finding Site;SRT;G-C0E3=Descending Aorta;SRT;T-42400
Ao Thoracic | cm mm | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Diameter;SRT;M-02550 | NUM |  | Finding Site;SRT;G-C0E3=Thoracic Aorta;SRT;T-42070 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Abd | cm mm | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Diameter;SRT;M-02550 | NUM |  | Finding Site;SRT;G-C0E3=Abdominal Aorta;SRT;T-42500 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Coarct Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Diameter;SRT;M-02550 | NUM |  | Finding Site;SRT;G-C0E3=Coarctation of the Aorta;SRT;D4-32014 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Coarct Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Thoracic Aorta Coarctation Systolic PeakVelocity;LN;29460-3 | NUM |  | Finding Site;SRT;G-C0E3=Coarctation of the Aorta;SRT;D4-32014
Left Main Diam | mm | 1 | Finding Site;SRT;G-C0E3=Coronary Artery;SRT;T-43000 | Diameter;SRT;M-02550 | NUM |  | Finding Site;SRT;G-C0E3=Left Main Coronary Artery;SRT;T-43107 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LAD Diam | mm | 1 | Finding Site;SRT;G-C0E3=Coronary Artery;SRT;T-43000 | Diameter;SRT;M-02550 | NUM |  | Finding Site;SRT;G-C0E3=Left Anterior Descending Coronary Artery;SRT;T-43110 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Circumflex Diam | mm | 1 | Finding Site;SRT;G-C0E3=Coronary Artery;SRT;T-43000 | Diameter;SRT;M-02550 | NUM |  | Finding Site;SRT;G-C0E3=Circumflex Coronary Artery;SRT;T-43120 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RCA Diam | mm | 1 | Finding Site;SRT;G-C0E3=Coronary Artery;SRT;T-43000 | Diameter;SRT;M-02550 | NUM |  | Finding Site;SRT;G-C0E3=Right Coronary Artery;SRT;T-43203 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
AoV Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Peak Systolic Velocity;LN;11726-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AoV AT | ms | 0 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Acceleration Time;LN;20168-1 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AoV Accel Slope | m/s² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Acceleration Slope;LN;20167-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AoV ET | ms | 0 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Aortic Valve Ejection Time;LN;18041-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AoV Area Planimetry | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Planimetry;DCM;125220 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVOT Diam s | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 |  | Cardiac Cycle Point;SRT;R-4089A=End Systole;SRT;R-FAB5B |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVOT VTI | cm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT Vmean | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Blood velocity, mean;DCM;122205 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Peak Systolic Velocity;LN;11726-7 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT Mean Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Mean Gradient;LN;20256-4 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT AT | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Acceleration Time;LN;20168-1 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
AoV VTI | cm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Velocity Time Integral;LN;20354-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AoV Vmean | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Blood velocity, mean;DCM;122205 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AoV Mean Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Mean Gradient;LN;20256-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AR Slope | m/s² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Deceleration Slope;LN;20216-8 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR PHT | ms | 0 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Pressure Half-Time;LN;20280-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR DT | ms | 0 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Deceleration Time;LN;20217-6 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR ED Velocity | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | End Diastolic Velocity;LN;11653-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR ED Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Pressure Gradient;SRT;F-023F7 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011 |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR VTI | cm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Velocity Time Integral;LN;20354-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Peak Velocity;LN;20351-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR Vmean | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Blood velocity, mean;DCM;122205 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR Mean Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Mean Gradient;LN;20256-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR Duration | ms | 0 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Duration;SRT;G-7290 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR VC Width | cm mm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Width;SRT;G-A220 | NUM |  | Finding Site;SRT;G-C0E3=Vena Contracta;SRT;R-421AA |  |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR Jet Area | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Jet Area;LN;59123-0 | NUM | [5 empty] | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
LVOT Area d | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 |  | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
AR Jet Ht | cm mm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | AR Jet Height;99SIEMENS;ARJetHeight | NUM | [5 empty] | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2
LVOT Ht d | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Height;DCM;121207 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 |  | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
AoV Site 2 Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
AoV Site 2 VTI | cm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR PISA Radius | cm mm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Flow Radius;LN;59102-4 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR Aliasing Velocity | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Alias Velocity;LN;59130-5 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AoV Annulus s | cm mm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Aortic Valve Ring;SRT;T-35410 |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
AoV AT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Acceleration Time;LN;20168-1 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV Accel Slope (Post) | m/s² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Acceleration Slope;LN;20167-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV ET (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Aortic Valve Ejection Time;LN;18041-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV Area Planimetry (Post) | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Planimetry;DCM;125220 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LVOT Diam s (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 |  | Cardiac Cycle Point;SRT;R-4089A=End Systole;SRT;R-FAB5B |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV VTI (Post) | cm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Velocity Time Integral;LN;20354-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV Vmax (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Peak Systolic Velocity;LN;11726-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV Vmean (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Blood velocity, mean;DCM;122205 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV Mean Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Mean Gradient;LN;20256-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LVOT Vmax (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Peak Systolic Velocity;LN;11726-7 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LVOT Peak Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Peak Gradient;LN;20247-3 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR Slope (Post) | m/s² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Deceleration Slope;LN;20216-8 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR PHT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Pressure Half-Time;LN;20280-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR DT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Deceleration Time;LN;20217-6 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR ED Velocity (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | End Diastolic Velocity;LN;11653-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR ED Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Pressure Gradient;SRT;F-023F7 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011 |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR VTI (Post) | cm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Velocity Time Integral;LN;20354-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR Vmax (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Peak Velocity;LN;20351-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR Vmean (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Blood velocity, mean;DCM;122205 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR Peak Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR Mean Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Mean Gradient;LN;20256-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR Duration (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Duration;SRT;G-7290 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR VC Width (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Width;SRT;G-A220 | NUM |  | Finding Site;SRT;G-C0E3=Vena Contracta;SRT;R-421AA |  |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR Jet Area (Post) | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Jet Area;LN;59123-0 | NUM | [5 empty] | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LVOT Area d (Post) | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 |  | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR Jet Ht (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | AR Jet Height;99SIEMENS;ARJetHeight | NUM | [5 empty] | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LVOT Ht d (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Height;DCM;121207 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 |  | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV Site 2 Diam (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV Site 2 VTI (Post) | cm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR PISA Radius (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Flow Radius;LN;59102-4 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR Aliasing Velocity (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Alias Velocity;LN;59130-5 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV Annulus s (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Aortic Valve Ring;SRT;T-35410 |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
ASD Diam | mm | 1 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Diameter;SRT;M-02550 | NUM |  | Finding Site;SRT;G-C0E3=Atrial Septal Defect;SRT;D4-31220 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
ASD Major | mm | 1 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Major Axis;SRT;G-A193 | NUM |  | Finding Site;SRT;G-C0E3=Atrial Septal Defect;SRT;D4-31220 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
ASD Minor | mm | 1 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Minor Axis;SRT;G-A194 | NUM |  | Finding Site;SRT;G-C0E3=Atrial Septal Defect;SRT;D4-31220 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
ASD Sup-Inf | mm | 1 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Superior-Inferior Diameter;99SIEMENS;SupInfDiam | NUM |  | Finding Site;SRT;G-C0E3=Atrial Septal Defect;SRT;D4-31220 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
ASD Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Peak Systolic Velocity;LN;11726-7 | NUM |  | Finding Site;SRT;G-C0E3=Atrial Septal Defect;SRT;D4-31220
PDA Diam | mm | 1 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Diameter;SRT;M-02550 | NUM |  | Finding Site;SRT;G-C0E3=Patent Ductus Arteriosus;SRT;D4-32012 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
PDA Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Peak Systolic Velocity;LN;11726-7 | NUM |  | Finding Site;SRT;G-C0E3=Patent Ductus Arteriosus;SRT;D4-32012
VSD Diam | mm | 1 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Diameter;SRT;M-02550 | NUM |  | Finding Site;SRT;G-C0E3=Ventricular Septal Defect;SRT;D4-31150 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
VSD Major | mm | 1 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Major Axis;SRT;G-A193 | NUM |  | Finding Site;SRT;G-C0E3=Ventricular Septal Defect;SRT;D4-31150 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
VSD Minor | mm | 1 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Minor Axis;SRT;G-A194 | NUM |  | Finding Site;SRT;G-C0E3=Ventricular Septal Defect;SRT;D4-31150 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
VSD Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Peak Systolic Velocity;LN;11726-7 | NUM |  | Finding Site;SRT;G-C0E3=Ventricular Septal Defect;SRT;D4-31150
Pulm Side Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Pulmonary Side;99SIEMENS;PulmonarySide
Pulm Side VTI | cm | 2 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Pulmonary Side;99SIEMENS;PulmonarySide | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
Sys Side Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Systemic Side;99SIEMENS;SystemicSide
Sys Side VTI | cm | 2 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Systemic Side;99SIEMENS;SystemicSide | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
Qp HR | bpm | 0 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Heart rate;LN;8867-4 | NUM |  | Finding Site;SRT;G-C0E3=Pulmonary Side;99SIEMENS;PulmonarySide
Qs HR | bpm | 0 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Heart rate;LN;8867-4 | NUM |  | Finding Site;SRT;G-C0E3=Systemic Side;99SIEMENS;SystemicSide
LA A-P Diam s PLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium Antero-posterior Systolic Dimension;LN;29469-4 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Major d A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Major Axis;SRT;G-A193 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Major s A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Major Axis;SRT;G-A193 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Minor d A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Minor Axis;SRT;G-A194 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Minor s A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Minor Axis;SRT;G-A194 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Major d A2C | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Major Axis;SRT;G-A193 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Major s A2C | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Major Axis;SRT;G-A193 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Area d A4C | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium Area A4C view;LN;17977-0 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol d A4C MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Area s A4C | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium Area A4C view;LN;17977-0 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol s A4C MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium Systolic Volume;SRT;G-0383 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Area d A2C | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Area;SRT;G-A166 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol d A2C MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Area s A2C | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Area;SRT;G-A166 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol s A2C MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium Systolic Volume;SRT;G-0383 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV E Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-Wave Peak Velocity;LN;18037-2 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV A Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve A-Wave Peak Velocity;LN;17978-8 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV E Vmax (Valsalva) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-Wave Peak Velocity;LN;18037-2 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Valsalva maneuver;SRT;R-40928 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV A Vmax (Valsalva) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve A-Wave Peak Velocity;LN;17978-8 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Valsalva maneuver;SRT;R-40928 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
LV IVRT | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Isovolumic Relaxation Time;LN;18071-1 | NUM
MV AT | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Acceleration Time;LN;20168-1 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Accel Slope | m/s² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Acceleration Slope;LN;20167-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV DT | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-Wave Deceleration Time;SRT;G-0384 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV PHT | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Pressure Half-Time;LN;20280-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV E Duration | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-Wave Duration;99SIEMENS;CA_MV_DURe | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV E Vmax, Annulus | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-Wave Peak Velocity;LN;18037-2 | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV A Vmax, Annulus | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve A-Wave Peak Velocity;LN;17978-8 | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV DT, Annulus | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Deceleration Time;LN;20217-6 | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV VTI, Annulus | cm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Vmax, Annulus | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Velocity;LN;20351-3 | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Vmean, Annulus | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Blood velocity, mean;DCM;122205 | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV DFT, Annulus | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Diastolic Filling Period;DFPm;SRT;R-0032C | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV E VTI, Annulus | cm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-Wave Velocity Time Integral;99SIEMENS;VTI_MVEWave | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV E Duration, Annulus | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-Wave Duration;99SIEMENS;CA_MV_DURe | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV A VTI, Annulus | cm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve A-Wave Velocity Time Integral;99SIEMENS;VTI_MVAWave | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV A Duration, Annulus | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve A-Wave Duration;SRT;G-0385 | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV e' Vmax (M) | m/s cm/s | 3 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Peak Early Diastolic Tissue Velocity;SRT;G-037A | NUM |  | Finding Site;SRT;G-C0E3=Medial Mitral Annulus;SRT;G-0391 |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV e' AT (M) | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Mitral Valve Acceleration time of Ea;99SIEMENS;MVEaAT | NUM |  | Finding Site;SRT;G-C0E3=Medial Mitral Annulus;SRT;G-0391 |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV e' DT (M) | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Mitral Valve Decceleration time of Ea;99SIEMENS;MVEaDT | NUM |  | Finding Site;SRT;G-C0E3=Medial Mitral Annulus;SRT;G-0391 |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV a' Vmax (M) | m/s cm/s | 3 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | LV Peak Diastolic Tissue Velocity During Atrial Systole;SRT;G-037C | NUM |  | Finding Site;SRT;G-C0E3=Medial Mitral Annulus;SRT;G-0391 |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV s' Vmax (M) | m/s cm/s | 3 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Peak Systolic Tissue Velocity;SRT;G-037D | NUM |  | Finding Site;SRT;G-C0E3=Medial Mitral Annulus;SRT;G-0391 |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV e' Vmax (L) | m/s cm/s | 3 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Peak Early Diastolic Tissue Velocity;SRT;G-037A | NUM |  | Finding Site;SRT;G-C0E3=Lateral Mitral Annulus;SRT;G-0392 |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV e' AT (L) | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Mitral Valve Acceleration time of Ea;99SIEMENS;MVEaAT | NUM |  | Finding Site;SRT;G-C0E3=Lateral Mitral Annulus;SRT;G-0392 |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV e' DT (L) | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Mitral Valve Decceleration time of Ea;99SIEMENS;MVEaDT | NUM |  | Finding Site;SRT;G-C0E3=Lateral Mitral Annulus;SRT;G-0392 |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV a' Vmax (L) | m/s cm/s | 3 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | LV Peak Diastolic Tissue Velocity During Atrial Systole;SRT;G-037C | NUM |  | Finding Site;SRT;G-C0E3=Lateral Mitral Annulus;SRT;G-0392 |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV s' Vmax (L) | m/s cm/s | 3 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Peak Systolic Tissue Velocity;SRT;G-037D | NUM |  | Finding Site;SRT;G-C0E3=Lateral Mitral Annulus;SRT;G-0392 |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
P Vein S Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Systolic Peak Velocity;LN;29450-4 | NUM
P Vein S VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein S-Wave Velocity Time Integral;SRT;G-038C | NUM
P Vein D Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Diastolic Peak Velocity;LN;29451-2 | NUM
P Vein D VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein D-Wave Velocity Time Integral;SRT;G-038D | NUM
P Vein A Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Atrial Contraction Reversal Peak Velocity;LN;29453-8 | NUM
P Vein A Duration | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein A-Wave Duration;SRT;G-038B | NUM
IVS d PLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Interventricular Septum Diastolic Thickness;LN;18154-5 | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
IVS s PLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Interventricular Septum Systolic Thickness;LN;18158-6 | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
IVS Bulge d PLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Interventricular Septum Diastolic Thickness;LN;18154-5 | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 | Finding Site;SRT;G-C0E3=Bulge;99SIEMENS;BULGE |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
IVS Bulge s PLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Interventricular Septum Systolic Thickness;LN;18158-6 | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 | Finding Site;SRT;G-C0E3=Bulge;99SIEMENS;BULGE |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVID d PLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Internal End Diastolic Dimension;LN;29436-3 | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVID s PLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Internal Systolic Dimension;LN;29438-9 | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVPW d PLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Posterior Wall Diastolic Thickness;LN;18152-9 | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVPW s PLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Posterior Wall Systolic Thickness;LN;18156-0 | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
IVS d PSAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Interventricular Septum Diastolic Thickness;LN;18154-5 | NUM | Image View;DCM;111031=Parasternal short axis;SRT;G-0397 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
IVS s PSAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Interventricular Septum Systolic Thickness;LN;18158-6 | NUM | Image View;DCM;111031=Parasternal short axis;SRT;G-0397 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVID d PSAX chord | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Internal End Diastolic Dimension;LN;29436-3 | NUM | Image View;DCM;111031=Parasternal short axis at the level of the mitral chords;SRT;G-0399 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVID s PSAX chord | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Internal Systolic Dimension;LN;29438-9 | NUM | Image View;DCM;111031=Parasternal short axis at the level of the mitral chords;SRT;G-0399 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVPW d PSAX chord | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Posterior Wall Diastolic Thickness;LN;18152-9 | NUM | Image View;DCM;111031=Parasternal short axis at the level of the mitral chords;SRT;G-0399 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVPW s PSAX chord | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Posterior Wall Systolic Thickness;LN;18156-0 | NUM | Image View;DCM;111031=Parasternal short axis at the level of the mitral chords;SRT;G-0399 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVID d A-P PSAX pap | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Internal Dimension A-P;99SIEMENS;LVID_AP | NUM | Image View;DCM;111031=Parasternal short axis at the Papillary Muscle level;SRT;G-039B |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVID s A-P PSAX pap | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Internal Dimension A-P;99SIEMENS;LVID_AP | NUM | Image View;DCM;111031=Parasternal short axis at the Papillary Muscle level;SRT;G-039B |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVID d M-L PSAX pap | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Internal Dimension M-L;99SIEMENS;LVID_ML | NUM | Image View;DCM;111031=Parasternal short axis at the Papillary Muscle level;SRT;G-039B |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVID s M-L PSAX pap | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Internal Dimension M-L;99SIEMENS;LVID_ML | NUM | Image View;DCM;111031=Parasternal short axis at the Papillary Muscle level;SRT;G-039B |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVPW d PSAX pap | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Posterior Wall Diastolic Thickness;LN;18152-9 | NUM | Image View;DCM;111031=Parasternal short axis at the Papillary Muscle level;SRT;G-039B |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVPW s PSAX pap | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Posterior Wall Systolic Thickness;LN;18156-0 | NUM | Image View;DCM;111031=Parasternal short axis at the Papillary Muscle level;SRT;G-039B |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Major d A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Diastolic Major axis;LN;18077-8 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Major s A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Systolic Major Axis;LN;18076-0 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Minor d A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Minor Axis;SRT;G-A194 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Minor s A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Minor Axis;SRT;G-A194 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Major d A2C | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Diastolic Major axis;LN;18077-8 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Major s A2C | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Systolic Major Axis;LN;18076-0 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Minor d A2C | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Minor Axis;SRT;G-A194 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Minor s A2C | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Minor Axis;SRT;G-A194 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Major s SCLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Systolic Major Axis;LN;18076-0 | NUM | Image View;DCM;111031=Subcostal long axis;SRT;G-039E |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV dP/dt | mmHg/s | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Regurgitation dP/dt derived from Mitral Reg. Velocity;LN;18035-6 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
LV PEP | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Pre-Ejection Period;LN;59085-1 | NUM
LV AT | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Acceleration Time;LN;20168-1 | NUM
LV ET | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Ejection Time;DCM;122211 | NUM
LV IVCT | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Isovolumic Contraction Time;LN;59084-4 | NUM
MV C-O Duration | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve Closure to Opening Time;SRT;G-0387 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
LV Area d A4C | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Diastolic Area;SRT;G-0375 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Semi-Major d | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Semi-major Axis Diastolic Dimension;SRT;G-0377 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Tr Semi-Major d | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Truncated Semi-major Axis Diastolic Dimension;SRT;G-0378 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d A4C MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Area s A4C | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Systolic Area;SRT;G-0374 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol s A4C MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Area d A2C | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Diastolic Area;SRT;G-0375 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d A2C MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Area s A2C | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Systolic Area;SRT;G-0374 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol s A2C MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Epi Area d PSAX pap | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Epicardial Diastolic Area, psax pap view;SRT;G-0379 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Area d PSAX pap | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Diastolic Area;SRT;G-0375 | NUM | Image View;DCM;111031=Parasternal short axis at the Papillary Muscle level;SRT;G-039B |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV Area Planimetry | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Planimetry;DCM;125220 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV Decel Slope | m/s² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Deceleration Slope;LN;20216-8 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Area (PHT) | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Area by Pressure Half-Time;DCM;125210 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV VTI | cm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Velocity Time Integral;LN;20354-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Velocity;LN;20351-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Vmean | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Blood velocity, mean;DCM;122205 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Mean Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mean Gradient;LN;20256-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV DFT | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Diastolic Filling Period;DFPm;SRT;R-0032C | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV PISA Radius | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Flow Radius;LN;59102-4 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Aliasing Velocity | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Alias Velocity;LN;59130-5 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Funnel Angle | ° | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Funnel Angle;99SIEMENS;FunnelAngle | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MR Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Velocity;LN;20351-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR VTI | cm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Velocity Time Integral;LN;20354-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR Vmean | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Blood velocity, mean;DCM;122205 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR Mean Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mean Gradient;LN;20256-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR AT | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Acceleration Time;LN;20168-1 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR Duration | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Duration;SRT;G-7290 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR VC Width | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Width;SRT;G-A220 | NUM |  | Finding Site;SRT;G-C0E3=Vena Contracta;SRT;R-421AA |  |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MV Site 2 Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV Site 2 VTI | cm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR PISA Radius | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Flow Radius;LN;59102-4 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR Aliasing Velocity | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Alias Velocity;LN;59130-5 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MV Annulus d | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM
MV Annulus s | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV Annulus d A-P | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Anterior-Posterior Diameter;99SIEMENS;APDiameter | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV Annulus s A-P | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Anterior-Posterior Diameter;99SIEMENS;APDiameter | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV Annulus d M-L | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Medial-Lateral Diameter;99SIEMENS;MLDiameter | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV Annulus s M-L | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Medial-Lateral Diameter;99SIEMENS;MLDiameter | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV E Vmax (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-Wave Peak Velocity;LN;18037-2 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV A Vmax (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve A-Wave Peak Velocity;LN;17978-8 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV E Vmax (Valsalva) (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-Wave Peak Velocity;LN;18037-2 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Valsalva maneuver;SRT;R-40928 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV A Vmax (Valsalva) (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve A-Wave Peak Velocity;LN;17978-8 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Valsalva maneuver;SRT;R-40928 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LV IVRT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Isovolumic Relaxation Time;LN;18071-1 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV AT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Acceleration Time;LN;20168-1 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Accel Slope (Post) | m/s² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Acceleration Slope;LN;20167-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV E DT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-Wave Deceleration Time;SRT;G-0384 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV PHT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Pressure Half-Time;LN;20280-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV E Duration (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-Wave Duration;99SIEMENS;CA_MV_DURe | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Area Planimetry (Post) | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Planimetry;DCM;125220 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Decel Slope (Post) | m/s² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Deceleration Slope;LN;20216-8 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Area (PHT) (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Area by Pressure Half-Time;DCM;125210 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV VTI (Post) | cm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Velocity Time Integral;LN;20354-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Vmax (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Velocity;LN;20351-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Vmean (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Blood velocity, mean;DCM;122205 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Mean Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mean Gradient;LN;20256-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV DFT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Diastolic Filling Period;DFPm;SRT;R-0032C | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV PISA Radius (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Flow Radius;LN;59102-4 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Aliasing Velocity (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Alias Velocity;LN;59130-5 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Funnel Angle (Post) | ° | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Funnel Angle;99SIEMENS;FunnelAngle | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR Vmax (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Velocity;LN;20351-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR Peak Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR VTI (Post) | cm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Velocity Time Integral;LN;20354-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR Vmean (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Blood velocity, mean;DCM;122205 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR Mean Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mean Gradient;LN;20256-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR AT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Acceleration Time;LN;20168-1 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR Duration (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Duration;SRT;G-7290 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LV dP/dt (Post) | mmHg/s | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Regurgitation dP/dt derived from Mitral Reg. Velocity;LN;18035-6 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR VC Width (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Width;SRT;G-A220 | NUM |  | Finding Site;SRT;G-C0E3=Vena Contracta;SRT;R-421AA |  |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Site 2 Diam (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Site 2 VTI (Post) | cm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR PISA Radius (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Flow Radius;LN;59102-4 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR Aliasing Velocity (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Alias Velocity;LN;59130-5 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Annulus d (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV VTI, Annulus (Post) | cm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Vmax, Annulus (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Velocity;LN;20351-3 | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Vmean, Annulus (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Blood velocity, mean;DCM;122205 | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Annulus s (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Annulus d A-P (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Anterior-Posterior Diameter;99SIEMENS;APDiameter | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Annulus s A-P (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Anterior-Posterior Diameter;99SIEMENS;APDiameter | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Annulus d M-L (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Medial-Lateral Diameter;99SIEMENS;MLDiameter | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Annulus s M-L (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Medial-Lateral Diameter;99SIEMENS;MLDiameter | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LVOT VTI (Post) | cm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LVOT Vmean (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Blood velocity, mean;DCM;122205 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LVOT Mean Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Mean Gradient;LN;20256-4 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LVOT AT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Acceleration Time;LN;20168-1 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MPA Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Artery;SRT;T-44000 | Main Pulmonary Artery Diameter;LN;18020-8 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MPA Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Artery;SRT;T-44000 | Peak Velocity;LN;20351-3 | NUM
MPA Peak Grad | mmHg | 2 | Finding site;SCT;363698007=Pulmonary artery structure;SCT;81040000 | Peak Gradient;LN;20247-3 | NUM
MPA VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Artery;SRT;T-44000 | Velocity Time Integral;LN;20354-7 | NUM
MPA Vmean | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Artery;SRT;T-44000 | Blood velocity, mean;DCM;122205 | NUM
MPA Mean Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonary Artery;SRT;T-44000 | Mean Gradient;LN;20256-4 | NUM
LPA Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Artery;SRT;T-44000 | Left Pulmonary Artery Diameter;LN;18019-0 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LPA Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Artery;SRT;T-44000 | Peak Velocity;LN;20351-3 | NUM |  | Finding Site;SRT;G-C0E3=Left Pulmonary Artery;SRT;T-44400
RPA Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Artery;SRT;T-44000 | Right Pulmonary Artery Diameter;LN;18021-6 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RPA Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Artery;SRT;T-44000 | Peak Velocity;LN;20351-3 | NUM |  | Finding Site;SRT;G-C0E3=Right Pulmonary Artery;SRT;T-44200
TR Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Velocity;LN;20351-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
RVOT VTI | cm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Peak Systolic Velocity;LN;11726-7 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT Vmean | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Blood velocity, mean;DCM;122205 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT Mean Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Mean Gradient;LN;20256-4 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT AT | ms | 0 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Acceleration Time;LN;20168-1 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT ET | ms | 0 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Ejection Time;DCM;122213 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
PE d, Ant PLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Pericardial Cavity;SRT;T-39050 | Diameter;SRT;M-02550 | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 | Finding Site;SRT;G-C0E3=Pericardial effusion;SRT;D3-90008 |  |  | Topographical modifier;SRT;G-A1F8=Anterior;SRT;R-404CC | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
PE d, Post PLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Pericardial Cavity;SRT;T-39050 | Diameter;SRT;M-02550 | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 | Finding Site;SRT;G-C0E3=Pericardial effusion;SRT;D3-90008 |  |  | Topographical modifier;SRT;G-A1F8=Posterior;SRT;R-404CE | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
PE d, Lt Lat A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Pericardial Cavity;SRT;T-39050 | Diameter;SRT;M-02550 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C | Finding Site;SRT;G-C0E3=Pericardial effusion;SRT;D3-90008 |  |  | Topographical modifier;SRT;G-A1F8=Left lateral;SRT;G-A101 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
PE d, Rt Lat A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Pericardial Cavity;SRT;T-39050 | Diameter;SRT;M-02550 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C | Finding Site;SRT;G-C0E3=Pericardial effusion;SRT;D3-90008 |  |  | Topographical modifier;SRT;G-A1F8=Right lateral;SRT;G-A100 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
PE d, Apical A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Pericardial Cavity;SRT;T-39050 | Diameter;SRT;M-02550 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C | Finding Site;SRT;G-C0E3=Pericardial effusion;SRT;D3-90008 |  |  | Topographical modifier;SRT;G-A1F8=Apical;SRT;G-A122 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
PE d, Ant SCLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Pericardial Cavity;SRT;T-39050 | Diameter;SRT;M-02550 | NUM | Image View;DCM;111031=Subcostal long axis;SRT;G-039E | Finding Site;SRT;G-C0E3=Pericardial effusion;SRT;D3-90008 |  |  | Topographical modifier;SRT;G-A1F8=Anterior;SRT;R-404CC | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
PE d, Post SCLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Pericardial Cavity;SRT;T-39050 | Diameter;SRT;M-02550 | NUM | Image View;DCM;111031=Subcostal long axis;SRT;G-039E | Finding Site;SRT;G-C0E3=Pericardial effusion;SRT;D3-90008 |  |  | Topographical modifier;SRT;G-A1F8=Posterior;SRT;R-404CE | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RA Diam d, insp | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Diameter;SRT;M-02550 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010
RVID d, insp | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Internal Diastolic Dimension;LN;20304-2 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010
LA Diam d, insp | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Diameter;SRT;M-02550 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010
LVID d, insp | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Internal End Diastolic Dimension;LN;29436-3 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010
MV E Vmax, insp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-Wave Peak Velocity;LN;18037-2 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV A Vmax, insp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve A-Wave Peak Velocity;LN;17978-8 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV DT, insp | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Deceleration Time;LN;20217-6 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
LV IVRT, insp | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Isovolumic Relaxation Time;LN;18071-1 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010
P Vein S Vmax, insp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Systolic Peak Velocity;LN;29450-4 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010
P Vein D Vmax, insp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Diastolic Peak Velocity;LN;29451-2 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010
P Vein A Vmax, insp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Atrial Contraction Reversal Peak Velocity;LN;29453-8 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010
TV E Vmax, insp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve E-Wave Peak Velocity;LN;18031-5 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV A Vmax, insp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve A-Wave Peak Velocity;LN;18030-7 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV DT, insp | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Deceleration Time;LN;20217-6 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
IVC Diam, insp | cm mm | 2 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | Inferior Vena Cava Diameter;LN;18006-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010
HV S Vmax, insp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | Hepatic Vein Systolic Peak Velocity;LN;29471-0 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010
HV D Vmax, insp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | Hepatic Vein Diastolic Peak Velocity;LN;29472-8 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010
HV A Vmax, insp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | Hepatic Vein Atrial Contraction Reversal Peak Velocity;LN;29474-4 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010
RA Diam d, exp | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Diameter;SRT;M-02550 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020
RVID d, exp | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Internal Diastolic Dimension;LN;20304-2 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020
LA Diam d, exp | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Diameter;SRT;M-02550 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020
LVID d, exp | cm mm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Internal End Diastolic Dimension;LN;29436-3 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020
MV E Vmax, exp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-Wave Peak Velocity;LN;18037-2 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV A Vmax, exp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve A-Wave Peak Velocity;LN;17978-8 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV DT, exp | ms | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Deceleration Time;LN;20217-6 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
LV IVRT, exp | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Isovolumic Relaxation Time;LN;18071-1 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020
P Vein S Vmax, exp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Systolic Peak Velocity;LN;29450-4 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020
P Vein D Vmax, exp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Diastolic Peak Velocity;LN;29451-2 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020
P Vein A Vmax, exp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Atrial Contraction Reversal Peak Velocity;LN;29453-8 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020
TV E Vmax, exp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve E-Wave Peak Velocity;LN;18031-5 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV A Vmax, exp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve A-Wave Peak Velocity;LN;18030-7 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV DT, exp | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Deceleration Time;LN;20217-6 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
IVC Diam, exp | cm mm | 2 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | Inferior Vena Cava Diameter;LN;18006-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020
HV S Vmax, exp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | Hepatic Vein Systolic Peak Velocity;LN;29471-0 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020
HV D Vmax, exp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | Hepatic Vein Diastolic Peak Velocity;LN;29472-8 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020
HV A Vmax, exp | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | Hepatic Vein Atrial Contraction Reversal Peak Velocity;LN;29474-4 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020
LUPV Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Diameter;SRT;M-02550 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Upper Segment;SRT;R-40491 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LUPV S Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Systolic Peak Velocity;LN;29450-4 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Upper Segment;SRT;R-40491
LUPV S VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein S-Wave Velocity Time Integral;SRT;G-038C | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Upper Segment;SRT;R-40491
LUPV D Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Diastolic Peak Velocity;LN;29451-2 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Upper Segment;SRT;R-40491
LUPV D VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein D-Wave Velocity Time Integral;SRT;G-038D | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Upper Segment;SRT;R-40491
LUPV A Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Atrial Contraction Reversal Peak Velocity;LN;29453-8 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Upper Segment;SRT;R-40491
LUPV A VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein A-Wave Velocity Time Integral;99SIEMENS;VTI_AWave | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Upper Segment;SRT;R-40491
LUPV A Duration | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein A-Wave Duration;SRT;G-038B | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Upper Segment;SRT;R-40491
LLPV Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Diameter;SRT;M-02550 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LLPV S Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Systolic Peak Velocity;LN;29450-4 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV S VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein S-Wave Velocity Time Integral;SRT;G-038C | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV D Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Diastolic Peak Velocity;LN;29451-2 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV D VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein D-Wave Velocity Time Integral;SRT;G-038D | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV A Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Atrial Contraction Reversal Peak Velocity;LN;29453-8 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV A VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein A-Wave Velocity Time Integral;99SIEMENS;VTI_AWave | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV A Duration | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein A-Wave Duration;SRT;G-038B | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
RUPV Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Diameter;SRT;M-02550 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Upper Segment;SRT;R-404A0 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RUPV S Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Systolic Peak Velocity;LN;29450-4 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Upper Segment;SRT;R-404A0
RUPV S VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein S-Wave Velocity Time Integral;SRT;G-038C | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Upper Segment;SRT;R-404A0
RUPV D Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Diastolic Peak Velocity;LN;29451-2 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Upper Segment;SRT;R-404A0
RUPV D VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein D-Wave Velocity Time Integral;SRT;G-038D | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Upper Segment;SRT;R-404A0
RUPV A Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Atrial Contraction Reversal Peak Velocity;LN;29453-8 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Upper Segment;SRT;R-404A0
RUPV A VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein A-Wave Velocity Time Integral;99SIEMENS;VTI_AWave | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Upper Segment;SRT;R-404A0
RUPV A Duration | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein A-Wave Duration;SRT;G-038B | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Upper Segment;SRT;R-404A0
RLPV Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Diameter;SRT;M-02550 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RLPV S Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Systolic Peak Velocity;LN;29450-4 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV S VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein S-Wave Velocity Time Integral;SRT;G-038C | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV D Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Diastolic Peak Velocity;LN;29451-2 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV D VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein D-Wave Velocity Time Integral;SRT;G-038D | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV A Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Atrial Contraction Reversal Peak Velocity;LN;29453-8 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV A VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein A-Wave Velocity Time Integral;99SIEMENS;VTI_AWave | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV A Duration | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein A-Wave Duration;SRT;G-038B | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
PV Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Peak Systolic Velocity;LN;11726-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV AT | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Acceleration Time;LN;20168-1 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV Accel Slope | m/s² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Acceleration Slope;LN;20167-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV ET | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Pulmonic Valve Ejection Time;LN;18042-2 | NUM
IVC Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | Inferior Vena Cava Diameter;LN;18006-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
IVC Collapsed Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | Inferior Vena Cava Diameter;LN;18006-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  | Blood vessel finding;SRT;F-3A000=Collapsed;SRT;G-A282
IVC S VTI | cm | 2 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | S-Wave Velocity Time Integral;99SIEMENS;VTI_SWave | NUM
IVC S Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | Systolic blood velocity, peak;DCM;122204 | NUM
IVC D VTI | cm | 2 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | D-Wave Velocity Time Integral;99SIEMENS;VTI_DWave | NUM
IVC D Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | Diastolic blood velocity, peak;DCM;122202 | NUM
IVC A Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | A-Wave Peak Velocity;LN;59081-0 | NUM
IVC A Duration | ms | 0 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | A-Wave Duration;LN;59105-7 | NUM
SVC S VTI | cm | 2 | Finding Site;SRT;G-C0E3=Superior Vena Cava;SRT;T-48610 | S-Wave Velocity Time Integral;99SIEMENS;VTI_SWave | NUM
SVC S Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Superior Vena Cava;SRT;T-48610 | Systolic blood velocity, peak;DCM;122204 | NUM
SVC D VTI | cm | 2 | Finding Site;SRT;G-C0E3=Superior Vena Cava;SRT;T-48610 | D-Wave Velocity Time Integral;99SIEMENS;VTI_DWave | NUM
SVC D Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Superior Vena Cava;SRT;T-48610 | Diastolic blood velocity, peak;DCM;122202 | NUM
SVC A Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Superior Vena Cava;SRT;T-48610 | A-Wave Peak Velocity;LN;59081-0 | NUM
SVC A Duration | ms | 0 | Finding Site;SRT;G-C0E3=Superior Vena Cava;SRT;T-48610 | A-Wave Duration;LN;59105-7 | NUM
HV Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | Diameter;SRT;M-02550 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
HV S VTI | cm | 2 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | S-Wave Velocity Time Integral;99SIEMENS;VTI_SWave | NUM
HV S Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | Hepatic Vein Systolic Peak Velocity;LN;29471-0 | NUM
HV D VTI | cm | 2 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | D-Wave Velocity Time Integral;99SIEMENS;VTI_DWave | NUM
HV D Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | Hepatic Vein Diastolic Peak Velocity;LN;29472-8 | NUM
HV A Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | Hepatic Vein Atrial Contraction Reversal Peak Velocity;LN;29474-4 | NUM
PV VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Velocity Time Integral;LN;20354-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV Vmean | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Blood velocity, mean;DCM;122205 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV Mean Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Mean Gradient;LN;20256-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
RVOT Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
PR Slope | m/s² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Deceleration Slope;LN;20216-8 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR PHT | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Pressure Half-Time;LN;20280-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR DT | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Deceleration Time;LN;20217-6 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR ED Velocity | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | End Diastolic Velocity;LN;11653-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Velocity Time Integral;LN;20354-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR Vmean | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Blood velocity, mean;DCM;122205 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR Mean Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Mean Gradient;LN;20256-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Peak Systolic Velocity;LN;11726-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR VC Width | cm mm | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Width;SRT;G-A220 | NUM |  | Finding Site;SRT;G-C0E3=Vena Contracta;SRT;R-421AA |  |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PV Site 2 Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2
PV Site 2 VTI | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PR PISA Radius | cm mm | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Flow Radius;LN;59102-4 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR Aliasing Velocity | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Alias Velocity;LN;59130-5 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PV Annulus s | cm mm | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Cardiac valve annulus;SRT;T-3500E |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
PV Vmax (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Peak Systolic Velocity;LN;11726-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV AT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Acceleration Time;LN;20168-1 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV Accel Slope (Post) | m/s² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Acceleration Slope;LN;20167-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV ET (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Pulmonic Valve Ejection Time;LN;18042-2 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV VTI (Post) | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Velocity Time Integral;LN;20354-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV Vmean (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Blood velocity, mean;DCM;122205 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV Mean Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Mean Gradient;LN;20256-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
RVOT Diam (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
RVOT VTI (Post) | cm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
RVOT Vmax (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Peak Systolic Velocity;LN;11726-7 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
RVOT Vmean (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Blood velocity, mean;DCM;122205 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
RVOT Mean Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Mean Gradient;LN;20256-4 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
RVOT AT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Acceleration Time;LN;20168-1 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
RVOT ET (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Ejection Time;DCM;122213 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR Slope (Post) | m/s² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Deceleration Slope;LN;20216-8 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR PHT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Pressure Half-Time;LN;20280-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR DT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Deceleration Time;LN;20217-6 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR ED Velocity (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | End Diastolic Velocity;LN;11653-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR ED Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Pressure Gradient;SRT;F-023F7 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011 |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR VTI (Post) | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Velocity Time Integral;LN;20354-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR Vmean (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Blood velocity, mean;DCM;122205 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR Mean Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Mean Gradient;LN;20256-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR Vmax (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Peak Systolic Velocity;LN;11726-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR VC Width (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Width;SRT;G-A220 | NUM |  | Finding Site;SRT;G-C0E3=Vena Contracta;SRT;R-421AA |  |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV Site 2 Diam (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV Site 2 VTI (Post) | cm | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR PISA Radius (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Flow Radius;LN;59102-4 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR Aliasing Velocity (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Alias Velocity;LN;59130-5 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV Annulus s (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Cardiac valve annulus;SRT;T-3500E |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
RA Major d A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Major Axis;SRT;G-A193 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RA Major s A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Major Axis;SRT;G-A193 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RA Minor d A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Minor Axis;SRT;G-A194 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RA Minor s A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Minor Axis;SRT;G-A194 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RA Area d A4C | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Right Atrium Area A4C view;LN;17988-7 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RA Vol d A4C MOD | ml | 1 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RA Area s A4C | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Right Atrium Area A4C view;LN;17988-7 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RA Vol s A4C MOD | ml | 1 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV E VTI, insp | cm | 3 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-Wave Velocity Time Integral;99SIEMENS;VTI_MVEWave | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV A VTI, insp | cm | 3 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve A-Wave Velocity Time Integral;99SIEMENS;VTI_MVAWave | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
P Vein S VTI, insp | cm | 3 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein S-Wave Velocity Time Integral;SRT;G-038C | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010
P Vein D VTI, insp | cm | 3 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein D-Wave Velocity Time Integral;SRT;G-038D | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010
P Vein A VTI, insp | cm | 3 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein A-Wave Velocity Time Integral;99SIEMENS;VTI_AWave | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010
P Vein A Duration, insp | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein A-Wave Duration;SRT;G-038B | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Inspiration;SRT;F-20010
MV E VTI, exp | cm | 3 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-Wave Velocity Time Integral;99SIEMENS;VTI_MVEWave | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV A VTI, exp | cm | 3 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve A-Wave Velocity Time Integral;99SIEMENS;VTI_MVAWave | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
P Vein S VTI, exp | cm | 3 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein S-Wave Velocity Time Integral;SRT;G-038C | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020
P Vein D VTI, exp | cm | 3 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein D-Wave Velocity Time Integral;SRT;G-038D | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020
P Vein A VTI, exp | cm | 3 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein A-Wave Velocity Time Integral;99SIEMENS;VTI_AWave | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020
P Vein A Duration, exp | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein A-Wave Duration;SRT;G-038B | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Expiration;SRT;F-20020
TV E Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve E-Wave Peak Velocity;LN;18031-5 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV A Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve A-Wave Peak Velocity;LN;18030-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV E Vmax, Annulus | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve E-Wave Peak Velocity;LN;18031-5 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV A Vmax, Annulus | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve A-Wave Peak Velocity;LN;18030-7 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV DT, Annulus | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Deceleration Time;LN;20217-6 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV A Duration, Annulus | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | A-Wave Duration;LN;59105-7 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
RV IVRT, Annulus | ms | 0 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Isovolumic Relaxation Time;LN;59083-6 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111
TV e' Vmax (L) | m/s cm/s | 3 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Peak Early Diastolic Tissue Velocity;99SIEMENS;RV_Ea | NUM |  | Finding Site;SRT;G-C0E3=Lateral Tricuspid Annulus;99SIEMENS;LatTriAnn |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
TV a' Vmax (L) | m/s cm/s | 3 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | RV Peak Diastolic Tissue Velocity During Atrial Systole;99SIEMENS;RV_Aa | NUM |  | Finding Site;SRT;G-C0E3=Lateral Tricuspid Annulus;99SIEMENS;LatTriAnn |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
TV s' Vmax (L) | m/s cm/s | 3 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Peak Systolic Tissue Velocity;99SIEMENS;RV_Sa | NUM |  | Finding Site;SRT;G-C0E3=Lateral Tricuspid Annulus;99SIEMENS;LatTriAnn |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
RVAW d PLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Anterior Wall Diastolic Thickness;LN;18153-7 | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVAW s PLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Anterior Wall Systolic Thickness;LN;18157-8 | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVID d PLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Internal Diastolic Dimension;LN;20304-2 | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVID s PLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Internal Systolic Dimension;LN;20305-9 | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Major d A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Major Axis;SRT;G-A193 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Major s A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Major Axis;SRT;G-A193 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Minor d A4C Base | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Internal Diastolic Dimension;LN;20304-2 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C | Finding Site;SRT;G-C0E3=Basal;SRT;G-A123 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Minor s A4C Base | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Internal Systolic Dimension;LN;20305-9 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C | Finding Site;SRT;G-C0E3=Basal;SRT;G-A123 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Minor d A4C Mid | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Internal Diastolic Dimension;LN;20304-2 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C | Finding Site;SRT;G-C0E3=Medial;SRT;R-404D5 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Minor s A4C Mid | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Internal Systolic Dimension;LN;20305-9 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C | Finding Site;SRT;G-C0E3=Medial;SRT;R-404D5 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV dP/dt | mmHg/s | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Regurgitation dP/dt;LN;18034-9 | NUM
RV PEP | ms | 0 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Pre-Ejection Period;LN;59085-1 | NUM
RV IVCT | ms | 0 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Isovolumic Contraction Time;LN;59084-4 | NUM
TV Close-Open Duration | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve Closure to Opening Time;SRT;G-0389 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
RV AT | ms | 0 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Acceleration Time;LN;20168-1 | NUM
RV ET | ms | 0 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Ejection Time;DCM;122213 | NUM
RVOT Diam PLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVOT 1 Diam PSAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Cardiovascular Orifice Diameter 1;99SIEMENS;CardOrificeDiam1 | NUM | Image View;DCM;111031=Parasternal short axis;SRT;G-0397 |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVOT 2 Diam PSAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Cardiovascular Orifice Diameter 2;99SIEMENS;CardOrificeDiam2 | NUM | Image View;DCM;111031=Parasternal short axis;SRT;G-0397 |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVOT Diam SCLAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | Image View;DCM;111031=Subcostal long axis;SRT;G-039E |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVOT Diam PSAX | cm mm | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | Image View;DCM;111031=Parasternal short axis;SRT;G-0397 |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Area d A4C | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Area;SRT;G-A166 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Vol d A4C MOD | ml | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Area s A4C | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Area;SRT;G-A166 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Vol s A4C MOD | ml | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
TV Area Planimetry | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Planimetry;DCM;125220 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
TV Decel Slope | m/s² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Deceleration Slope;LN;20216-8 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV DT | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Deceleration Time;LN;20217-6 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV PHT | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Pressure Half-Time;LN;20280-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV VTI | cm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Velocity Time Integral;LN;20354-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Vmax | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Velocity;LN;20351-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Vmean | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Blood velocity, mean;DCM;122205 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Mean Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Mean Gradient;LN;20256-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV AT | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Acceleration Time;LN;20168-1 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV DFT | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Diastolic Filling Period;DFPt;SRT;R-003A9 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV PISA Radius | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Flow Radius;LN;59102-4 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Aliasing Velocity | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Alias Velocity;LN;59130-5 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Funnel Angle | ° | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Funnel Angle;99SIEMENS;FunnelAngle | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TR VTI | cm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Velocity Time Integral;LN;20354-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR Vmean | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Blood velocity, mean;DCM;122205 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR Mean Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Mean Gradient;LN;20256-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR AT | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Acceleration Time;LN;20168-1 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR Duration | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Duration;SRT;G-7290 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR VC Width | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Width;SRT;G-A220 | NUM |  | Finding Site;SRT;G-C0E3=Vena Contracta;SRT;R-421AA |  |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TV Site 2 Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
TV Site 2 VTI | cm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TR PISA Radius | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Flow Radius;LN;59102-4 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR PISA | cm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Flow Area;LN;20226-7 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR Aliasing Velocity | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Alias Velocity;LN;59130-5 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TV Diam | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM
TV Annulus d | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
TV Annulus d A-P | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Anterior-Posterior Diameter;99SIEMENS;APDiameter | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
TV Annulus d M-L | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Medial-Lateral Diameter;99SIEMENS;MLDiameter | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
TV Annulus d A4C | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
TV VTI, Annulus | cm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Vmax, Annulus | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Velocity;LN;20351-3 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Vmean, Annulus | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Blood velocity, mean;DCM;122205 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Mean Grad, Annulus | mmHg | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Mean Gradient;LN;20256-4 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV AT, Annulus | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Acceleration Time;LN;20168-1 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV DFT, Annulus | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Diastolic Filling Period;DFPt;SRT;R-003A9 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV E Vmax (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve E-Wave Peak Velocity;LN;18031-5 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV A Vmax (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve A-Wave Peak Velocity;LN;18030-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Area Planimetry (Post) | cm² mm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Planimetry;DCM;125220 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Decel Slope (Post) | m/s² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Deceleration Slope;LN;20216-8 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV DT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Deceleration Time;LN;20217-6 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV PHT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Pressure Half-Time;LN;20280-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV VTI (Post) | cm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Velocity Time Integral;LN;20354-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Vmax (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Velocity;LN;20351-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Vmean (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Blood velocity, mean;DCM;122205 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Mean Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Mean Gradient;LN;20256-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV AT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Acceleration Time;LN;20168-1 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV DFT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Diastolic Filling Period;DFPt;SRT;R-003A9 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV PISA Radius (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Flow Radius;LN;59102-4 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Aliasing Velocity (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Alias Velocity;LN;59130-5 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Funnel Angle (Post) | ° | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Funnel Angle;99SIEMENS;FunnelAngle | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR Vmax (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Velocity;LN;20351-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR Peak Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR VTI (Post) | cm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Velocity Time Integral;LN;20354-7 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR Vmean (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Blood velocity, mean;DCM;122205 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR Mean Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Mean Gradient;LN;20256-4 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR AT (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Acceleration Time;LN;20168-1 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR Duration (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Duration;SRT;G-7290 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR VC Width (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Width;SRT;G-A220 | NUM |  | Finding Site;SRT;G-C0E3=Vena Contracta;SRT;R-421AA |  |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
RV dP/dt (Post) | mmHg/s | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Regurgitation dP/dt;LN;18034-9 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Site 2 Diam (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Site 2 VTI (Post) | cm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR PISA Radius (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Flow Radius;LN;59102-4 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR PISA (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Flow Area;LN;20226-7 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR Aliasing Velocity (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Alias Velocity;LN;59130-5 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Diam (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Annulus d (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Annulus d A-P (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Anterior-Posterior Diameter;99SIEMENS;APDiameter | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Annulus d M-L (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Medial-Lateral Diameter;99SIEMENS;MLDiameter | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Annulus d A4C (Post) | cm mm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV VTI, Annulus (Post) | cm | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Vmax, Annulus (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Velocity;LN;20351-3 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Vmean, Annulus (Post) | m/s cm/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Blood velocity, mean;DCM;122205 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Mean Grad, Annulus (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Mean Gradient;LN;20256-4 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV AT, Annulus (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Acceleration Time;LN;20168-1 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV DFT, Annulus (Post) | ms | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Diastolic Filling Period;DFPt;SRT;R-003A9 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
RA Pressure (Post) | mmHg | 0 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Right Atrium Systolic Pressure;LN;18070-3 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao d/LA s (M-mode) | unitless | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | End Diastolic Aortic Root to End Systolic Left Atrium Ratio;99SIEMENS;RATIOaod_las | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LA s/Ao d (M-mode) | unitless | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium to Aortic Root Ratio;LN;17985-3 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
CA/CE | unitless | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | A-Wave Amplitude to E-Wave Amplitude;99SIEMENS;CA2CE | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
CE/CA | unitless | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | E-Wave Amplitude to A-Wave Amplitude;99SIEMENS;CE2CA | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
EPSS/LVID d | unitless | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | EPSS to Diastolic LVID Ratio;99SIEMENS;RATIOepss_lvidd | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
IVS/LVPW (M-mode) | unitless | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Interventricular Septum to Posterior Wall Thickness Ratio;LN;18155-2 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
IVS %Thick (M-mode) | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Interventricular Septum % Thickening;LN;18054-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LVPW %Thick (M-mode) | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Posterior Wall % Thickening;LN;18053-9 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV %FS | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Fractional Shortening;LN;18051-3 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV EF | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Ejection Fraction by US;LN;18043-0 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV Vol d | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV Vol s | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV SV | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Volume;SRT;F-32120 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV SI | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Index;SRT;F-00078 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=M mode;SRT;G-0394 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV CO | l/min | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Output;SRT;F-32100 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV CI | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Index;SRT;F-32110 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV Mass ASE (M-mode) | g | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Mass;LN;18087-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV Mass ASE/BSA (M-mode) | g/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Mass;LN;18087-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Mass ASE/Ht (M-mode) | g/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Mass;LN;18087-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
Mean VCFc | circs/s | 3 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | HR-Corrected Mean Velocity of Circumferential Fiber Shortening;LN;59118-0 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
Wall Stress Circum | g/cm² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Circumferential Wall Stress;99SIEMENS;WScircum | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
Wall Stress Merid | g/cm² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Meridional Wall Stress;LN;59097-6 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV ETc | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Heart Rate-Corrected Ejection Time;LN;59086-9 | NUM
LV PEPc | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Heart Rate-Corrected Pre-Ejection Period;LN;59087-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV PEP/ET | unitless | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Pre-Ejection Period/Ejection Time Ratio;LN;59088-5 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
RV ETc | ms | 0 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Heart Rate-Corrected Ejection Time;LN;59086-9 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
RV PEPc | ms | 0 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Heart Rate-Corrected Pre-Ejection Period;LN;59087-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
RV PEP/ET | unitless | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Pre-Ejection Period/Ejection Time Ratio;LN;59088-5 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
IVC Diam % change (M-mode) | % | 1 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | Inferior Vena Cava % Collapse;LN;18050-5 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
IVC Diam change (M-mode) | cm | 2 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | Change in Diameter;99SIEMENS;DiamChange | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
Ao Root d/BSA | cm/m² | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Aortic Root Diameter;LN;18015-8 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
Ao Sinus d/BSA | cm/m² | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Aortic Root Diameter;LN;18015-8 | NUM |  | Finding Site;SRT;G-C0E3=Structure Sinus of Valsalva;SRT;T-42200 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
Ao STJ d/BSA | cm/m² | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Aortic Root Diameter;LN;18015-8 | NUM |  | Finding Site;SRT;G-C0E3=Aortic Sinotubular Junction;SRT;T-42102 |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
Ao Asc d/BSA | cm/m² | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Ascending Aortic Diameter;LN;18012-5 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
Ao Arch d/BSA | cm/m² | 2 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Aortic Arch Diameter;LN;18011-7 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
AoV Ann s/BSA | cm/m² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Aortic Valve Ring;SRT;T-35410 |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
AoV AT/ET | unitless | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Ratio of Aortic Valve Acceleration Time to Ejection Time;SRT;G-0382 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
LVOT/AoV Vmax | unitless | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Ratio of LVOT Maximum velocity to Aortic Valve Maximum velocity;99SIEMENS;RatioVmaxLVOT_AV | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
LVOT/AoV VTI | unitless | 3 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Ratio of LVOT VTI to Aortic Valve VTI;99SIEMENS;RatioVTILVOT_AV | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AVA (Vmax) | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Peak Velocity;DCM;125214 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AVA (Vmean) | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Mean Velocity;DCM;125213 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AVA (VTI) | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AVA/BSA (Vmax) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Peak Velocity;DCM;125214 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6
AVA/BSA (Vmean) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Mean Velocity;DCM;125213 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6
AVA/BSA (VTI) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Pressure ED | mmHg | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End-Diastolic Pressure;SRT;F-03E0E | NUM
AR Jet/LVOT Area d | unitless | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | AR Jet Area/LVOT Area Ratio;99SIEMENS;RatioARA_LVOT | NUM | [5 empty] | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2
AR Jet/LVOT Ht d | unitless | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | AR Jet Height/LVOT Height Ratio;99SIEMENS;RatioARHt_LVOT | NUM | [5 empty] | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2
AoV Site 2 Area | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
AoV Site 2 SV | ml | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR Volume (Doppler) | ml | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Volume;SRT;G-D705 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR RF (Doppler) | % | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Regurgitant Fraction;SRT;G-0390 | NUM
AR PISA | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Flow Area;LN;20226-7 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR Inst Flow Rate | ml/s | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Peak Instantaneous Flow Rate;LN;34141-2 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR EROA | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR Volume (PISA) | ml | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Volume;SRT;G-D705 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR RF (PISA) | % | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Regurgitant Fraction;SRT;G-0390 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216
AoV Ann Area | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Aortic Valve Ring;SRT;T-35410 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
AoV SV | ml | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Stroke Volume;SRT;F-32120 | NUM
AoV SI | ml/m² | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Stroke Index;SRT;F-00078 | NUM
AoV CO | l/min | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiac Output;SRT;F-32100 | NUM
AoV CI | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiac Index;SRT;F-32110 | NUM
AoV Ann s/BSA (Post) | cm/m² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM |  | Finding Site;SRT;G-C0E3=Aortic Valve Ring;SRT;T-35410 |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV AT/ET (Post) | unitless | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Ratio of Aortic Valve Acceleration Time to Ejection Time;SRT;G-0382 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LVOT Area s (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 |  | Cardiac Cycle Point;SRT;R-4089A=End Systole;SRT;R-FAB5B |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LVOT SV (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LVOT/AoV Vmax (Post) | unitless | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Ratio of LVOT Maximum velocity to Aortic Valve Maximum velocity;99SIEMENS;RatioVmaxLVOT_AV | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LVOT/AoV VTI (Post) | unitless | 3 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Ratio of LVOT VTI to Aortic Valve VTI;99SIEMENS;RatioVTILVOT_AV | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AVA (Vmax) (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Peak Velocity;DCM;125214 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AVA (Vmean) (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Mean Velocity;DCM;125213 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AVA (VTI) (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AVA/BSA (Vmax) (Post) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Peak Velocity;DCM;125214 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AVA/BSA (Vmean) (Post) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Mean Velocity;DCM;125213 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AVA/BSA (VTI) (Post) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LV Pressure s (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Systolic Pressure;SRT;F-03E0D | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LV Pressure ED (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End-Diastolic Pressure;SRT;F-03E0E | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR Jet/LVOT Area d (Post) | unitless | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | AR Jet Area/LVOT Area Ratio;99SIEMENS;RatioARA_LVOT | NUM | [5 empty] | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR Jet/LVOT Ht d (Post) | unitless | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | AR Jet Height/LVOT Height Ratio;99SIEMENS;RatioARHt_LVOT | NUM | [5 empty] | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV Site 2 Area (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV Site 2 SV (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR Volume (Doppler) (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Volume;SRT;G-D705 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR RF (Doppler) (Post) | % | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Regurgitant Fraction;SRT;G-0390 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR PISA (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Flow Area;LN;20226-7 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR Inst Flow Rate (Post) | ml/s | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Peak Instantaneous Flow Rate;LN;34141-2 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR EROA (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR Volume (PISA) (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Volume;SRT;G-D705 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR RF (PISA) (Post) | % | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Regurgitant Fraction;SRT;G-0390 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 | [7 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV Ann Area (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Aortic Valve Ring;SRT;T-35410 |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV SV (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Stroke Volume;SRT;F-32120 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV SI (Post) | ml/m² | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Stroke Index;SRT;F-00078 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV CO (Post) | l/min | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiac Output;SRT;F-32100 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AoV CI (Post) | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiac Index;SRT;F-32110 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
RVSP (VSD) | mmHg | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Peak Systolic Pressure;SRT;G-0380 | NUM |  | Finding Site;SRT;G-C0E3=Ventricular Septal Defect;SRT;D4-31150
Pulm Side Area | cm² | 3 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Pulmonary Side;99SIEMENS;PulmonarySide
Pulm Side SV | ml | 1 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Pulmonary Side;99SIEMENS;PulmonarySide
Pulm Side SI | ml/m² | 1 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Stroke Index;SRT;F-00078 | NUM |  | Finding Site;SRT;G-C0E3=Pulmonary Side;99SIEMENS;PulmonarySide
Pulm Side CO | l/min | 2 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Cardiac Output;SRT;F-32100 | NUM |  | Finding Site;SRT;G-C0E3=Pulmonary Side;99SIEMENS;PulmonarySide
Pulm Side CI | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Cardiac Index;SRT;F-32110 | NUM |  | Finding Site;SRT;G-C0E3=Pulmonary Side;99SIEMENS;PulmonarySide
Sys Side Area | cm² | 3 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Systemic Side;99SIEMENS;SystemicSide
Sys Side SV | ml | 1 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Systemic Side;99SIEMENS;SystemicSide
Sys Side SI | ml/m² | 1 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Stroke Index;SRT;F-00078 | NUM |  | Finding Site;SRT;G-C0E3=Systemic Side;99SIEMENS;SystemicSide
Sys Side CO | l/min | 2 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Cardiac Output;SRT;F-32100 | NUM |  | Finding Site;SRT;G-C0E3=Systemic Side;99SIEMENS;SystemicSide
Sys Side CI | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Cardiac Index;SRT;F-32110 | NUM |  | Finding Site;SRT;G-C0E3=Systemic Side;99SIEMENS;SystemicSide
Qp / Qs | unitless | 2 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Pulmonary-to-Systemic Shunt Flow Ratio;LN;29462-9 | NUM
Qp-Qs | l/min | 2 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Difference between Pulmonary Valve CO and Aortic Valve CO;99SIEMENS;QpMinusQs | NUM
Ao d/LA s | unitless | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | End Diastolic Aortic Root to End Systolic Left Atrium Ratio;99SIEMENS;RATIOaod_las | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA s/Ao d | unitless | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium to Aortic Root Ratio;LN;17985-3 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA s/BSA | cm/m² | 2 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Diameter;SRT;M-02550 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LA %FAC A4C | % | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrial Fractional Area Change;99SIEMENS;LAFAC | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA %FAC A2C | % | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrial Fractional Area Change;99SIEMENS;LAFAC | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA EF A4C A-L | % | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Cardiac ejection fraction;SRT;F-32070 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol d A4C A-L | ml | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol s A4C A-L | ml | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium Systolic Volume;SRT;G-0383 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol d A4C A-L/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LA Vol s A4C A-L/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium Systolic Volume;SRT;G-0383 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LA EF A2C A-L | % | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Cardiac ejection fraction;SRT;F-32070 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol d A2C A-L | ml | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol s A2C A-L | ml | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium Systolic Volume;SRT;G-0383 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol d A2C A-L/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LA Vol s A2C A-L/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium Systolic Volume;SRT;G-0383 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LA EF Biplane A-L | % | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Cardiac ejection fraction;SRT;F-32070 | NUM |  |  | Measurement Method;SRT;G-C036=Area-Length Biplane;DCM;125204 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol d Biplane A-L | ml | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Volume;SRT;G-D705 | NUM |  |  | Measurement Method;SRT;G-C036=Area-Length Biplane;DCM;125204 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol s Biplane A-L | ml | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium Systolic Volume;SRT;G-0383 | NUM |  |  | Measurement Method;SRT;G-C036=Area-Length Biplane;DCM;125204 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol d Biplane A-L/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Volume;SRT;G-D705 | NUM |  |  | Measurement Method;SRT;G-C036=Area-Length Biplane;DCM;125204 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LA Vol s Biplane A-L/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium Systolic Volume;SRT;G-0383 | NUM |  |  | Measurement Method;SRT;G-C036=Area-Length Biplane;DCM;125204 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LA EF A4C MOD | % | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Cardiac ejection fraction;SRT;F-32070 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol d A4C MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LA Vol s A4C MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium Systolic Volume;SRT;G-0383 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LA EF A2C MOD | % | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Cardiac ejection fraction;SRT;F-32070 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol d A2C MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LA Vol s A2C MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium Systolic Volume;SRT;G-0383 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LA EF Biplane MOD | % | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Cardiac ejection fraction;SRT;F-32070 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Biplane;DCM;125207 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol d Biplane MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Volume;SRT;G-D705 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Biplane;DCM;125207 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol s Biplane MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium Systolic Volume;SRT;G-0383 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Biplane;DCM;125207 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Vol d Biplane MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Volume;SRT;G-D705 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Biplane;DCM;125207 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LA Vol s Biplane MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Left Atrium Systolic Volume;SRT;G-0383 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Biplane;DCM;125207 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
P Vein S/D Vmax | unitless | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Systolic to Diastolic Ratio;LN;29452-0 | NUM
P Vein A/MV A Dur | unitless | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein A Duration Mitral Valve A Duration Ratio;LN;59112-3 | NUM
IVS/LVPW | unitless | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Interventricular Septum to Posterior Wall Thickness Ratio;LN;18155-2 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
IVS %Thick | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Interventricular Septum % Thickening;LN;18054-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVPW %Thick | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Posterior Wall % Thickening;LN;18053-9 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVID d/BSA | cm/m² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Internal End Diastolic Dimension;LN;29436-3 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LVID s/BSA | cm/m² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Internal Systolic Dimension;LN;29438-9 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Relative Wall Thickness (LV RWT) | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Relative Wall Thickness;99SIEMENS;LVRelWallThick | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Eccentricity Index d (LVEI d) | unitless | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | LV Eccentricity Index;99SIEMENS;LVEccentrictyIdx | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Eccentricity Index s (LVEI s) | unitless | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | LV Eccentricity Index;99SIEMENS;LVEccentrictyIdx | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=End Systole;SRT;R-FAB5B |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV MPI | unitless | 3 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Index of Myocardial Performance;SRT;G-037F | NUM
LV PEP/LV ET | unitless | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Pre-Ejection Period/Ejection Time Ratio;LN;59088-5 | NUM
dPEP | ms | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Interventricular Delay;99SIEMENS;LVRVdPEP | NUM
LV AT/LV ET | unitless | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Ratio of Left Ventricular Acceleration Time to Ejection Time;99SIEMENS;RatioLVAT_ET | NUM
LV Pressure s | mmHg | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Systolic Pressure;SRT;F-03E0D | NUM
LVOT Area s | cm² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 |  | Cardiac Cycle Point;SRT;R-4089A=End Systole;SRT;R-FAB5B |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVOT SV | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT SI | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Index;SRT;F-00078 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT CO | l/min | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Output;SRT;F-32100 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT CI | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Index;SRT;F-32110 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LV %FS Cubed | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Fractional Shortening;LN;18051-3 | NUM |  |  | Measurement Method;SRT;G-C036=Cube Method;DCM;125206 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d Cubed | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  |  | Measurement Method;SRT;G-C036=Cube Method;DCM;125206 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol s Cubed | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  |  | Measurement Method;SRT;G-C036=Cube Method;DCM;125206 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV EF PSAX Cubed | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Ejection Fraction by US;LN;18043-0 | NUM | Image View;DCM;111031=Parasternal short axis;SRT;G-0397 |  | Measurement Method;SRT;G-C036=Cube Method;DCM;125206 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV EF Cubed | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Ejection Fraction by US;LN;18043-0 | NUM |  |  | Measurement Method;SRT;G-C036=Cube Method;DCM;125206 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV SV Cubed | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Volume;SRT;F-32120 | NUM |  |  | Measurement Method;SRT;G-C036=Cube Method;DCM;125206 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV SI Cubed | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Index;SRT;F-00078 | NUM |  |  | Measurement Method;SRT;G-C036=Cube Method;DCM;125206 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CO Cubed | l/min | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Output;SRT;F-32100 | NUM |  |  | Measurement Method;SRT;G-C036=Cube Method;DCM;125206 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CI Cubed | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Index;SRT;F-32110 | NUM |  |  | Measurement Method;SRT;G-C036=Cube Method;DCM;125206 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV %FS (Teichholz) | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Fractional Shortening;LN;18051-3 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV EF (Teichholz) | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Ejection Fraction by US;LN;18043-0 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV EF PSAX (Teichholz) | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Ejection Fraction by US;LN;18043-0 | NUM | Image View;DCM;111031=Parasternal short axis;SRT;G-0397 |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d (Teichholz) | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol s (Teichholz) | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d (Teichholz)/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Vol s (Teichholz)/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Vol d (Teichholz)/HT | ml/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
LV Vol s (Teichholz)/HT | ml/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
LV SV (Teichholz) | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Volume;SRT;F-32120 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV SI (Teichholz) | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Index;SRT;F-00078 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CO (Teichholz) | l/min | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Output;SRT;F-32100 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CI (Teichholz) | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Index;SRT;F-32110 | NUM |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV %FAC A4C | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Fractional Area Change;SRT;G-0376 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV EF A4C A-L | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Ejection Fraction by US;LN;18043-0 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d A4C A-L | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol s A4C A-L | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d A4C A-L/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Vol s A4C A-L/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Vol d A4C A-L/HT | ml/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
LV Vol s A4C A-L/HT | ml/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
LV %FAC A2C | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Fractional Area Change;SRT;G-0376 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV EF A2C A-L | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Ejection Fraction by US;LN;18043-0 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d A2C A-L | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol s A2C A-L | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d A2C A-L/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Vol s A2C A-L/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Vol d A2C A-L/HT | ml/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
LV Vol s A2C A-L/HT | ml/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
LV EF Biplane A-L | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Ejection Fraction by US;LN;18043-0 | NUM |  |  | Measurement Method;SRT;G-C036=Area-Length Biplane;DCM;125204 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d Biplane A-L | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  |  | Measurement Method;SRT;G-C036=Area-Length Biplane;DCM;125204 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol s Biplane A-L | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  |  | Measurement Method;SRT;G-C036=Area-Length Biplane;DCM;125204 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d Biplane A-L/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  |  | Measurement Method;SRT;G-C036=Area-Length Biplane;DCM;125204 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Vol s Biplane A-L/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  |  | Measurement Method;SRT;G-C036=Area-Length Biplane;DCM;125204 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Vol d Biplane A-L/HT | ml/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  |  | Measurement Method;SRT;G-C036=Area-Length Biplane;DCM;125204 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
LV Vol s Biplane A-L/HT | ml/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  |  | Measurement Method;SRT;G-C036=Area-Length Biplane;DCM;125204 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
LV EF A4C MOD | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Ejection Fraction by US;LN;18043-0 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d A4C MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Vol s A4C MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Vol d A4C MOD/HT | ml/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
LV Vol s A4C MOD/HT | ml/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
LV SV A4C MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV SI A4C MOD | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Index;SRT;F-00078 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CO A4C MOD | l/min | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Output;SRT;F-32100 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CI A4C MOD | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Index;SRT;F-32110 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV EF A2C MOD | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Ejection Fraction by US;LN;18043-0 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d A2C MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Vol s A2C MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Vol d A2C MOD/HT | ml/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
LV Vol s A2C MOD/HT | ml/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
LV SV A2C MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV SI A2C MOD | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Index;SRT;F-00078 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CO A2C MOD | l/min | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Output;SRT;F-32100 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CI A2C MOD | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Index;SRT;F-32110 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Major s max | cm | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Systolic Major Axis;LN;18076-0 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV EF Biplane MOD | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Ejection Fraction by US;LN;18043-0 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Biplane;DCM;125207 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d Biplane MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Biplane;DCM;125207 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol s Biplane MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Biplane;DCM;125207 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d Biplane MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Biplane;DCM;125207 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Vol s Biplane MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Biplane;DCM;125207 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Vol d Biplane MOD/HT | ml/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Biplane;DCM;125207 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
LV Vol s Biplane MOD/HT | ml/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Biplane;DCM;125207 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
LV SV Biplane MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Volume;SRT;F-32120 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Biplane;DCM;125207 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV SI Biplane MOD | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Index;SRT;F-00078 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Biplane;DCM;125207 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CO Biplane MOD | l/min | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Output;SRT;F-32100 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Biplane;DCM;125207 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CI Biplane MOD | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Index;SRT;F-32110 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks, Biplane;DCM;125207 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Mass ASE | g | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Mass;LN;18087-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Mass ASE/BSA | g/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Mass;LN;18087-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Mass ASE/Ht | g/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Mass;LN;18087-7 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
LV Mass A-L | g | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Mass;LN;18087-7 | NUM |  |  | Measurement Method;SRT;G-C036=Left Ventricle Mass by Area Length;DCM;125270 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Mean Wall Thickness d | cm | 3 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Myocardial Thickness;99SIEMENS;THI_Myo | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Mass/BSA (A-L) | g/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Mass;LN;18087-7 | NUM |  |  | Measurement Method;SRT;G-C036=Left Ventricle Mass by Area Length;DCM;125270 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Mass/Ht (A-L) | g/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Mass;LN;18087-7 | NUM |  |  | Measurement Method;SRT;G-C036=Left Ventricle Mass by Area Length;DCM;125270 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
LV Mass T-E | g | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Mass;LN;18087-7 | NUM |  |  | Measurement Method;SRT;G-C036=Left Ventricle Mass by Truncated Ellipse;DCM;125222 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Minor Rad d PSAX pap | cm | 3 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Minor Radius;99SIEMENS;MinorRadius | NUM | Image View;DCM;111031=Parasternal short axis at the Papillary Muscle level;SRT;G-039B |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Mass/BSA (T-E) | g/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Mass;LN;18087-7 | NUM |  |  | Measurement Method;SRT;G-C036=Left Ventricle Mass by Truncated Ellipse;DCM;125222 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
LV Mass/Ht (T-E) | g/cm | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle Mass;LN;18087-7 | NUM |  |  | Measurement Method;SRT;G-C036=Left Ventricle Mass by Truncated Ellipse;DCM;125222 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Patient Height;LN;8302-2
MV E/A | unitless | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E to A Ratio;LN;18038-0 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV E/A (Valsalva) | unitless | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E to A Ratio;LN;18038-0 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Valsalva maneuver;SRT;R-40928 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV A/E | unitless | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve A to E Ratio;99SIEMENS;MVAtoERatio | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Area (VTI) | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Area/BSA (VTI) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6
MV PISA | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Flow Area;LN;20226-7 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Inst Flow Rate | l/min | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Instantaneous Flow Rate;LN;34141-2 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Area (PISA) | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Area/BSA (PISA) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6
MV Site 2 Area | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV Site 2 SV | ml | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR Volume (Doppler) | ml | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Volume;SRT;G-D705 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR RF (Doppler) | % | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Regurgitant Fraction;SRT;G-0390 | NUM
MR PISA | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Flow Area;LN;20226-7 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR Inst Flow Rate | ml/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Instantaneous Flow Rate;LN;34141-2 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR EROA | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR Volume (PISA) | ml | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Volume;SRT;G-D705 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR RF (PISA) | % | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Regurgitant Fraction;SRT;G-0390 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MV Annulus Area | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV SV | ml | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Stroke Volume;SRT;F-32120 | NUM
MV SI | ml/m² | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Stroke Index;SRT;F-00078 | NUM
MV CO | l/min | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiac Output;SRT;F-32100 | NUM
MV CI | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiac Index;SRT;F-32110 | NUM
MV E/A (Annulus) | unitless | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E to A Ratio;LN;18038-0 | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV A/E (Annulus) | unitless | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve A to E Ratio;99SIEMENS;MVAtoERatio | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV E/MV VTI (Annulus) | unitless | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Ratio of MV VTIe Annulus to MV VTI Annulus;99SIEMENS;VTIe2VTIAnn | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV A/MV VTI (Annulus) | unitless | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Ratio of MV VTIa Annulus to MV VTI Annulus;99SIEMENS;VTIa2VTIAnn | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV E/e' (M) | unitless | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave;SRT;G-037B | NUM |  | Finding Site;SRT;G-C0E3=Medial Mitral Annulus;SRT;G-0391
MV E/e' (L) | unitless | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave;SRT;G-037B | NUM |  | Finding Site;SRT;G-C0E3=Lateral Mitral Annulus;SRT;G-0392
MV E/E' Vmax (Average) | unitless | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave;SRT;G-037B | NUM
MV E/A (Post) | unitless | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E to A Ratio;LN;18038-0 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV E/A (Valsalva) (Post) | unitless | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E to A Ratio;LN;18038-0 | NUM | [6 empty] | Respiratory Cycle Point;SRT;R-40899=Valsalva maneuver;SRT;R-40928 | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV A/E (Post) | unitless | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve A to E Ratio;99SIEMENS;MVAtoERatio | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Area (VTI) (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Area/BSA (VTI) (Post) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV PISA (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Flow Area;LN;20226-7 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Inst Flow Rate (Post) | l/min | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Instantaneous Flow Rate;LN;34141-2 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Area (PISA) (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Area/BSA (PISA) (Post) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Site 2 Area (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Site 2 SV (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR Volume (Doppler) (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Volume;SRT;G-D705 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR RF (Doppler) (Post) | % | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Regurgitant Fraction;SRT;G-0390 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR PISA (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Flow Area;LN;20226-7 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR Inst Flow Rate (Post) | ml/s | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Instantaneous Flow Rate;LN;34141-2 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR EROA (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR Volume (PISA) (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Volume;SRT;G-D705 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MR RF (PISA) (Post) | % | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Regurgitant Fraction;SRT;G-0390 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Annulus Area (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV SV (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Stroke Volume;SRT;F-32120 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV SI (Post) | ml/m² | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Stroke Index;SRT;F-00078 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV CO (Post) | l/min | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiac Output;SRT;F-32100 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV CI (Post) | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiac Index;SRT;F-32110 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV e'/a' (M) | unitless | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle E to A Tissue Velocity Ratio;LN;59129-7 | NUM |  | Finding Site;SRT;G-C0E3=Medial Mitral Annulus;SRT;G-0391 |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV e'/a' (L) | unitless | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricle E to A Tissue Velocity Ratio;LN;59129-7 | NUM |  | Finding Site;SRT;G-C0E3=Lateral Mitral Annulus;SRT;G-0392 |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
RA d Diam % change | % | 1 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | % Change in Diameter;99SIEMENS;DiamPercntChange | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVID d % change | % | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | % Change in Diameter;99SIEMENS;DiamPercntChange | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA d Diam % change | % | 1 | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | % Change in Diameter;99SIEMENS;DiamPercntChange | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVID d % change | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | % Change in Diastolic minor axis Dimension;99SIEMENS;LVIDdPrcntChange | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
TV E Vmax % change | % | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | % Change in E-Wave Peak Velocity;99SIEMENS;VMAXePrcntChange | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV A Vmax % change | % | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | % Change in A-Wave Peak Velocity;99SIEMENS;VMAXaPrcntChange | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV DT % change | % | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | % Change in Deceleration Time;99SIEMENS;DTPercentChange | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
HV S Vmax % change | % | 1 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | % Change in Systolic Peak Velocity;99SIEMENS;VMAXsPrcntChange | NUM
HV D Vmax % change | % | 1 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | % Change in Diastolic Peak Velocity;99SIEMENS;VMAXdPrcntChange | NUM
HV A Vmax % change | % | 1 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | % Change in Peak Reversal Velocity during Atrial Contraction;99SIEMENS;AVmaxPrcntChange | NUM
RV PEP/RV ET | unitless | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Pre-Ejection Period/Ejection Time Ratio;LN;59088-5 | NUM
RV AT/RV ET | unitless | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Ratio of Right Ventricular Acceleration Time to Ejection Time;99SIEMENS;RatioRVAT_ET | NUM
PASP | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonary Artery;SRT;T-44000 | Pulmonary Artery Pressure;SRT;F-0212C | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020
PASP/SBP | unitless | 2 | Finding Site;SRT;G-C0E3=Pulmonary Artery;SRT;T-44000 | Systolic Pulmonary Artery to Systolic Blood Pressure Ratio;99SIEMENS;PASP2SBP | NUM
PAEDP | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonary Artery;SRT;T-44000 | Pulmonary Artery Pressure;SRT;F-0212C | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011
Mean PAP | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonary Artery;SRT;T-44000 | Mean Pulmonary Artery Pressure;99SIEMENS;MPAP | NUM
PVR | WU | 2 | Finding Site;SRT;G-C0E3=Pulmonary Artery;SRT;T-44000 | Pulmonary Vascular Resistance;SRT;F-03E86 | NUM
LUPV A Dur - MV A Dur | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulm Vein A duration to MV A duration difference;LN;59114-9 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Upper Segment;SRT;R-40491
LUPV S Vmax / D Vmax | unitless | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Systolic to Diastolic Ratio;LN;29452-0 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Upper Segment;SRT;R-40491
LUPV Systolic Filling Fraction | % | 1 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Systolic Filling Fraction;99SIEMENS;SysFillFrac | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Upper Segment;SRT;R-40491
LLPV A Dur - MV A Dur | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulm Vein A duration to MV A duration difference;LN;59114-9 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV S Vmax / D Vmax | unitless | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Systolic to Diastolic Ratio;LN;29452-0 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV Systolic Filling Fraction | % | 1 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Systolic Filling Fraction;99SIEMENS;SysFillFrac | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
RUPV A Dur - MV A Dur | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulm Vein A duration to MV A duration difference;LN;59114-9 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Upper Segment;SRT;R-404A0
RUPV S Vmax / D Vmax | unitless | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Systolic to Diastolic Ratio;LN;29452-0 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Upper Segment;SRT;R-404A0
RUPV Systolic Filling Fraction | % | 1 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Systolic Filling Fraction;99SIEMENS;SysFillFrac | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Upper Segment;SRT;R-404A0
RLPV A Dur - MV A Dur | ms | 0 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulm Vein A duration to MV A duration difference;LN;59114-9 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV S Vmax / D Vmax | unitless | 2 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Pulmonary Vein Systolic to Diastolic Ratio;LN;29452-0 | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV Systolic Filling Fraction | % | 1 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | Systolic Filling Fraction;99SIEMENS;SysFillFrac | NUM |  |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
IVC Diam % change | % | 1 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | Inferior Vena Cava % Collapse;LN;18050-5 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
IVC Diam change | cm | 2 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | Change in Diameter;99SIEMENS;DiamChange | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
IVC S/D Vmax | unitless | 2 | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM
HV S/D Vmax | unitless | 2 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | Hepatic Vein Systolic to Diastolic Ratio;LN;29473-6 | NUM
HV Systolic Filling Fraction | % | 1 | Finding Site;SRT;G-C0E3=Hepatic Vein;SRT;T-48720 | Systolic Filling Fraction;99SIEMENS;SysFillFrac | NUM
PV AT/ET | unitless | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Ratio of Pulmonic Valve Acceleration Time to Ejection Time;SRT;G-0388 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
RVOT Area | cm² | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVOT SV | ml | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
PV Area (Vmax) | cm² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Peak Velocity;DCM;125214 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV Area/BSA (Vmax) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Peak Velocity;DCM;125214 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6
PV Area (VTI) | cm² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV Area/BSA (VTI) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6
PV Site 2 Area | cm² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
PV Site 2 SV | ml | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PR Volume (Doppler) | ml | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Volume;SRT;G-D705 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR RF (Doppler) | % | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Regurgitant Fraction;SRT;G-0390 | NUM
PR PISA | cm² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Flow Area;LN;20226-7 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR Inst Flow Rate | ml/s | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Peak Instantaneous Flow Rate;LN;34141-2 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR EROA | cm² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR Volume (PISA) | ml | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Volume;SRT;G-D705 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR RF (PISA) | % | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Regurgitant Fraction;SRT;G-0390 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PV Ann s Area | cm² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Cardiac valve annulus;SRT;T-3500E |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
PV SV | ml | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Stroke Volume;SRT;F-32120 | NUM
PV SI | ml/m² | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Stroke Index;SRT;F-00078 | NUM
PV CO | l/min | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiac Output;SRT;F-32100 | NUM
PV CI | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiac Index;SRT;F-32110 | NUM
PV AT/ET (Post) | unitless | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Ratio of Pulmonic Valve Acceleration Time to Ejection Time;SRT;G-0388 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
RVOT Area (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
RVOT SV (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV Area (Vmax) (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Peak Velocity;DCM;125214 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV Area/BSA (Vmax) (Post) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Peak Velocity;DCM;125214 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV Area (VTI) (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV Area/BSA (VTI) (Post) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV Site 2 Area (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV Site 2 SV (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR Volume (Doppler) (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Volume;SRT;G-D705 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR RF (Doppler) (Post) | % | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Regurgitant Fraction;SRT;G-0390 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR PISA (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Flow Area;LN;20226-7 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR Inst Flow Rate (Post) | ml/s | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Peak Instantaneous Flow Rate;LN;34141-2 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR EROA (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR Volume (PISA) (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Volume;SRT;G-D705 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR RF (PISA) (Post) | % | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Regurgitant Fraction;SRT;G-0390 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV Ann s Area (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Cardiac valve annulus;SRT;T-3500E |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV SV (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Stroke Volume;SRT;F-32120 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV SI (Post) | ml/m² | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Stroke Index;SRT;F-00078 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV CO (Post) | l/min | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiac Output;SRT;F-32100 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PV CI (Post) | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiac Index;SRT;F-32110 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
RA Vol d A4C A-L | ml | 1 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RA Vol s A4C A-L | ml | 1 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RA Vol d A4C A-L/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
RA Vol s A4C A-L/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
RA Vol d A4C MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
RA Vol s A4C MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
MV E Vmax % change | % | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | % Change in E-Wave Peak Velocity;99SIEMENS;VMAXePrcntChange | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV A Vmax % change | % | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | % Change in A-Wave Peak Velocity;99SIEMENS;VMAXaPrcntChange | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV E VTI % change | % | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | % Change in E-Wave Velocity Time Integral;99SIEMENS;VTIePrcntChange | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV A VTI % change | % | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | % Change in A-Wave Velocity Time Integral;99SIEMENS;VTIaPrcntChange | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV DT % change | % | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | % Change in Deceleration Time;99SIEMENS;DTPercentChange | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
LV IVRT % change | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | % Change in Isovolumic Relaxation Time;99SIEMENS;IVRTPrcntChange | NUM
P Vein S Vmax % change | % | 1 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | % Change in Systolic Peak Velocity;99SIEMENS;VMAXsPrcntChange | NUM
P Vein D Vmax % change | % | 1 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | % Change in Diastolic Peak Velocity;99SIEMENS;VMAXdPrcntChange | NUM
P Vein A Vmax % change | % | 1 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | % Change in Peak Reversal Velocity during Atrial Contraction;99SIEMENS;AVmaxPrcntChange | NUM
P Vein S VTI % change | % | 1 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | % Change in S-wave Velocity Time Integral;99SIEMENS;SVTIPrcntChange | NUM
P Vein D VTI % change | % | 1 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | % Change in D-Wave Velocity Time Integral;99SIEMENS;DVTIPrcntChange | NUM
P Vein A VTI % change | % | 1 | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 | % Change in A-Wave Velocity Time Integral;99SIEMENS;AVTIPrcntChange | NUM
TV E/A | unitless | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve E to A Ratio;LN;18039-8 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
RVID d/BSA | cm/m² | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Internal Diastolic Dimension;LN;20304-2 | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
RVID d/LVID d | unitless | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right-to-left ventricular Diastolic Diameter Ratio;99SIEMENS;RatioRVID_LVIDd | NUM | Image View;DCM;111031=Parasternal long axis;SRT;G-0396 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV %FAC A4C | % | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Fractional Area Change;99SIEMENS;RVFAC | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV MPI | unitless | 3 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Index of Myocardial Performance;SRT;G-0381 | NUM
RVSP (TR) | mmHg | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Peak Systolic Pressure;SRT;G-0380 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
RVOT/LVOT VTI | unitless | 2 | Right Ventricle Outflow Tract;SRT;T-32550 | Ratio of RVOT VTI to LVOT VTI;99SIEMENS;RVOT2LVOT_VTI | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
RVOT Area PSAX | cm² | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Image View;DCM;111031=Parasternal short axis;SRT;G-0397 | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVOT SV PSAX | ml | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Parasternal short axis;SRT;G-0397 | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT SI PSAX | ml/m² | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Stroke Index;SRT;F-00078 | NUM | Image View;DCM;111031=Parasternal short axis;SRT;G-0397 | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT CO PSAX | l/min | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Cardiac Output;SRT;F-32100 | NUM | Image View;DCM;111031=Parasternal short axis;SRT;G-0397 | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT CI PSAX | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Cardiac Index;SRT;F-32110 | NUM | Image View;DCM;111031=Parasternal short axis;SRT;G-0397 | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT SI | ml/m² | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Stroke Index;SRT;F-00078 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT CO | l/min | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Cardiac Output;SRT;F-32100 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT CI | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Cardiac Index;SRT;F-32110 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RV EF A4C A-L | % | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | RV Ejection Fraction;LN;10231-9 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Area d A4C/BSA | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Area;SRT;G-A166 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
RV Area s A4C/BSA | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Area;SRT;G-A166 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
RV Vol d A4C A-L | ml | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Vol s A4C A-L | ml | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Vol d A4C A-L/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
RV Vol s A4C A-L/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Area-Length Single Plane;DCM;125205 | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
RV EF A4C MOD | % | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | RV Ejection Fraction;LN;10231-9 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Vol d A4C MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
RV Vol s A4C MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks, Single Plane;DCM;125208 | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6
TV A/E | unitless | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve A to E Ratio;99SIEMENS;TVAtoERatio | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Area (PHT) | cm² | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Area by Pressure Half-Time;DCM;125210 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Area (VTI) | cm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Area/BSA (VTI) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6
TV PISA | cm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Flow Area;LN;20226-7 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Inst Flow Rate | l/min | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Instantaneous Flow Rate;LN;34141-2 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Area (PISA) | cm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Area/BSA (PISA) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6
TV Site 2 Area | cm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
TV Site 2 SV | ml | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TR Volume (Doppler) | ml | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Volume;SRT;G-D705 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR RF (Doppler) | % | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Regurgitant Fraction;SRT;G-0390 | NUM
TR Inst Flow Rate | ml/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Instantaneous Flow Rate;LN;34141-2 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR EROA | cm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR Volume (PISA) | ml | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Volume;SRT;G-D705 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR RF (PISA) | % | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Regurgitant Fraction;SRT;G-0390 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TV Ann Area | cm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
TV SV | ml | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Stroke Volume;SRT;F-32120 | NUM
TV SI | ml/m² | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Stroke Index;SRT;F-00078 | NUM
TV CO | l/min | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiac Output;SRT;F-32100 | NUM
TV CI | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiac Index;SRT;F-32110 | NUM
TV E/A (Annulus) | unitless | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve E to A Ratio;LN;18039-8 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV A/E (Annulus) | unitless | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve A to E Ratio;99SIEMENS;TVAtoERatio | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV E/e' (L) | unitless | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Ratio of TV E Vmax to RV Free Wall Ea;99SIEMENS;RatioRVEePrime | NUM |  | Finding Site;SRT;G-C0E3=Lateral Tricuspid Annulus;99SIEMENS;LatTriAnn
TV e'/a' (L) | unitless | 2 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Tricuspid Valve Ea to Aa Ratio;99SIEMENS;CA_TV_Ea2Aa | NUM |  | Finding Site;SRT;G-C0E3=Lateral Tricuspid Annulus;99SIEMENS;LatTriAnn |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
TV E/A (Post) | unitless | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve E to A Ratio;LN;18039-8 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV A/E (Post) | unitless | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Tricuspid Valve A to E Ratio;99SIEMENS;TVAtoERatio | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Area (PHT) (Post) | cm² | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Area by Pressure Half-Time;DCM;125210 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Area (VTI) (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Area/BSA (VTI) (Post) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV PISA (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Flow Area;LN;20226-7 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Inst Flow Rate (Post) | l/min | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Instantaneous Flow Rate;LN;34141-2 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Area (PISA) (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Area/BSA (PISA) (Post) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Site 2 Area (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Site 2 SV (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR Volume (Doppler) (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Volume;SRT;G-D705 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR RF (Doppler) (Post) | % | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Regurgitant Fraction;SRT;G-0390 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR Inst Flow Rate (Post) | ml/s | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Instantaneous Flow Rate;LN;34141-2 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR EROA (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR Volume (PISA) (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Volume;SRT;G-D705 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TR RF (PISA) (Post) | % | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Regurgitant Fraction;SRT;G-0390 | NUM |  |  | Measurement Method;SRT;G-C036=Proximal Isovelocity Surface Area;DCM;125216 |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Ann Area (Post) | cm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV SV (Post) | ml | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Stroke Volume;SRT;F-32120 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV SI (Post) | ml/m² | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Stroke Index;SRT;F-00078 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV CO (Post) | l/min | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiac Output;SRT;F-32100 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV CI (Post) | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiac Index;SRT;F-32110 | NUM | [10 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
HR-LV | bpm | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Heart rate;LN;8867-4 | NUM | [5 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
HR-LV | bpm | 0 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Heart rate;LN;8867-4 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
HR-RV | bpm | 0 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Heart rate;LN;8867-4 | NUM | [5 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
HR-AV | bpm | 0 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Heart rate;LN;8867-4 | NUM | [5 empty] | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2
HR-AV (Post) | bpm | 0 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Heart rate;LN;8867-4 | NUM | [5 empty] | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
HR-MV | bpm | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Heart rate;LN;8867-4 | NUM | [5 empty] | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2
HR-MV (Post) | bpm | 0 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Heart rate;LN;8867-4 | NUM | [5 empty] | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
HR-PV | bpm | 0 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Heart rate;LN;8867-4 | NUM | [5 empty] | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2
HR-PV (Post) | bpm | 0 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Heart rate;LN;8867-4 | NUM | [5 empty] | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
HR-TV | bpm | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Heart rate;LN;8867-4 | NUM | [5 empty] | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2
HR-TV (Post) | bpm | 0 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Heart rate;LN;8867-4 | NUM | [5 empty] | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Asc Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Peak Gradient;LN;20247-3 | NUM
Ao Desc Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Peak Gradient;LN;20247-3 | NUM
Ao Coarct Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Aorta;SRT;T-42000 | Thoracic Aorta Coarctation Systolic Peak Instantaneous Gradient;LN;17995-2 | NUM
AoV Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
LVOT Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Peak Gradient;LN;20247-3 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
AR ED Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Pressure Gradient;SRT;F-023F7 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AoV Peak Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LVOT Peak Grad(Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Peak Gradient;LN;20247-3 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR ED Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Pressure Gradient;SRT;F-023F7 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AR Peak Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
ASD Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Peak Gradient;LN;20247-3 | NUM |  | Finding Site;SRT;G-C0E3=Atrial Septal Defect;SRT;D4-31220
PDA Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Peak Gradient;LN;20247-3 | NUM |  | Finding Site;SRT;G-C0E3=Patent Ductus Arteriosus;SRT;D4-32012
VSD Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 | Peak Gradient;LN;20247-3 | NUM
MV Peak Grad, Annulus | mmHg | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Gradient;LN;20247-3 | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Peak Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
MV Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Gradient;LN;20247-3 | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
LPA Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonary Artery;SRT;T-44000 | Peak Gradient;LN;20247-3 | NUM |  | Finding Site;SRT;G-C0E3=Left Pulmonary Artery;SRT;T-44400
RPA Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonary Artery;SRT;T-44000 | Peak Gradient;LN;20247-3 | NUM |  | Finding Site;SRT;G-C0E3=Right Pulmonary Artery;SRT;T-44200
TR Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
RVOT Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Peak Gradient;LN;20247-3 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
PV Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PR ED Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Pressure Gradient;SRT;F-023F7 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PV Peak Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
RVOT Peak Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Peak Gradient;LN;20247-3 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 | [8 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR ED Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Pressure Gradient;SRT;F-023F7 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
PR Peak Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Peak Grad (Post) | mmHg | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Gradient;LN;20247-3 | NUM | [7 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
TV Peak Grad | mmHg | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Gradient;LN;20247-3 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047

## Valves Measurements
Measurement | EnglishLabel | Precision | MaxInstances | Unit | ConceptCoding | ValueType | ImageView | CardiacCyclePoint | Index | Container | Subcontainer | Finding Site | TopographicalModifier | Additional Modifier
Aortic Valve Analysis | Ann Area | 1 | 5 | mm² | Area;SRT;G-A166 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Aortic Valve Ring;SRT;T-35410 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Ann Perimeter | 1 | 5 | mm | Perimeter;SRT;G-A197 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Aortic Valve Ring;SRT;T-35410 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Ann Min Diam | 1 | 5 | mm | Minimum Diameter;99SIEMENS;MinDiam | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Aortic Valve Ring;SRT;T-35410 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Ann Max Diam | 1 | 5 | mm | Maximum Diameter;99SIEMENS;MaxDiam | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Aortic Valve Ring;SRT;T-35410 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Ann Diam - Area Derived | 1 | 5 | mm | Diameter;SRT;M-02550 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Aortic Valve Ring;SRT;T-35410 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Ann Diam - Perimeter Derived | 1 | 5 | mm | Diameter;SRT;M-02550 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Aortic Valve Ring;SRT;T-35410 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Ann Avg Diam | 1 | 5 | mm | Average Diameter;99SIEMENS;AvgDiam | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Aortic Valve Ring;SRT;T-35410 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Root SoV Area | 1 | 5 | mm² | Vessel lumen cross sectional area;SRT;G-0366 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Structure Sinus of Valsalva;SRT;T-42200 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Root SoV Perimeter | 1 | 5 | mm | Perimeter;SRT;G-A197 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Structure Sinus of Valsalva;SRT;T-42200 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Min SoV Diam | 1 | 5 | mm | Minimum Diameter;99SIEMENS;MinDiam | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Structure Sinus of Valsalva;SRT;T-42200 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Max SoV Diam | 1 | 5 | mm | Maximum Diameter;99SIEMENS;MaxDiam | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Structure Sinus of Valsalva;SRT;T-42200 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Root SoV Diam - Perimeter Derived | 1 | 5 | mm | Vessel lumen diameter;SRT;G-0364 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Structure Sinus of Valsalva;SRT;T-42200 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Root STJ Area | 1 | 5 | mm² | Vessel lumen cross sectional area;SRT;G-0366 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Aortic Sinotubular Junction;SRT;T-42102 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Root STJ Perimeter | 1 | 5 | mm | Perimeter;SRT;G-A197 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Aortic Sinotubular Junction;SRT;T-42102 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Min STJ Diam | 1 | 5 | mm | Minimum Diameter;99SIEMENS;MinDiam | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Aortic Sinotubular Junction;SRT;T-42102 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Max STJ Diam | 1 | 5 | mm | Maximum Diameter;99SIEMENS;MaxDiam | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Aortic Sinotubular Junction;SRT;T-42102 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Root STJ Perimeter | 1 | 5 | mm | Vessel lumen diameter;SRT;G-0364 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Aortic Sinotubular Junction;SRT;T-42102 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Min Root Height | 1 | 5 | mm | Minimum Height;99SIEMENS;MinHeight | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Aortic Root;SRT;F-04403 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Max Root Height | 1 | 5 | mm | Maximum Height;99SIEMENS;MaxHeight | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Aortic Root;SRT;F-04403 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Avg Root Height | 1 | 5 | mm | Average Height;99SIEMENS;AvgHeight | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Finding Site;SRT;G-C0E3=Aortic Root;SRT;F-04403 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| L Ostium Height | 1 | 5 | mm | Coronary Ostium Height;99SIEMENS;CorOstiumHeight | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Topographical modifier;SRT;G-A1F8=Left;SRT;G_A101 | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| R Ostium Height | 1 | 5 | mm | Coronary Ostium Height;99SIEMENS;CorOstiumHeight | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Topographical modifier;SRT;G-A1F8=Right;SRT;G_A100 | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Interostia Angle | 1 | 5 | ° | Aortic Valve Interostia Angle;99SIEMENS;AVIntrostiaAngle | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Valve Orifice Area | 1 | 5 | mm² | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Effective Height | 1 | 5 | mm | Aortic Valve Effective Height;99SIEMENS;AVEffectiveHt | NUM |  | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011 |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| L Leaflet Height | 1 | 5 | mm | Height;DCM;121207 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| R Leaflet Height | 1 | 5 | mm | Height;DCM;121207 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| N Leaflet Height | 1 | 5 | mm | Height;DCM;121207 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Aorto-Mitral Angle | 1 | 5 | ° | AV MV Angle;99SIEMENS;AVMVAngle | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
Mitral Valve Analysis | Ann Area | 1 | 5 | mm² | Area;SRT;G-A166 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Total Ann Perimeter | 1 | 5 | mm | Perimeter;SRT;G-A197 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Ant Ann Perimeter | 1 | 5 | mm | Perimeter;SRT;G-A197 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | Topographical modifier;SRT;G-A1F8=Anterior;SRT;R-404CC | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Post Ann Perimeter | 1 | 5 | mm | Perimeter;SRT;G-A197 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 | Topographical modifier;SRT;G-A1F8=Posterior;SRT;R-404CE | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Ann AP Diam | 1 | 5 | mm | Diameter;SRT;M-02550 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Ann AL-PM Diam | 1 | 5 | mm | Mitral Valve Anterolateral Posteromedial Diameter;99SIEMENS;ALPMDiam | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Ann Height | 1 | 5 | mm | Height;DCM;121207 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Intertrigonal distance | 1 | 5 | mm | Mitral Valve Intertrigonal Distance;99SIEMENS;IntertrigDist | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Ann Height : Commisural Width | 1 | 5 | unitless | Mitral Valve Annulus Height AL PM Commissural Diameter Ratio;99SIEMENS;AHtALPMDiamRatio | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Valve Orifice Area | 1 | 5 | mm² | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Intercomm Dist | 1 | 5 | mm | Mitral Valve Intercommissural Distance;99SIEMENS;InterCommDist | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| A1-P1 Coapt Height | 1 | 5 | mm | Mitral Valve A1 P1 Coaptation Segmental Height;99SIEMENS;A1P1CoaptSegHt | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| A2-P2 Coapt Height | 1 | 5 | mm | Mitral Valve A2 P2 Coaptation Segmental Height;99SIEMENS;A2P2CoaptSegHt | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| A3-P3 Coapt Height | 1 | 5 | mm | Mitral Valve A3 P3 Coaptation Segmental Height;99SIEMENS;A3P3CoaptSegHt | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Avg Leaflet Coapt Height | 1 | 5 | mm | Mitral Valve Coaptation Average Height;99SIEMENS;MVCoaptAvgHt | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| A1-P1 Flail Gap | 1 | 5 | mm | Mitral Valve A1 P1 Flail Segmental Gap;99SIEMENS;A1P1FlailSegGap | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| A2-P2 Flail Gap | 1 | 5 | mm | Mitral Valve A2 P2 Flail Segmental Gap;99SIEMENS;A2P2FlailSegGap | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| A3-P3 Flail Gap | 1 | 5 | mm | Mitral Valve A3 P3 Flail Segmental Gap;99SIEMENS;A3P3FlailSegGap | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Max Flail Gap | 1 | 5 | mm | Mitral Valve Maximum Flail Gap;99SIEMENS;MVMaxFlailGap | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Avg Flail Gap | 1 | 5 | mm | Mitral Valve Average Flail Gap;99SIEMENS;MVAvgFlailGap | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| A1-P1 Tenting Height | 1 | 5 | mm | Mitral Valve A1 P1 Tenting Segmental Height;99SIEMENS;A1P1TentingSegHt | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| A2-P2 Tenting Height | 1 | 5 | mm | Mitral Valve A2 P2 Coaptation Segmental Height;99SIEMENS;A2P2CoaptSegHt | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| A3-P3 Tenting Height | 1 | 5 | mm | Mitral Valve A3 P3 Tenting Segmental Height;99SIEMENS;A3P3TentingSegHt | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Max Tenting Height | 1 | 5 | mm | Mitral Valve Maximum Tenting Height;99SIEMENS;MVMaxTentingHt | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Avg Tenting Height | 1 | 5 | mm | Mitral Valve Average Tenting Height;99SIEMENS;MVAvgTentingHt | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Ant Prolapse Area | 1 | 5 | mm² | Mitral Valve Prolapse Area Anterior;99SIEMENS;MVProlapseAreaAn | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Post Prolapse Area | 1 | 5 | mm² | Mitral Valve Prolapse Area Posterior;99SIEMENS;MVProlapseAreaPo | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Ant Prolapse Volume | 1 | 5 | ml | Mitral Valve Prolapse Volume Anterior;99SIEMENS;MVProlapseVolAn | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Post Prolapse Volume | 1 | 5 | ml | Mitral Valve Prolapse Volume Posterior;99SIEMENS;MVProlapseVolPo | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Max Ant Leaflet Length | 1 | 5 | mm | Mitral Valve Leaflet Maximum Length;99SIEMENS;MVLeafletMaxLen | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Topographical modifier;SRT;G-A1F8=Anterior;SRT;R-404CC | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Avg Ant Leaflet Length | 1 | 5 | mm | Mitral Valve Leaflet Average Length;99SIEMENS;MVLeafletAvgLen | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Topographical modifier;SRT;G-A1F8=Anterior;SRT;R-404CC | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| P1 Length | 1 | 5 | mm | Length;SRT;G-A22A | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=A1 Leaflet Segment;99SIEMENS;A1LeafletSegment |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| P2 Length | 1 | 5 | mm | Length;SRT;G-A22A | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=A2 Leaflet Segment;99SIEMENS;A2LeafletSegment |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| P3 Length | 1 | 5 | mm | Length;SRT;G-A22A | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=A3 Leaflet Segment;99SIEMENS;A3LeafletSegment |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Max Post Leaflet Length | 1 | 5 | mm | Mitral Valve Leaflet Maximum Length;99SIEMENS;MVLeafletMaxLen | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Topographical modifier;SRT;G-A1F8=Posterior;SRT;R-404CE | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Avg Post Leaflet Length | 1 | 5 | mm | Mitral Valve Leaflet Average Length;99SIEMENS;MVLeafletAvgLen | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Topographical modifier;SRT;G-A1F8=Posterior;SRT;R-404CE | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| A1-P1 Coapt Height | 1 | 5 | mm | Length;SRT;G-A22A | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=P1 Leaflet Segment;99SIEMENS;P1LeafletSegment |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| A2-P2 Coapt Height | 1 | 5 | mm | Length;SRT;G-A22A | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=P2 Leaflet Segment;99SIEMENS;P2LeafletSegment |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| A3-P3 Coapt Height | 1 | 5 | mm | Length;SRT;G-A22A | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=P3 Leaflet Segment;99SIEMENS;P3LeafletSegment |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Leaflet Area | 1 | 5 | mm² | Mitral Valve Total Leaflet Area;99SIEMENS;MVTotalLeafArea | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| A1 Length | 1 | 5 | mm² | Area;SRT;G-A166 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=A1 Leaflet Segment;99SIEMENS;A1LeafletSegment |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| A2 Length | 1 | 5 | mm² | Area;SRT;G-A166 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=A2 Leaflet Segment;99SIEMENS;A2LeafletSegment |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| A3 Length | 1 | 5 | mm² | Area;SRT;G-A166 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=A3 Leaflet Segment;99SIEMENS;A3LeafletSegment |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| P1 Area | 1 | 5 | mm² | Area;SRT;G-A166 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=P1 Leaflet Segment;99SIEMENS;P1LeafletSegment |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| P2 Area | 1 | 5 | mm² | Area;SRT;G-A166 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=P2 Leaflet Segment;99SIEMENS;P2LeafletSegment |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| P3 Area | 1 | 5 | mm² | Area;SRT;G-A166 | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Finding Site;SRT;G-C0E3=P3 Leaflet Segment;99SIEMENS;P3LeafletSegment |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES
| Aorto-Mitral Angl | 1 | 5 | ° | AV MV Angle;99SIEMENS;AVMVAngle | NUM |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  |  | Clinical Application Package;99SIEMENS;CAP=Valves;99SIEMENS;VALVES

## Strain Gauge Measurements
MeasType | EnglishLabel | Precision | MaxInstances | Unit | ConceptCoding | ValueType | ImageView | MeasurementMethod | CardiacCyclePoint | Index | Subcontainer | FindingSite | Modifier
Left Ventricle A4C Analysis | HR | 0 | 5 | bpm | Heart Rate;LN;8867-4 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| CO | 1 | 5 | l/min | Cardiac Output;SRT;F-32100 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ED Axis | 1 | 5 | cm | Left Ventricle diastolic major axis;LN;18077-8 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EF | 1 | 5 | % | Left Ventricular Ejection Fraction;LN;18043-0 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDV | 1 | 5 | ml | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDV/BSA | 1 | 5 | ml/m² | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV | 1 | 5 | ml | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV/BSA | 1 | 5 | ml/m² | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ES Axis | 1 | 5 | cm | Left Ventricle systolic major axis;LN;18076-0 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS | 1 | 5 | % | Global Longitudinal Strain;99SIEMENS;GLS | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| Mass | 1 | 5 | g | Left Ventricle Mass;LN;18087-7 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| Mass/BSA | 1 | 5 | g/m² | Left Ventricle Mass;LN;18087-7 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| SV | 1 | 5 | ml | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| SV/BSA | 1 | 5 | ml/m² | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
Left Atrium A4C Analysis | HR | 0 | 5 | bpm | Heart Rate;LN;8867-4 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| CO | 1 | 5 | l/min | Cardiac Output;SRT;F-32100 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ED Axis | 1 | 5 | cm | Major Axis;SRT;G-A193 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011 |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EF | 1 | 5 | % | Left Atrial Ejection Fraction;99SIEMENS;LAEF | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDV | 1 | 5 | ml | Left Atrial ED Volume;DCM;122407 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDV/BSA | 1 | 5 | ml/m² | Left Atrial ED Volume;DCM;122407 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV | 1 | 5 | ml | Left Atrial ES Volume;DCM;122408 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV/BSA | 1 | 5 | ml/m² | Left Atrial ES Volume;DCM;122408 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ES Axis | 1 | 5 | cm | Major Axis;SRT;G-A193 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Cardiac Cycle Point;SRT;R-4089A=End Systole;SRT;R-FAB5B |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS | 1 | 5 | % | Global Longitudinal Strain;99SIEMENS;GLS | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS Res | 1 | 5 | % | Global Longitudinal Strain Reservoir;99SIEMENS;GLSRes | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS Cond | 1 | 5 | % | Global Longitudinal Strain Conduit;99SIEMENS;GLSCond | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS Cont | 1 | 5 | % | Global Longitudinal Strain Contraction;99SIEMENS;GLSCont | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| SV | 1 | 5 | ml | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| SV/BSA | 1 | 5 | ml/m² | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
Right Ventricle A4C Analysis | HR | 0 | 5 | bpm | Heart Rate;LN;8867-4 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ED Axis | 1 | 5 | cm | Major Axis;SRT;G-A193 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011 |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDA | 1 | 5 | cm² | Right Ventricular Diastolic Area;99SIEMENS;RVDA | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDA/BSA | 1 | 5 | cm²/m² | Right Ventricular Diastolic Area;99SIEMENS;RVDA | NUM |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESA | 1 | 5 | cm² | Right Ventricular Systolic Area;99SIEMENS;RVSA | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESA/BSA | 1 | 5 | cm²/m² | Right Ventricular Systolic Area;99SIEMENS;RVSA | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ES Axis | 1 | 5 | cm | Major Axis;SRT;G-A193 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Cardiac Cycle Point;SRT;R-4089A=End Systole;SRT;R-FAB5B |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| FAC | 1 | 5 | % | Right Ventricular Fractional Area Change;99SIEMENS;RVFAC | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS FW | 1 | 5 | % | Global Longitudinal Strain Free Wall;99SIEMENS;GLSFW | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS Sept | 1 | 5 | % | Global Longitudinal Strain Septum;99SIEMENS;GLSSept | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS All | 1 | 5 | % | Global Longitudinal Strain All;99SIEMENS;GLSAll | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| TAPSE | 1 | 5 | cm | Tricuspid Annular Plane Systolic Excursion;99SIEMENS;TAPSE | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ED Base | 1 | 5 | cm | Base Axis;99SIEMENS;BaseAxis | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ES Base | 1 | 5 | cm | Base Axis;99SIEMENS;BaseAxis | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ED Mid | 1 | 5 | cm | Mid Axis;99SIEMENS;MidAxis | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ES Mid | 1 | 5 | cm | Mid Axis;99SIEMENS;MidAxis | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
Right Atrium A4C Analysis | HR | 0 | 5 | bpm | Heart Rate;LN;8867-4 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| CO | 1 | 5 | l/min | Cardiac Output;SRT;F-32100 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ED Axis | 1 | 5 | cm | Major Axis;SRT;G-A193 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011 |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EF | 1 | 5 | % | Cardiac ejection fraction;SRT;F-32070 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDV | 1 | 5 | ml | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDV/BSA | 1 | 5 | ml/m² | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV | 1 | 5 | ml | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV/BSA | 1 | 5 | ml/m² | Volume;SRT;G-D705 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ES Axis | 1 | 5 | cm | Major Axis;SRT;G-A193 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Cardiac Cycle Point;SRT;R-4089A=End Systole;SRT;R-FAB5B |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS | 1 | 5 | % | Global Longitudinal Strain;99SIEMENS;GLS | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| SV | 1 | 5 | ml | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| SV/BSA | 1 | 5 | ml/m² | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDA | 1 | 5 | cm² | Right Atrial Diastolic Area;99SIEMENS;RADA | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDA/BSA | 1 | 5 | cm²/m² | Right Atrial Diastolic Area;99SIEMENS;RADA | NUM |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESA | 1 | 5 | cm² | Right Atrial Systolic Area;99SIEMENS;RASA | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESA/BSA | 1 | 5 | cm²/m² | Right Atrial Systolic Area;99SIEMENS;RASA | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS Res | 1 | 5 | % | Global Longitudinal Strain Reservoir;99SIEMENS;GLSRes | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS Cond | 1 | 5 | % | Global Longitudinal Strain Conduit;99SIEMENS;GLSCond | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS Cont | 1 | 5 | % | Global Longitudinal Strain Contraction;99SIEMENS;GLSCont | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
Left Ventricle A2C Analysis | HR | 0 | 5 | bpm | Heart Rate;LN;8867-4 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| CO | 1 | 5 | l/min | Cardiac Output;SRT;F-32100 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ED Axis | 1 | 5 | cm | Left Ventricle diastolic major axis;LN;18077-8 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EF | 1 | 5 | % | Left Ventricular Ejection Fraction;LN;18043-0 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDV | 1 | 5 | ml | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDV/BSA | 1 | 5 | ml/m² | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV | 1 | 5 | ml | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV/BSA | 1 | 5 | ml/m² | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ES Axis | 1 | 5 | cm | Left Ventricle systolic major axis;LN;18076-0 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS | 1 | 5 | % | Global Longitudinal Strain;99SIEMENS;GLS | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| Mass | 1 | 5 | g | Left Ventricle Mass;LN;18087-7 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| Mass/BSA | 1 | 5 | g/m² | Left Ventricle Mass;LN;18087-7 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| SV | 1 | 5 | ml | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| SV/BSA | 1 | 5 | ml/m² | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
Left Atrium A2C Analysis | HR | 0 | 5 | bpm | Heart Rate;LN;8867-4 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| CO | 1 | 5 | l/min | Cardiac Output;SRT;F-32100 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ED Axis | 1 | 5 | cm | Major Axis;SRT;G-A193 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011 |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EF | 1 | 5 | % | Left Atrial Ejection Fraction;99SIEMENS;LAEF | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDV | 1 | 5 | ml | Left Atrial ED Volume;DCM;122407 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDV/BSA | 1 | 5 | ml/m² | Left Atrial ED Volume;DCM;122407 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV | 1 | 5 | ml | Left Atrial ES Volume;DCM;122408 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV/BSA | 1 | 5 | ml/m² | Left Atrial ES Volume;DCM;122408 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ES Axis | 1 | 5 | cm | Major Axis;SRT;G-A193 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Cardiac Cycle Point;SRT;R-4089A=End Systole;SRT;R-FAB5B |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS | 1 | 5 | % | Global Longitudinal Strain;99SIEMENS;GLS | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS Res | 1 | 5 | % | Global Longitudinal Strain Reservoir;99SIEMENS;GLSRes | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS Cond | 1 | 5 | % | Global Longitudinal Strain Conduit;99SIEMENS;GLSCond | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS Cont | 1 | 5 | % | Global Longitudinal Strain Contraction;99SIEMENS;GLSCont | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| SV | 1 | 5 | ml | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| SV/BSA | 1 | 5 | ml/m² | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
Left Ventricle A3C Analysis | HR | 0 | 5 | bpm | Heart Rate;LN;8867-4 | NUM | Image View;DCM;111031=Apical three chamber;99SIEMENS;A3C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| CO | 1 | 5 | l/min | Cardiac Output;SRT;F-32100 | NUM | Image View;DCM;111031=Apical three chamber;99SIEMENS;A3C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ED Axis | 1 | 5 | cm | Left Ventricle diastolic major axis;LN;18077-8 | NUM | Image View;DCM;111031=Apical three chamber;99SIEMENS;A3C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EF | 1 | 5 | % | Left Ventricular Ejection Fraction;LN;18043-0 | NUM | Image View;DCM;111031=Apical three chamber;99SIEMENS;A3C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDV | 1 | 5 | ml | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical three chamber;99SIEMENS;A3C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDV/BSA | 1 | 5 | ml/m² | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical three chamber;99SIEMENS;A3C |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV | 1 | 5 | ml | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical three chamber;99SIEMENS;A3C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV/BSA | 1 | 5 | ml/m² | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical three chamber;99SIEMENS;A3C |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ES Axis | 1 | 5 | cm | Left Ventricle systolic major axis;LN;18076-0 | NUM | Image View;DCM;111031=Apical three chamber;99SIEMENS;A3C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS | 1 | 5 | % | Global Longitudinal Strain;99SIEMENS;GLS | NUM | Image View;DCM;111031=Apical three chamber;99SIEMENS;A3C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| Mass | 1 | 5 | g | Left Ventricle Mass;LN;18087-7 | NUM | Image View;DCM;111031=Apical three chamber;99SIEMENS;A3C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| Mass/BSA | 1 | 5 | g/m² | Left Ventricle Mass;LN;18087-7 | NUM | Image View;DCM;111031=Apical three chamber;99SIEMENS;A3C |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| SV | 1 | 5 | ml | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Apical three chamber;99SIEMENS;A3C |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| SV/BSA | 1 | 5 | ml/m² | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Apical three chamber;99SIEMENS;A3C |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
Left Ventricle Biplane Analysis | HR | 0 | 5 | bpm | Heart Rate;LN;8867-4 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| CO | 1 | 5 | l/min | Cardiac Output;SRT;F-32100 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ED Axis | 1 | 5 | cm | Left Ventricle diastolic major axis;LN;18077-8 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EF | 1 | 5 | % | Left Ventricular Ejection Fraction;LN;18043-0 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDV | 1 | 5 | ml | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDV/BSA | 1 | 5 | ml/m² | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV | 1 | 5 | ml | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV/BSA | 1 | 5 | ml/m² | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ES Axis | 1 | 5 | cm | Left Ventricle systolic major axis;LN;18076-0 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS | 1 | 5 | % | Global Longitudinal Strain;99SIEMENS;GLS | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| Mass | 1 | 5 | g | Left Ventricle Mass;LN;18087-7 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| Mass/BSA | 1 | 5 | g/m² | Left Ventricle Mass;LN;18087-7 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| SV | 1 | 5 | ml | Stroke Volume;SRT;F-32120 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| SV/BSA | 1 | 5 | ml/m² | Stroke Volume;SRT;F-32120 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
Left Atrium Biplane Analysis | HR | 0 | 5 | bpm | Heart Rate;LN;8867-4 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| CO | 1 | 5 | l/min | Cardiac Output;SRT;F-32100 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ED Axis | 1 | 5 | cm | Major Axis;SRT;G-A193 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011 |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EF | 1 | 5 | % | Left Atrial Ejection Fraction;99SIEMENS;LAEF | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDV | 1 | 5 | ml | Left Atrial ED Volume;DCM;122407 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| EDV/BSA | 1 | 5 | ml/m² | Left Atrial ED Volume;DCM;122407 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV | 1 | 5 | ml | Left Atrial ES Volume;DCM;122408 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV/BSA | 1 | 5 | ml/m² | Left Atrial ES Volume;DCM;122408 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ES Axis | 1 | 5 | cm | Major Axis;SRT;G-A193 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P | Cardiac Cycle Point;SRT;R-4089A=End Systole;SRT;R-FAB5B |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS | 1 | 5 | % | Global Longitudinal Strain;99SIEMENS;GLS | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS Res | 1 | 5 | % | Global Longitudinal Strain Reservoir;99SIEMENS;GLSRes | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS Cond | 1 | 5 | % | Global Longitudinal Strain Conduit;99SIEMENS;GLSCond | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS Cont | 1 | 5 | % | Global Longitudinal Strain Contraction;99SIEMENS;GLSCont | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| SV | 1 | 5 | ml | Stroke Volume;SRT;F-32120 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| SV/BSA | 1 | 5 | ml/m² | Stroke Volume;SRT;F-32120 | NUM |  | Measurement Method;SRT;G-C036=Biplane;99SIEMENS;2P |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
Left Ventricle Summary | AVC | 0 | 5 | ms | Aortic Valve Closure;99SIEMENS;AVC | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| A4C GLS | 1 | 5 | % | Global Longitudinal Strain A4C;99SIEMENS;GLSA4C | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| A3C GLS | 1 | 5 | % | Global Longitudinal Strain A3C;99SIEMENS;GLSA3C | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| A2C GLS | 1 | 5 | % | Global Longitudinal Strain A2C;99SIEMENS;GLSA2C | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| AVG EF | 1 | 5 | % | Averaged Ejection Fraction;99SIEMENS;AvgEF | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| Overall GLS | 1 | 5 | % | Overall Global Longitudinal Strain;99SIEMENS;GLSOverall | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| PSD | 0 | 5 | ms | Peak Strain Dispersion;99SIEMENS;PSD | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
Left Atrium Summary | AVG ESV | 1 | 5 | ml | Averaged End Systolic Volume;99SIEMENS;AvgESV | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| AVG ESV/BSA | 1 | 5 | ml/m² | Averaged End Systolic Volume;99SIEMENS;AvgESV | NUM |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| AVG GLS Cond | 1 | 5 | % | Averaged Global Longitudinal Strain Conduit;99SIEMENS;AvgGLSCond | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| AVG GLS Cont | 1 | 5 | % | Averaged Global Longitudinal Strain Contraction;99SIEMENS;AvgGLSCont | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| AVG GLS Res | 1 | 5 | % | Averaged Global Longitudinal Strain Reservoir;99SIEMENS;AvgGLSRes | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| PreA | 0 | 5 | ms | PreA Time;99SIEMENS;PreA | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
Right Atrium Summary | ESV | 1 | 5 | ml | Right Atrial ES Volume;DCM;122408 | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| ESV/BSA | 1 | 5 | ml/m² | Right Atrial ES Volume;DCM;122408 | NUM |  |  |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS Cond | 1 | 5 | % | Global Longitudinal Strain Conduit;99SIEMENS;GLSCond | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS Cont | 1 | 5 | % | Global Longitudinal Strain Contraction;99SIEMENS;GLSCont | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS Res | 1 | 5 | % | Global Longitudinal Strain Reservoir;99SIEMENS;GLSRes | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| PreA | 0 | 5 | ms | PreA Time;99SIEMENS;PreA | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
Right Ventricle Summary | GLS All | 1 | 5 | % | Global Longitudinal Strain All;99SIEMENS;GLSAll | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| GLS FW | 1 | 5 | % | Global Longitudinal Strain Free Wall;99SIEMENS;GLSFW | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| FAC | 1 | 5 | % | Right Ventricular Fractional Area Change;99SIEMENS;RVFAC | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| TAPSE | 1 | 5 | cm | Right Ventricular Fractional Area Change;99SIEMENS;RVFAC | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG
| PVC | 0 | 5 | ms | Pulmonary Valve Closure;99SIEMENS;PVC | NUM |  |  |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Strain Gauge;99SIEMENS;SG

## Unified Heart Measurements
Measurement | EnglishLabel | Precision | MaxInstances | Unit | ConceptCoding | ValueType | CardiacCyclePoint | Index | Subcontainer | Finding Site | Modifier
Left Ventricle Analysis | HR | 0 | 5 | bpm | Heart Rate;LN;8867-4 | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| CO | 1 | 5 | l/min | Cardiac Output;SRT;F-32100 | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| EF | 1 | 5 | % | Left Ventricular Ejection Fraction;LN;18043-0 | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| EDV | 1 | 5 | ml | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| EDV/BSA | 1 | 5 | ml/m² | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| ESV | 1 | 5 | ml | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| ESV/BSA | 1 | 5 | ml/m² | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| GLS | 1 | 5 | % | Global Longitudinal Strain;99SIEMENS;GLS | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| Mass | 1 | 5 | g | Left Ventricle Mass;LN;18087-7 | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| Mass/BSA | 1 | 5 | g/m² | Left Ventricle Mass;LN;18087-7 | NUM |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| SV | 1 | 5 | ml | Stroke Volume;SRT;F-32120 | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| SV/BSA | 1 | 5 | ml/m² | Stroke Volume;SRT;F-32120 | NUM |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| AVC | 0 | 5 | ms | Aortic Valve Closure;99SIEMENS;AVC | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
Left Atrium Analysis | HR | 0 | 5 | bpm | Heart Rate;LN;8867-4 | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| EF | 1 | 5 | % | Left Atrial Ejection Fraction;99SIEMENS;LAEF | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| EDV | 1 | 5 | ml | Left Atrial ED Volume;DCM;122407 | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| EDV/BSA | 1 | 5 | ml/m² | Left Atrial ED Volume;DCM;122407 | NUM |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| ESV | 1 | 5 | ml | Left Atrial ES Volume;DCM;122408 | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| ESV/BSA | 1 | 5 | ml/m² | Left Atrial ES Volume;DCM;122408 | NUM |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| GLS | 1 | 5 | % | Global Longitudinal Strain;99SIEMENS;GLS | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| GLS Res | 1 | 5 | % | Global Longitudinal Strain Reservoir;99SIEMENS;GLSRes | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| GLS Cond | 1 | 5 | % | Global Longitudinal Strain Conduit;99SIEMENS;GLSCond | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| GLS Cont | 1 | 5 | % | Global Longitudinal Strain Contraction;99SIEMENS;GLSCont | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| PreA | 0 | 5 | ms | PreA Time;99SIEMENS;PreA | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
Right Ventricle Analysis | HR | 0 | 5 | bpm | Heart Rate;LN;8867-4 | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| EF | 1 | 5 | % | Cardiac ejection fraction;SRT;F-32070 | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| EDV | 1 | 5 | ml | Volume;SRT;G-D705 | NUM | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| EDV/BSA | 1 | 5 | ml/m² | Volume;SRT;G-D705 | NUM | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| ESV | 1 | 5 | ml | Volume;SRT;G-D705 | NUM | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| ESV/BSA | 1 | 5 | ml/m² | Volume;SRT;G-D705 | NUM | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| GLS | 1 | 5 | % | Global Longitudinal Strain;99SIEMENS;GLS | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| SV | 1 | 5 | ml | Stroke Volume;SRT;F-32120 | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| SV/BSA | 1 | 5 | ml/m² | Stroke Volume;SRT;F-32120 | NUM |  | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
Right Atrium Analysis | HR | 0 | 5 | bpm | Heart Rate;LN;8867-4 | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| EF | 1 | 5 | % | Cardiac ejection fraction;SRT;F-32070 | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| EDV | 1 | 5 | ml | Volume;SRT;G-D705 | NUM | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| EDV/BSA | 1 | 5 | ml/m² | Volume;SRT;G-D705 | NUM | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| ESV | 1 | 5 | ml | Volume;SRT;G-D705 | NUM | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| ESV/BSA | 1 | 5 | ml/m² | Volume;SRT;G-D705 | NUM | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Index;DCM;121425=Body Surface Area;LN;8277-6 | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH
| GLS | 1 | 5 | % | Global Longitudinal Strain;99SIEMENS;GLS | NUM |  |  | Findings (DCM, 121070) | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 | Clinical Application Package;99SIEMENS;CAP=Unified Heart;99SIEMENS;UH

## Stress Echo Measurements
English Label | MaxInstances | Precision | Possible Units | Parent Container | Concept Coding | Value Type | Image view | Finding Site | Measurement Method | Cardiac Cycle Point | Image Mode | Flow Direction | Additional Modifier
AoV VTI | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Velocity Time Integral;LN;20354-7 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
AoV Mean Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Mean Gradient;LN;20256-4 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
AoV Vmax | 5 | 2 | m/s cm/s | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Peak Systolic Velocity;LN;11726-7 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
AoV Peak Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Peak Gradient;LN;20247-3 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
AoV Site 2 VTI | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
LVOT VTI | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 |  |  |  |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
LVOT Mean Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Mean Gradient;LN;20256-4 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 |  |  |  |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
LVOT Vmax | 5 | 2 | m/s cm/s | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Peak Systolic Velocity;LN;11726-7 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 |  |  |  |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
LVOT Peak Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Peak Gradient;LN;20247-3 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 |  |  |  |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
HR-AV | 5 | 0 | bpm | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Heart_Rate_LN_8867_4 | NUM |  |  |  |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2 |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
AR VTI | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Velocity Time Integral;LN;20354-7 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
AR Mean Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Mean Gradient;LN;20256-4 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
AR Vmax | 5 | 2 | m/s cm/s | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Peak Velocity;LN;20351-3 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
AR Peak Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Peak Gradient;LN;20247-3 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
AR Slope | 5 | 2 | m/s² | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Deceleration Slope;LN;20216-8 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
AR PHT | 5 | 0 | ms | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Pressure Half-Time;LN;20280-4 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
AR DT | 5 | 0 | ms | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Deceleration Time;LN;20217-6 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
MV VTI | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Velocity Time Integral;LN;20354-7 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
MV Mean Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mean Gradient;LN;20256-4 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
MV Site 2 VTI | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
MR Vmax | 5 | 2 | m/s cm/s | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Velocity;LN;20351-3 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
MR Peak Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Gradient;LN;20247-3 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
MR VTI | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Velocity Time Integral;LN;20354-7 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
MR Mean Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mean Gradient;LN;20256-4 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
MV VTI, Annulus | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
MV Vmax, Annulus | 5 | 2 | m/s cm/s | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Velocity;LN;20351-3 | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
MV Peak Grad, Annulus | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Gradient;LN;20247-3 | NUM |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TV VTI | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Velocity Time Integral;LN;20354-7 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TV Mean Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Mean Gradient;LN;20256-4 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TV Decel Slope | 5 | 2 | m/s² | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Deceleration Slope;LN;20216-8 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TV DT | 5 | 0 | ms | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Deceleration Time;LN;20217-6 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TV PHT | 5 | 0 | ms | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Pressure Half-Time;LN;20280-4 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TV Area (PHT) | 5 | 0 | cm² mm² | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Area by Pressure Half-Time;DCM;125210 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TR Vmax | 5 | 2 | m/s cm/s | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Velocity;LN;20351-3 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TR Peak Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Gradient;LN;20247-3 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TR VTI | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Velocity Time Integral;LN;20354-7 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TR Mean Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Mean Gradient;LN;20256-4 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TV Site 2 VTI | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TV VTI, Annulus | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TV Vmax, Annulus | 5 | 2 | m/s cm/s | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Velocity;LN;20351-3 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TV Peak Grad, Annulus | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Gradient;LN;20247-3 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TV Mean Grad, Annulus | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Mean Gradient;LN;20256-4 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
PV Site 2 VTI | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
PV Vmax | 5 | 2 | m/s cm/s | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Peak Systolic Velocity;LN;11726-7 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
PV Peak Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Peak Gradient;LN;20247-3 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
PV VTI | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Velocity Time Integral;LN;20354-7 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
PV Mean Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Mean Gradient;LN;20256-4 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
RVOT VTI | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Velocity Time Integral;LN;20354-7 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 |  |  |  |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
RVOT Mean Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Mean Gradient;LN;20256-4 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 |  |  |  |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
RVOT Vmax | 5 | 2 | m/s cm/s | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Peak Systolic Velocity;LN;11726-7 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 |  |  |  |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
RVOT Peak Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Peak Gradient;LN;20247-3 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 |  |  |  |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
PR Slope | 5 | 2 | m/s² | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Deceleration Slope;LN;20216-8 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
PR PHT | 5 | 0 | ms | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Pressure Half-Time;LN;20280-4 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
PR DT | 5 | 0 | ms | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Deceleration Time;LN;20217-6 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
PR ED Velocity | 5 | 2 | m/s cm/s | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | End Diastolic Velocity;LN;11653-3 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
PR ED Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Pressure Gradient;SRT;F-023F7 | NUM |  |  |  | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011 |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
LV Area d A4C | 5 | 2 | cm² mm² | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Diastolic Area;SRT;G-0375 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
LV Major d A4C | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
LV Vol d A4C MOD | 5 | 1 | ml | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
LV Area s A4C | 5 | 2 | cm² mm² | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Systolic Area;SRT;G-0374 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
LV Major s A4C | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
LV Vol s A4C MOD | 5 | 1 | ml | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
LV Area d A2C | 5 | 2 | cm² mm² | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Diastolic Area;SRT;G-0375 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
LV Major d A2C | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
LV Vol d A2C MOD | 5 | 1 | ml | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
LV Area s A2C | 5 | 2 | cm² mm² | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Systolic Area;SRT;G-0374 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
LV Major s A2C | 5 | 2 | cm mm | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
LV Vol s A2C MOD | 5 | 1 | ml | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
HR-LV | 5 | 0 | bpm | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Heart_Rate_LN_8867_4 | NUM |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 |  | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
MV Vmax | 5 | 2 | m/s cm/s | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Velocity;LN;20351-3 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
MV Peak Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Peak Gradient;LN;20247-3 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TV Vmax | 5 | 2 | m/s cm/s | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Velocity;LN;20351-3 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
TV Peak Grad | 5 | 1 | mmHg | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Peak Gradient;LN;20247-3 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
MV DT | 5 | 0 | ms | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Mitral Valve E-Wave Deceleration Time;SRT;G-0384 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047 | Procedure reported;DCM;121058=Stress Echo;99SIEMENS;StressEcho
Stage Number | Subject ID;DCM;121030
Stage Name | Stage;LN;18139-6

## Stress Echo Calculations
English Label | Unit | Precision | Parent Container | Concept Coding | Value Type | Image view | Finding Site | Measurement Method | Cardiac Cycle Point | Image Mode | Flow Direction
LVOT/AoV Vmax | unitless | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Ratio of LVOT Maximum velocity to Aortic Valve Maximum velocity;99SIEMENS;RatioVmaxLVOT_AV | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
LVOT/AoV VTI | unitless | 3 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Ratio of LVOT VTI to Aortic Valve VTI;99SIEMENS;RatioVTILVOT_AV | NUM | [5 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
LV EF A4C MOD | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Ejection Fraction;LN;18043-0 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d A4C MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol s A4C MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV SV A4C MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV SI A4C MOD | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Index;SRT;F-00078 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CO A4C MOD | l/min | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Output;SRT;F-32100 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CI A4C MOD | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Index;SRT;F-32110 | NUM | Image View;DCM;111031=Apical four chamber;SRT;G-A19C |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV EF A2C MOD | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Ejection Fraction;LN;18043-0 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d A2C MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol s A2C MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV SV A2C MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Volume;SRT;F-32120 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV SI A2C MOD | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Index;SRT;F-00078 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CO A2C MOD | l/min | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Output;SRT;F-32100 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CI A2C MOD | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Index;SRT;F-32110 | NUM | Image View;DCM;111031=Apical two chamber;SRT;G-A19B |  | Measurement Method;SRT;G-C036=Method of Disks Single Plane;DCM;125208 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV EF Biplane MOD | % | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular Ejection Fraction;LN;18043-0 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks Biplane;DCM;125207 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d Biplane MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks Biplane;DCM;125207 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol s Biplane MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks Biplane;DCM;125207 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d Biplane MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Diastolic Volume;LN;18026-5 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks Biplane;DCM;125207 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol s Biplane MOD/BSA | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Left Ventricular End Systolic Volume;LN;18148-7 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks Biplane;DCM;125207 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV SV Biplane MOD | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Volume;SRT;F-32120 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks Biplane;DCM;125207 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV SI Biplane MOD | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Index;SRT;F-00078 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks Biplane;DCM;125207 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CO Biplane MOD | l/min | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Output;SRT;F-32100 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks Biplane;DCM;125207 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CI Biplane MOD | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Index;SRT;F-32110 | NUM |  |  | Measurement Method;SRT;G-C036=Method of Disks Biplane;DCM;125207 |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVSP (TR) | mmHg | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Right Ventricular Peak Systolic Pressure;SRT;G-0380 | NUM |  | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
LVOT SV | ml | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
AVA (Vmax) | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Peak Velocity;DCM;125214 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AVA (Vmean) | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Mean Velocity;DCM;125213 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AVA (VTI) | cm² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AVA/BSA (Vmax) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Peak Velocity;DCM;125214 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AVA/BSA (Vmean) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Mean Velocity;DCM;125213 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AVA/BSA (VTI) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AoV Site 2 SV | ml | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR Volume (Doppler) | ml | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Volume;SRT;G-D705 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR RF (Doppler) | % | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Regurgitant Fraction;SRT;G-0390 | NUM
AoV SV | ml | 1 | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 | Stroke Volume;SRT;F-32120 | NUM
LVOT SI | ml/m² | 1 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Stroke Index;SRT;F-00078 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT CO | l/min | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Output;SRT;F-32100 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT CI | l/min/m² | 2 | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 | Cardiac Index;SRT;F-32110 | NUM |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
MV Area (VTI) | cm² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Area/BSA (VTI) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Site 2 SV | ml | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR Volume (Doppler) | ml | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Volume;SRT;G-D705 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR RF (Doppler) | % | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Regurgitant Fraction;SRT;G-0390 | NUM
MV SV | ml | 1 | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 | Stroke Volume;SRT;F-32120 | NUM
RVOT SV | ml | 1 | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
PV Area (Vmax) | cm² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Peak Velocity;DCM;125214 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV Area/BSA (Vmax) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Peak Velocity;DCM;125214 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV Area (VTI) | cm² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV Area/BSA (VTI) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV Site 2 SV | ml | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PR Volume (Doppler) | ml | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Volume;SRT;G-D705 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR RF (Doppler) | % | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Regurgitant Fraction;SRT;G-0390 | NUM
PV SV | ml | 1 | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 | Stroke Volume;SRT;F-32120 | NUM
TV Area (VTI) | cm² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Area/BSA (VTI) | cm²/m² | 2 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Cardiovascular Orifice Area;SRT;G-038E | NUM |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Site 2 SV | ml | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Stroke Volume;SRT;F-32120 | NUM |  | Finding Site;SRT;G-C0E3=Site 2;99SIEMENS;SITE2 |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TR Volume (Doppler) | ml | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Volume;SRT;G-D705 | NUM | [5 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR RF (Doppler) | % | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Regurgitant Fraction;SRT;G-0390 | NUM
TV SV | ml | 1 | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 | Stroke Volume;SRT;F-32120 | NUM
|  |  | Stage Number | Subject ID;DCM;121030
|  |  | Stage Name | Stage;LN;18139-6

## Ped & Fetal Echo Measurements
EnglishLabel | Description | DetailEnglishLabel | PossibleUnits | MaxInstances | ConceptCoding | ValueType | ParentContainer | ParentLaterality | SRSection | AdditionalModifiers | MeasurementMethod | FindingSite | TopographicalModifier | FlowDirection | CardiacCyclePoint | ImageMode | ImageView
Fetal HR | Fetal Heart Rate |  | bpm | 5 | Heart Rate;LN;8867-4 | NUM | Fetus Characteristics;DCM;125015 |  | Fetus Characteristics;DCM;125015
Atrial FHR | Atrial Fetal Heart Rate |  | bpm | 5 | Heart Rate;LN;8867-4 | NUM | Fetus Characteristics;DCM;125015 |  | Fetus Characteristics;DCM;125015 |  |  | Finding Site;SRT;G-C0E3=Atrium;SRT;T-32101
Vent FHR | Ventricular Fetal Heart Rate |  | bpm | 5 | Heart Rate;LN;8867-4 | NUM | Fetus Characteristics;DCM;125015 |  | Fetus Characteristics;DCM;125015 |  |  | Finding Site;SRT;G-C0E3=Cardiac ventricle;SRT;T-32401
LVOT FHR | Left Ventricular Outflow Tract Fetal Heart Rate |  | bpm | 5 | Heart Rate;LN;8867-4 | NUM | Fetus Characteristics;DCM;125015 |  | Fetus Characteristics;DCM;125015 |  |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
RVOT FHR | Right Ventricular Outflow Tract Fetal Heart Rate |  | bpm | 5 | Heart Rate;LN;8867-4 | NUM | Fetus Characteristics;DCM;125015 |  | Fetus Characteristics;DCM;125015 |  |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
Pulm side FHR | Pulmonic Fetal Heart Rate |  | bpm | 5 | Heart Rate;LN;8867-4 | NUM | Fetus Characteristics;DCM;125015 |  | Fetus Characteristics;DCM;125015 |  |  | Finding Site;SRT;G-C0E3=Pulmonary_side_99SIEMENS_PulmonarySide
Sys Side FHR | Systemic Fetal Heart Rate |  | bpm | 5 | Heart Rate;LN;8867-4 | NUM | Fetus Characteristics;DCM;125015 |  | Fetus Characteristics;DCM;125015 |  |  | Finding Site;SRT;G-C0E3=Systemic_side_99SIEMENS_SystemicSide
AC | Abdominal Circumference |  | cm mm | 5 | Abdominal Circumference;LN;11979-2 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
APTD | Anterior-Posterior Trunk Diameter |  | cm mm | 5 | Anterior-Posterior Trunk Diameter;LN;11819-0 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
BPD | Biparietal Diameter |  | cm mm | 5 | Biparietal Diameter;LN;11820-8 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
Ductus V A Vmax | Ductus Venosus A-wave Vmax | Vmax | cm/s m/s | 5 | A-Wave Peak Velocity;LN;59081-0 | NUM | Finding Site;SRT;G-C0E3=Ductus venosus;SRT;T-F6806 |  | Fetal Measurements;DCM;125016
Ductus V A Vmax | Doppler Angle Property of Ductus Venosus A-wave Vmax | Angle | ° | 5 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Ductus V D Vmax | Ductus Venosus D-wave Vmax | Vmax | cm/s m/s | 5 | Diastolic blood velocity, peak;DCM;122202 | NUM | Finding Site;SRT;G-C0E3=Ductus venosus;SRT;T-F6806 |  | Fetal Measurements;DCM;125016
Ductus V D Vmax | Doppler Angle Property of Ductus Venosus D-wave Vmax | Angle | ° | 5 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Ductus V S Vmax | Ductus Venosus S-wave Vmax | Vmax | cm/s m/s | 5 | Systolic blood velocity, peak;DCM;122204 | NUM | Finding Site;SRT;G-C0E3=Ductus venosus;SRT;T-F6806 |  | Fetal Measurements;DCM;125016
Ductus V S Vmax | Doppler Angle Property of Ductus Venosus S-wave Vmax | Angle | ° | 5 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
FL | Femur Length |  | cm mm | 5 | Femur Length;LN;11963-6 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
Ao Abd Diam | Abdominal Aorta Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Abdominal Aorta;SRT;T-42500 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Arch Diam | Aortic Arch Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Aortic Arch;SRT;T-42300 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Arch Trans Diam | Aortic Arch Transverse Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Aortic Arch;SRT;T-42300 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Transverse Aortic Arch;SRT;T-42304 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Arch | Aortic Arch Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Aortic Arch;SRT;T-42300 |  | Fetal Measurements;DCM;125016
Ao Arch | Aortic Arch Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Aortic Arch;SRT;T-42300 |  | Fetal Measurements;DCM;125016
Ao Coarct AT | Aortic Coarctation Acceleration Time |  | ms | 5 | Acceleration Time;LN;20168-1 | NUM | Finding Site;SRT;G-C0E3=Coarctation of the Aorta;SRT;D4-32014 |  | Fetal Measurements;DCM;125016
Ao Coarct Diam | Aortic Coarctation Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Coarctation of the Aorta;SRT;D4-32014 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Coarct | Aortic Coarctation Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Coarctation of the Aorta;SRT;D4-32014 |  | Fetal Measurements;DCM;125016
Ao Coarct | Aortic Coarctation Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Coarctation of the Aorta;SRT;D4-32014 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod
Ao Coarct | Aortic Coarctation Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Coarctation of the Aorta;SRT;D4-32014 |  | Fetal Measurements;DCM;125016
Ao Coarct | Aortic Coarctation Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Coarctation of the Aorta;SRT;D4-32014 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod
Ao Coarct | Aortic Coarctation Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Coarctation of the Aorta;SRT;D4-32014 |  | Fetal Measurements;DCM;125016
Ao Coarct | Aortic Coarctation Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Coarctation of the Aorta;SRT;D4-32014 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod
Ao Coarct | Aortic Coarctation Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Coarctation of the Aorta;SRT;D4-32014 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod
Ao Coarct ET | Aortic Coarctation Ejection Time |  | ms | 5 | Ejection Time;LN;20222-6 | NUM | Finding Site;SRT;G-C0E3=Coarctation of the Aorta;SRT;D4-32014 |  | Fetal Measurements;DCM;125016
Ao Isthmus 3VV Diam | Aortic Isthmus 3 Vessel View Diameter | Diam | mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Aortic Isthmus;SRT;T-42310 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 | Image View;DCM;111031=3-vessel view;99SIEMENS;3VesselView
Ao Isthmus Sagittal Diam | Aortic Isthmus Sagittal Diameter | Diam | mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Aortic Isthmus;SRT;T-42310 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 | Image View;DCM;111031=Suprasternal sagittal view;SRT;R-40B01
AR | Aortic Regurgitation Velocity Time Integral | VTI | cm mm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR | Aortic Regurgitation Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR | Aortic Regurgitation Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR | Aortic Regurgitation Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR | Aortic Regurgitation Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR Decel | Aortic Regurgitation Deceleration Slope | Slope | cm/s² m/s² | 5 | Deceleration Slope;LN;20216-8 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR Decel | Aortic Regurgitation Deceleration Time | Time | ms | 5 | Deceleration Time;LN;20217-6 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR Decel | Aortic Regurgitation Deceleration Velocity 1 | V1 | cm/s m/s | 5 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR Decel | Aortic Regurgitation Deceleration Velocity 2 | V2 | cm/s m/s | 5 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR ED | Aortic Regurgitation End Diastolic Velocity | Vel | cm/s m/s | 5 | End Diastolic Velocity;LN;11653-3 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
AR ED | Aortic Regurgitation End Diastolic Gradient | Grad | mmHg | 5 | Pressure Gradient;SRT;F-023F7 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011
Ao Root Diam | Aortic Root Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Root of Aorta;SRT;T-42110 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
AoV Annulus Circum s | Aortic Valve Annulus Circumference systole |  | cm mm | 5 | Circumference;SRT;M-02560 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Cardiac valve annulus;SRT;T-3500E |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
AoV Annulus Diam s | Aortic Valve Annulus Diameter systole |  | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Cardiac valve annulus;SRT;T-3500E |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
AoV ET | Aortic Valve Ejection Time |  | ms | 5 | Ejection Time;LN;20222-6 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016
AoV Accel | Aortic Valve Acceleration Slope | Slope | cm/s² m/s² | 5 | Acceleration Slope;LN;20167-3 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016
AoV Accel | Aortic Valve Acceleration Time | Time | ms | 5 | Acceleration Time;LN;20168-1 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016
AoV Cusp Sep | Aortic Valve Cusp Separation |  | cm mm | 5 | Aortic Valve Cusp Separation;LN;17996-0 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
AoV | Aortic Valve Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AoV | Aortic Valve Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AoV | Aortic Valve Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AoV | Aortic Valve Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AoV | Aortic Valve Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AoV | Aortic Valve Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AoV | Aortic Valve Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
Ao Asc Diam | Ascending Aorta Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Ascending aorta;SRT;T-42100 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Asc | Ascending Aorta Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Ascending aorta;SRT;T-42100 |  | Fetal Measurements;DCM;125016
Ao Asc | Ascending Aorta Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Ascending aorta;SRT;T-42100 |  | Fetal Measurements;DCM;125016
ASD | Atrial Septal Defect Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Atrial Septal Defect;SRT;D4-31220 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
ASD | Atrial Septal Defect Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Atrial Septal Defect;SRT;D4-31220 |  | Fetal Measurements;DCM;125016
ASD | Atrial Septal Defect Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Atrial Septal Defect;SRT;D4-31220 |  | Fetal Measurements;DCM;125016
A Septum Length | Atrial Septum Length |  | cm mm | 5 | Length;SRT;G-A22A | NUM | Finding Site;SRT;G-C0E3=Interatrial Septum Structure;SRT;T-32150 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
AVSD | Atrioventricular Septal Defect Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Common atrioventricular-type ventricular septal defect;SRT;D4-31042 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
AVSD | Atrioventricular Septal Defect Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Common atrioventricular-type ventricular septal defect;SRT;D4-31042 |  | Fetal Measurements;DCM;125016
AVSD | Atrioventricular Septal Defect Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Common atrioventricular-type ventricular septal defect;SRT;D4-31042 |  | Fetal Measurements;DCM;125016
Cardiac Axis | Cardiac Axis |  | ° | 5 | Cardiac axis;SRT;F-0258C | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
Cor Sinus | Coronary Sinus Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Coronary sinus;SRT;T-48410 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Cor Sinus Length | Coronary Sinus Length |  | cm mm | 5 | Length;SRT;G-A22A | NUM | Finding Site;SRT;G-C0E3=Coronary sinus;SRT;T-48410 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Desc AT | Descending Aorta Acceleration Time |  | ms | 5 | Acceleration Time;LN;20168-1 | NUM | Finding Site;SRT;G-C0E3=Descending Aorta;SRT;T-D0765 |  | Fetal Measurements;DCM;125016
Ao Desc Diam | Descending Aorta Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Descending Aorta;SRT;T-D0765 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Desc ET | Descending Aorta Ejection Time |  | ms | 5 | Ejection Time;LN;20222-6 | NUM | Finding Site;SRT;G-C0E3=Descending Aorta;SRT;T-D0765 |  | Fetal Measurements;DCM;125016
Ao Desc | Descending Aorta Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Descending Aorta;SRT;T-D0765 |  | Fetal Measurements;DCM;125016
Ao Desc | Descending Aorta Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Descending Aorta;SRT;T-D0765 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod
Ao Desc | Descending Aorta Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Descending Aorta;SRT;T-D0765 |  | Fetal Measurements;DCM;125016
Ao Desc | Descending Aorta Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Descending Aorta;SRT;T-D0765 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod
Ao Desc | Descending Aorta Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Descending Aorta;SRT;T-D0765 |  | Fetal Measurements;DCM;125016
Ao Desc | Descending Aorta Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Descending Aorta;SRT;T-D0765 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod
Ao Desc | Descending Aorta Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Descending Aorta;SRT;T-D0765 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod
Ductus A | Ductus Arteriosus Peak Systolic Velocity | PS | cm/s m/s | 5 | Peak Systolic Velocity;LN;11726-7 | NUM | Finding Site;SRT;G-C0E3=Ductus arteriosus;SRT;T-F6845 |  | Fetal Measurements;DCM;125016
Ductus A | Ductus Arteriosus End Diastoic Velocity | ED | cm/s m/s | 5 | End Diastolic Velocity;LN;11653-3 | NUM | Finding Site;SRT;G-C0E3=Ductus arteriosus;SRT;T-F6845 |  | Fetal Measurements;DCM;125016
Ductus A | Ductus Arteriosus Pulsatility Index | PI | unitless | 5 | Pulsatility Index;LN;12008-9 | NUM | Finding Site;SRT;G-C0E3=Ductus arteriosus;SRT;T-F6845 |  | Fetal Measurements;DCM;125016
Ductus A | Ductus Arteriosus Resistivity Index | RI | unitless | 5 | Resistivity Index;LN;12023-8 | NUM | Finding Site;SRT;G-C0E3=Ductus arteriosus;SRT;T-F6845 |  | Fetal Measurements;DCM;125016
Ductus A | Ductus Arteriosus Systolic to Diastolic Ratio | S/D | unitless | 5 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Finding Site;SRT;G-C0E3=Ductus arteriosus;SRT;T-F6845 |  | Fetal Measurements;DCM;125016
Ductus A | Doppler Angle Property of PSV, EDV | Angle | ° | 5 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Ductus A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | Time averaged peak velocity;LN;11692-1 | NUM | Finding Site;SRT;G-C0E3=Ductus arteriosus;SRT;T-F6845 |  | Fetal Measurements;DCM;125016
Ductus A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | Time averaged mean velocity;LN;20352-1 | NUM | Finding Site;SRT;G-C0E3=Ductus arteriosus;SRT;T-F6845 |  | Fetal Measurements;DCM;125016
Ductus A 3VV | Ductus Arteriosus 3 Vessel View Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Ductus arteriosus;SRT;T-F6845 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2 | Image View;DCM;111031=3-vessel view;99SIEMENS;3VesselView
Ductus V | Ductus Venosus Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Ductus venosus;SRT;T-F6806 |  | Fetal Measurements;DCM;125016
Ductus V | Ductus Venosus Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Ductus venosus;SRT;T-F6806 |  | Fetal Measurements;DCM;125016
Ductus V | Ductus Venosus Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Ductus venosus;SRT;T-F6806 |  | Fetal Measurements;DCM;125016
Ductus V | Ductus Venosus Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Ductus venosus;SRT;T-F6806 |  | Fetal Measurements;DCM;125016
Ductus V | Ductus Venosus Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Ductus venosus;SRT;T-F6806 |  | Fetal Measurements;DCM;125016
IVC | Inferior Vena Cava Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
IVC | Inferior Vena Cava Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 |  | Fetal Measurements;DCM;125016
IVC | Inferior Vena Cava Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Inferior Vena Cava;SRT;T-48710 |  | Fetal Measurements;DCM;125016
IVS d | Interventricular Septal Thickness diastole |  | cm mm | 5 | ROI Thickness by US;LN;59089-3 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Interventricular Septum;SRT;T-32410 |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
IVS d | Interventricular Septal Thickness diastole (M-mode) |  | cm mm | 5 | ROI Thickness by US;LN;59089-3 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Interventricular Septum;SRT;T-32410 |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=M mode;SRT;G-0394
IVS s | Interventricular Septal Thickness systole |  | cm mm | 5 | ROI Thickness by US;LN;59089-3 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Interventricular Septum;SRT;T-32410 |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
IVS s | Interventricular Septal Thickness systole (M-mode) |  | cm mm | 5 | ROI Thickness by US;LN;59089-3 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Interventricular Septum;SRT;T-32410 |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LAD | Left Anterior Descending Coronary Artery Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Anterior Descending Branch of Left Coronary Artery;SRT;T-43110 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LAD D Vmax | Left Anterior Descending Coronary Artery D-wave Vmax |  | cm/s m/s | 5 | Diastolic blood velocity, peak;DCM;122202 | NUM | Finding Site;SRT;G-C0E3=Anterior Descending Branch of Left Coronary Artery;SRT;T-43110 |  | Fetal Measurements;DCM;125016
LAD S Vmax | Left Anterior Descending Coronary Artery S-wave Vmax |  | cm/s m/s | 5 | Systolic blood velocity, peak;DCM;122204 | NUM | Finding Site;SRT;G-C0E3=Anterior Descending Branch of Left Coronary Artery;SRT;T-43110 |  | Fetal Measurements;DCM;125016
LA A-P Diam s | Left Atrium Anterior-Posterior Diameter systole |  | cm mm | 5 | Left Atrium Antero-posterior Systolic Dimension;LN;29469-4 | NUM | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Area d | Left Atrium Area diastole |  | cm² mm² | 5 | Area;SRT;G-A166 | NUM | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Area s | Left Atrium Area systole |  | cm² mm² | 5 | Area;SRT;G-A166 | NUM | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Diam s | Left Atrium Diameter systole (M-mode) |  | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LA Major d | Left Atrium Major Axis diastole |  | cm mm | 5 | Major Axis;SRT;G-A193 | NUM | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Major s | Left Atrium Major Axis systole |  | cm mm | 5 | Major Axis;SRT;G-A193 | NUM | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Minor d | Left Atrium Minor Axis diastole |  | cm mm | 5 | Minor Axis;SRT;G-A194 | NUM | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA Minor s | Left Atrium Minor Axis systole |  | cm mm | 5 | Minor Axis;SRT;G-A194 | NUM | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LHV Vmax | Left Hepatic Vein Vmax |  | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Left Hepatic Vein;SRT;T-48727 |  | Fetal Measurements;DCM;125016
LLPV | Left Lower Pulmonary Vein Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LLPV A Duration | Left Lower Pulmonary Vein A-wave Duration |  | ms | 5 | A-Wave Duration;LN;59105-7 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV A Vmax | Left Lower Pulmonary Vein A-wave Vmax |  | cm/s m/s | 5 | Peak Reversal Velocity during Atrial Contraction;LN;59079-4 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV D | Left Lower Pulmonary Vein D-wave VTI | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010
LLPV D | Left Lower Pulmonary Vein D-wave Vmax | Vmax | cm/s m/s | 5 | Diastolic blood velocity, peak;DCM;122202 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV D | Left Lower Pulmonary Vein D-wave Vmean | Vmean | cm/s m/s | 5 | Diastolic blood velocity, mean;DCM;122201 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV D | Left Lower Pulmonary Vein D-wave Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010
LLPV D | Left Lower Pulmonary Vein D-wave Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010
LLPV D Vmax | Left Lower Pulmonary Vein D-wave Vmax |  | cm/s m/s | 5 | Diastolic blood velocity, peak;DCM;122202 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV S | Left Lower Pulmonary Vein S-wave VTI | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020
LLPV S | Left Lower Pulmonary Vein S-wave Vmax | Vmax | cm/s m/s | 5 | Systolic blood velocity, peak;DCM;122204 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV S | Left Lower Pulmonary Vein S-wave Vmean | Vmean | cm/s m/s | 5 | Systolic blood velocity, mean;DCM;122203 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV S | Left Lower Pulmonary Vein S-wave Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020
LLPV S | Left Lower Pulmonary Vein S-wave Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020
LLPV S Vmax | Left Lower Pulmonary Vein S-wave Vmax |  | cm/s m/s | 5 | Systolic blood velocity, peak;DCM;122204 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
Left Main | Left Main Coronary Artery Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Left Main Coronary Artery;SRT;T-43107 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Left Main D Vmax | Left Main Coronary Artery D-wave Vmax |  | cm/s m/s | 5 | Diastolic blood velocity, peak;DCM;122202 | NUM | Finding Site;SRT;G-C0E3=Left Main Coronary Artery;SRT;T-43107 |  | Fetal Measurements;DCM;125016
Left Main S Vmax | Left Main Coronary Artery S-wave Vmax |  | cm/s m/s | 5 | Systolic blood velocity, peak;DCM;122204 | NUM | Finding Site;SRT;G-C0E3=Left Main Coronary Artery;SRT;T-43107 |  | Fetal Measurements;DCM;125016
LPA AT | Left Pulmonary Artery Acceleration Time |  | ms | 5 | Acceleration Time;LN;20168-1 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Left Pulmonary Artery;SRT;T-44400
LPA ET | Left Pulmonary Artery Ejection Time |  | ms | 5 | Ejection Time;LN;20222-6 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Left Pulmonary Artery;SRT;T-44400
LPA | Left Pulmonary Artery Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Left Pulmonary Artery;SRT;T-44400 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LPA | Left Pulmonary Artery Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Left Pulmonary Artery;SRT;T-44400
LPA | Left Pulmonary Artery Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod | Finding Site;SRT;G-C0E3=Left Pulmonary Artery;SRT;T-44400
LPA | Left Pulmonary Artery Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Left Pulmonary Artery;SRT;T-44400
LPA | Left Pulmonary Artery Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod | Finding Site;SRT;G-C0E3=Left Pulmonary Artery;SRT;T-44400
LPA | Left Pulmonary Artery Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Left Pulmonary Artery;SRT;T-44400
LPA | Left Pulmonary Artery Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod | Finding Site;SRT;G-C0E3=Left Pulmonary Artery;SRT;T-44400
LPA | Left Pulmonary Artery Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod | Finding Site;SRT;G-C0E3=Left Pulmonary Artery;SRT;T-44400
LUPV | Left Upper Pulmonary Vein Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LUPV A Duration | Left Upper Pulmonary Vein A-wave Duration |  | ms | 5 | A-Wave Duration;LN;59105-7 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LUPV A Vmax | Left Upper Pulmonary Vein A-wave Vmax |  | cm/s m/s | 5 | Peak Reversal Velocity during Atrial Contraction;LN;59079-4 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LUPV D | Left Upper Pulmonary Vein D-wave VTI | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010
LUPV D | Left Upper Pulmonary Vein D-wave Vmax | Vmax | cm/s m/s | 5 | Diastolic blood velocity, peak;DCM;122202 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LUPV D | Left Upper Pulmonary Vein D-wave Vmean | Vmean | cm/s m/s | 5 | Diastolic blood velocity, mean;DCM;122201 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LUPV D | Left Upper Pulmonary Vein D-wave Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010
LUPV D | Left Upper Pulmonary Vein D-wave Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010
LUPV D Vmax | Left Upper Pulmonary Vein D-wave Vmax |  | cm/s m/s | 5 | Diastolic blood velocity, peak;DCM;122202 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LUPV S | Left Upper Pulmonary Vein S-wave VTI | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020
LUPV S | Left Upper Pulmonary Vein S-wave Vmax | Vmax | cm/s m/s | 5 | Systolic blood velocity, peak;DCM;122204 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LUPV S | Left Upper Pulmonary Vein S-wave Vmean | Vmean | cm/s m/s | 5 | Systolic blood velocity, mean;DCM;122203 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LUPV S | Left Upper Pulmonary Vein S-wave Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020
LUPV S | Left Upper Pulmonary Vein S-wave Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020
LUPV S Vmax | Left Upper Pulmonary Vein S-wave Vmax |  | cm/s m/s | 5 | Systolic blood velocity, peak;DCM;122204 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LV Area d | Left Ventricle Area diastole |  | cm² mm² | 5 | Area;SRT;G-A166 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Area s | Left Ventricle Area systole |  | cm² mm² | 5 | Area;SRT;G-A166 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV ET | Left Ventricle Ejection Time |  | ms | 5 | Ejection Time;LN;20222-6 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016
LV ET | Left Ventricle Ejection Time (M-mode) |  | ms | 5 | Ejection Time;LN;20222-6 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LVID d | Left Ventricle Internal Dimension diastole |  | cm mm | 5 | ROI Internal Dimension by US;LN;59090-1 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVID d | Left Ventricle Internal Dimension diastole (M-mode) |  | cm mm | 5 | ROI Internal Dimension by US;LN;59090-1 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LVID s | Left Ventricle Internal Dimension systole |  | cm mm | 5 | ROI Internal Dimension by US;LN;59090-1 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVID s | Left Ventricle Internal Dimension systole (M-mode) |  | cm mm | 5 | ROI Internal Dimension by US;LN;59090-1 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV IVCT | Left Ventricle IVCT |  | ms | 5 | Isovolumic Contraction Time;LN;59084-4 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016
LV IVRT | Left Ventricle IVRT |  | ms | 5 | Isovolumic Relaxation Time;LN;59083-6 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016
LV Major d | Left Ventricle Major Axis diastole |  | cm mm | 5 | Major Axis;SRT;G-A193 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Major s | Left Ventricle Major Axis systole |  | cm mm | 5 | Major Axis;SRT;G-A193 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Minor d | Left Ventricle Minor Axis diastole |  | cm mm | 5 | Minor Axis;SRT;G-A194 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Minor s | Left Ventricle Minor Axis systole |  | cm mm | 5 | Minor Axis;SRT;G-A194 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV PEP | Left Ventricle Pre-ejection Period |  | ms | 5 | Pre-Ejection Period;LN;59085-1 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016
LV PEP | Left Ventricle Pre-ejection Period (M-mode) |  | ms | 5 | Pre-Ejection Period;LN;59085-1 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LVPW d | Left Ventricle Posterior Wall Thickness diastole |  | cm mm | 5 | ROI Thickness by US;LN;59089-3 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Posterior Wall;SRT;R-42175 |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVPW d | Left Ventricle Posterior Wall Thickness diastole (M-mode) |  | cm mm | 5 | ROI Thickness by US;LN;59089-3 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Posterior Wall;SRT;R-42175 |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LVPW s | Left Ventricle Posterior Wall Thickness systole |  | cm mm | 5 | ROI Thickness by US;LN;59089-3 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Posterior Wall;SRT;R-42175 |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVPW s | Left Ventricle Posterior Wall Thickness systole (M-mode) |  | cm mm | 5 | ROI Thickness by US;LN;59089-3 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Posterior Wall;SRT;R-42175 |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LVOT | Left Ventricular Outflow Tract Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT | Left Ventricular Outflow Tract Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT | Left Ventricular Outflow Tract Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT | Left Ventricular Outflow Tract Peak Grad | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT | Left Ventricular Outflow Tract Mean Grad | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT | Left Ventricular Outflow Tract Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT | Left Ventricular Outflow Tract Peak Grad | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT IVCT | LeftVentricular Outflow Tract IVCT |  | ms | 5 | Isovolumic Contraction Time;LN;59084-4 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT IVRT | LeftVentricular Outflow Tract IVRT |  | ms | 5 | Isovolumic Relaxation Time;LN;59083-6 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT Diam s | Left Ventricular Outflow Tract Diameter systole |  | cm mm | 5 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 |  |  | Cardiac Cycle Point;SRT;R-4089A=End Systole;SRT;R-FAB5B | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV C-O Duration | Mitral Valve Closure to Opening Time |  | ms | 5 | Closure to Opening Time;LN;59082-8 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016
MPA | Main Pulmonary Artery Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MPA | Main Pulmonary Artery Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016
MPA | Main Pulmonary Artery Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016
MHV Vmax | Middle Heptatic Vein Vmax |  | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Middle Hepatic Vein;SRT;T-48726 |  | Fetal Measurements;DCM;125016
MAPSE | Mitral Annular Plane Systolic Excursion |  | cm mm | 5 | Mitral annular plane systolic excursion;99SIEMENS;MAPSE | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
MV a' Vmax (L) | Mitral Valve Lateral: A' Vmax |  | cm/s m/s | 5 | Peak Tissue Velocity;LN;59133-9 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Lateral Mitral Annulus;SRT;G-0392 |  |  | Cardiac Cycle Point;SRT;R-4089A=Atrial Systole;SRT;F-32030 | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV a' Vmax (M) | Mitral Valve Medial: A' Vmax |  | cm/s m/s | 5 | Peak Tissue Velocity;LN;59133-9 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Medial Mitral Annulus;SRT;G-0391 |  |  | Cardiac Cycle Point;SRT;R-4089A=Atrial Systole;SRT;F-32030 | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV A Duration | Mitral Valve A-wave Duration |  | ms | 5 | A-Wave Duration;LN;59105-7 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016
MV A Vmax | Mitral Valve A-wave Vmax |  | cm/s m/s | 5 | A-Wave Peak Velocity;LN;59081-0 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Annulus | Mitral Valve Annulus Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Annulus | Mitral Valve Annulus Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Annulus | Mitral Valve Annulus Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Annulus | Mitral Valve Annulus Mean Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Annulus | Mitral Valve Annulus Peak Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Annulus Circum d | Mitral Valve Annulus Circumference diastole |  | cm mm | 5 | Circumference;SRT;M-02560 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV Annulus Diam d | Mitral Valve Annulus Diameter diastole |  | cm mm | 5 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV Planimetry | Mitral Valve Area Planimetry |  | cm² mm² | 5 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Planimetry;DCM;125220 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV Decel | Mitral Valve Deceleration Slope | Slope | cm/s² m/s² | 5 | Deceleration Slope;LN;20216-8 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Decel | Mitral Valve Deceleration Time | Time | ms | 5 | Deceleration Time;LN;20217-6 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Decel | Mitral Valve Deceleration Velocity 1 | V1 | cm/s m/s | 5 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016
MV Decel | Mitral Valve Deceleration Velocity 2 | V2 | cm/s m/s | 5 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016
MV EPSS | E-Point Septal Separation (M-mode) |  | cm mm | 5 | Mitral Valve EPSS, E wave;LN;18036-4 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
MV e' Vmax (L) | Mitral Valve Lateral: E' Vmax |  | cm/s m/s | 5 | Peak Tissue Velocity;LN;59133-9 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Lateral Mitral Annulus;SRT;G-0392 |  |  | Cardiac Cycle Point;SRT;R-4089A=Early Diastole;SRT;R-40B1B | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV e' Vmax (M) | Mitral Valve Medial: E' Vmax |  | cm/s m/s | 5 | Peak Tissue Velocity;LN;59133-9 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Medial Mitral Annulus;SRT;G-0391 |  |  | Cardiac Cycle Point;SRT;R-4089A=Early Diastole;SRT;R-40B1B | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV E Vmax | Mitral Valve E-wave Vmax |  | cm/s m/s | 5 | E-Wave Peak Velocity;LN;59080-2 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016
MV | Mitral Valve Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV | Mitral Valve Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV | Mitral Valve Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV | Mitral Valve Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV | Mitral Valve Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV | Mitral Valve Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV | Mitral Valve Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MR | Mitral Valve Regurgitation Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR | Mitral Valve Regurgitation Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR | Mitral Valve Regurgitation Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR | Mitral Valve Regurgitation Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR | Mitral Valve Regurgitation Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR | Mitral Valve Regurgitation Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR | Mitral Valve Regurgitation Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
MR VC Width | Mitral Valve Regurgitation Vena Contracta |  | cm mm | 5 | Width;SRT;G-A220 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Vena Contracta;SRT;R-421AA |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 |  | Image Mode;SRT;G-0373=Doppler Color Flow;SRT;R-409E2
MV s' Vmax (L) | Mitral Valve Lateral: S' Vmax |  | cm/s m/s | 5 | Peak Tissue Velocity;LN;59133-9 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Lateral Mitral Annulus;SRT;G-0392 |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV s' Vmax (M) | Mitral Valve Medial: S' Vmax |  | cm/s m/s | 5 | Peak Tissue Velocity;LN;59133-9 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Medial Mitral Annulus;SRT;G-0391 |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
PFO | Patent Foramen Ovale Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Patent Foramen Ovale;SRT;D4-31012 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
PFO | Patent Foramen Ovale Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Patent Foramen Ovale;SRT;D4-31012 |  | Fetal Measurements;DCM;125016
PFO | Patent Foramen Ovale Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Patent Foramen Ovale;SRT;D4-31012 |  | Fetal Measurements;DCM;125016
PE d | Pericardial Effusion diastole |  | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Pericardium;SRT;T-39000 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Effusion;SRT;M-36700 |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
PR Interval | PR Interval |  | ms | 5 | PR Interval;99SIEMENS;PRInterval | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
PR Interval | PR Interval |  | ms | 5 | PR Interval;99SIEMENS;PRInterval | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
PR | Pulmonary Regurgitation VTI | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR | Pulmonary Regurgitation Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR | Pulmonary Regurgitation Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR | Pulmonary Regurgitation Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR | Pulmonary Regurgitation Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR | Pulmonary Regurgitation Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR | Pulmonary Regurgitation Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR Decel | Pulmonary Regurgitation Deceleration Slope | Slope | cm/s² m/s² | 5 | Deceleration Slope;LN;20216-8 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR Decel | Pulmonary Regurgitation Deceleration Time | Time | ms | 5 | Deceleration Time;LN;20217-6 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR Decel | Pulmonary Regurgitation Deceleration Velocity 1 | V1 | cm/s m/s | 5 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016
PR Decel | Pulmonary Regurgitation Deceleration Velocity 2 | V2 | cm/s m/s | 5 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016
PR ED | Pulmonary Regurgitation End Diastolic Velocity | Vel | cm/s m/s | 5 | End Diastolic Velocity;LN;11653-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PR ED | Pulmonary Regurgitation End Diastolic Gradient | Grad | mmHg | 5 | Pressure Gradient;SRT;F-023F7 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011
PV Annulus Diam s | Pulmonary Valve Annulus Diameter systole |  | cm mm | 5 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Cardiac valve annulus;SRT;T-3500E |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
PV ET | Pulmonary Valve Ejection Time |  | ms | 5 | Ejection Time;LN;20222-6 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016
PV AT | Pulmonary Valve Acceleration Time |  | ms | 5 | Acceleration Time;LN;20168-1 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV | Pulmonary Valve VTI | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV | Pulmonary Valve Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV | Pulmonary Valve Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV | Pulmonary Valve Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV | Pulmonary Valve Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV | Pulmonary Valve Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
PV | Pulmonary Valve Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
Pulm Side | Pulmonic Diameter | Diam | cm mm | 5 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Pulmonary side;99SIEMENS;PulmonarySide |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Pulm Side | Pulmonic Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Pulmonary side;99SIEMENS;PulmonarySide |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
Pulm Side | Pulmonic Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Pulmonary side;99SIEMENS;PulmonarySide |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
Pulm Side | Pulmonic Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Pulmonary side;99SIEMENS;PulmonarySide |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
Pulm Side | Pulmonic Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Pulmonary side;99SIEMENS;PulmonarySide |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
Pulm Side | Pulmonic Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Pulmonary side;99SIEMENS;PulmonarySide |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
RA Area d | Right Atrium Area diastole |  | cm² mm² | 5 | Area;SRT;G-A166 | NUM | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RA Area s | Right Atrium Area systole |  | cm² mm² | 5 | Area;SRT;G-A166 | NUM | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RA Major d | Right Atrium Major Axis diastole |  | cm mm | 5 | Major Axis;SRT;G-A193 | NUM | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RA Major s | Right Atrium Major Axis systole |  | cm mm | 5 | Major Axis;SRT;G-A193 | NUM | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RA Minor d | Right Atrium Minor Axis diastole |  | cm mm | 5 | Minor Axis;SRT;G-A194 | NUM | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RA Minor s | Right Atrium Minor Axis systole |  | cm mm | 5 | Minor Axis;SRT;G-A194 | NUM | Finding Site;SRT;G-C0E3=Right Atrium;SRT;T-32200 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RCA | Right Coronary Artery Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Right Coronary Artery;SRT;T-43203 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RCA D Vmax | Right Coronary Artery D-wave Vmax |  | cm/s m/s | 5 | Diastolic blood velocity, peak;DCM;122202 | NUM | Finding Site;SRT;G-C0E3=Right Coronary Artery;SRT;T-43203 |  | Fetal Measurements;DCM;125016
RCA S Vmax | Right Coronary Artery S-wave Vmax |  | cm/s m/s | 5 | Systolic blood velocity, peak;DCM;122204 | NUM | Finding Site;SRT;G-C0E3=Right Coronary Artery;SRT;T-43203 |  | Fetal Measurements;DCM;125016
RHV Vmax | Right Hepatic Vein Vmax |  | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Right Hepatic Vein;SRT;T-48725 |  | Fetal Measurements;DCM;125016
RLPV | Right Lower Pulmonary Vein Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RLPV A Duration | Right Lower Pulmonary Vein A-wave Duration |  | ms | 5 | A-Wave Duration;LN;59105-7 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV A Vmax | Right Lower Pulmonary Vein A-wave Vmax |  | cm/s m/s | 5 | Peak Reversal Velocity during Atrial Contraction;LN;59079-4 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV D | Right Lower Pulmonary Vein D-wave VTI | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010
RLPV D | Right Lower Pulmonary Vein D-wave Vmax | Vmax | cm/s m/s | 5 | Diastolic blood velocity, peak;DCM;122202 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV D | Right Lower Pulmonary Vein D-wave Vmean | Vmean | cm/s m/s | 5 | Diastolic blood velocity, mean;DCM;122201 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV D | Right Lower Pulmonary Vein D-wave Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010
RLPV D | Right Lower Pulmonary Vein D-wave Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010
RLPV D Vmax | Right Lower Pulmonary Vein D-wave Vmax |  | cm/s m/s | 5 | Diastolic blood velocity, peak;DCM;122202 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV S | Right Lower Pulmonary Vein S-wave VTI | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020
RLPV S | Right Lower Pulmonary Vein S-wave Vmax | Vmax | cm/s m/s | 5 | Systolic blood velocity, peak;DCM;122204 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV S | Right Lower Pulmonary Vein S-wave Vmean | Vmean | cm/s m/s | 5 | Systolic blood velocity, mean;DCM;122203 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV S | Right Lower Pulmonary Vein S-wave Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020
RLPV S | Right Lower Pulmonary Vein S-wave Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020
RLPV S Vmax | Right Lower Pulmonary Vein S-wave Vmax |  | cm/s m/s | 5 | Systolic blood velocity, peak;DCM;122204 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RPA AT | Right Pulmonary Artery Acceleration Time |  | ms | 5 | Acceleration Time;LN;20168-1 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Right Pulmonary Artery;SRT;T-44200
RPA | Right Pulmonary Artery Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Right Pulmonary Artery;SRT;T-44200 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RPA ET | Right Pulmonary Artery Ejection Time |  | ms | 5 | Ejection Time;LN;20222-6 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Right Pulmonary Artery;SRT;T-44200
RPA | Right Pulmonary Artery Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod | Finding Site;SRT;G-C0E3=Right Pulmonary Artery;SRT;T-44200
RPA | Right Pulmonary Artery Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod | Finding Site;SRT;G-C0E3=Right Pulmonary Artery;SRT;T-44200
RPA | Right Pulmonary Artery Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Right Pulmonary Artery;SRT;T-44200
RPA | Right Pulmonary Artery Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod | Finding Site;SRT;G-C0E3=Right Pulmonary Artery;SRT;T-44200
RPA | Right Pulmonary Artery Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Right Pulmonary Artery;SRT;T-44200
RPA | Right Pulmonary Artery Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod | Finding Site;SRT;G-C0E3=Right Pulmonary Artery;SRT;T-44200
RPA | Right Pulmonary Artery Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Pulmonary artery;SRT;T-44000 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Right Pulmonary Artery;SRT;T-44200
RUPV | Right Upper Pulmonary Vein Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RUPV A Duration | Right Upper Pulmonary Vein A-wave Duration |  | ms | 5 | A-Wave Duration;LN;59105-7 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RUPV A Vmax | Right Upper Pulmonary Vein A-wave Vmax |  | cm/s m/s | 5 | Peak Reversal Velocity during Atrial Contraction;LN;59079-4 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RUPV D | Right Upper Pulmonary Vein D-wave VTI | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010
RUPV D | Right Upper Pulmonary Vein D-wave Vmax | Vmax | cm/s m/s | 5 | Diastolic blood velocity, peak;DCM;122202 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RUPV D | Right Upper Pulmonary Vein D-wave Vmean | Vmean | cm/s m/s | 5 | Diastolic blood velocity, mean;DCM;122201 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RUPV D | Right Upper Pulmonary Vein D-wave Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010
RUPV D | Right Upper Pulmonary Vein D-wave Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010
RUPV D Vmax | Right Upper Pulmonary Vein D-wave Vmax |  | cm/s m/s | 5 | Diastolic blood velocity, peak;DCM;122202 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RUPV S | Right Upper Pulmonary Vein S-wave VTI | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020
RUPV S | Right Upper Pulmonary Vein S-wave Vmax | Vmax | cm/s m/s | 5 | Systolic blood velocity, peak;DCM;122204 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RUPV S | Right Upper Pulmonary Vein S-wave Vmean | Vmean | cm/s m/s | 5 | Systolic blood velocity, mean;DCM;122203 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RUPV S | Right Upper Pulmonary Vein S-wave Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020
RUPV S | Right Upper Pulmonary Vein S-wave Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  |  |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020
RUPV S Vmax | Right Upper Pulmonary Vein S-wave Vmax |  | cm/s m/s | 5 | Systolic blood velocity, peak;DCM;122204 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RVID d | Right Ventricle Diameter diastole |  | cm mm | 5 | ROI Internal Dimension by US;LN;59090-1 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVID d | Right Ventricle Diameter diastole (M-mode) |  | cm mm | 5 | ROI Internal Dimension by US;LN;59090-1 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=M mode;SRT;G-0394
RV ET | Right Ventricle Ejection Time |  | ms | 5 | Ejection Time;LN;20222-6 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016
RV ET | Right Ventricle Ejection Time (M-mode) |  | ms | 5 | Ejection Time;LN;20222-6 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
RVFW A' Vmax | Right Ventricle Free Wall A' Vmax |  | cm/s m/s | 5 | Peak Tissue Velocity;LN;59133-9 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  |  | Cardiac Cycle Point;SRT;R-4089A=Atrial Systole;SRT;F-32030 | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
RVFW d | Right Ventricle Free Wall diastole |  | cm mm | 5 | ROI Thickness by US;LN;59089-3 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Anterior Wall;SRT;R-4210B |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVFW d | Right Ventricle Free Wall diastole (M-mode) |  | cm mm | 5 | ROI Thickness by US;LN;59089-3 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Anterior Wall;SRT;R-4210B |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=M mode;SRT;G-0394
RVFW E' Vmax | Right Ventricle Free Wall E' Vmax |  | cm/s m/s | 5 | Peak Tissue Velocity;LN;59133-9 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  |  | Cardiac Cycle Point;SRT;R-4089A=Early Diastole;SRT;R-40B1B | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
RVFW S' Vmax | Right Ventricle Free Wall S' Vmax |  | cm/s m/s | 5 | Peak Tissue Velocity;LN;59133-9 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  |  | Cardiac Cycle Point;SRT;R-4089A=Peak Systolic;SRT;F-32021 | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
RVFW s | Right Ventricle Free Wall systole |  | cm mm | 5 | ROI Thickness by US;LN;59089-3 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Anterior Wall;SRT;R-4210B |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVFW s | Right Ventricle Free Wall systole (M-mode) |  | cm mm | 5 | ROI Thickness by US;LN;59089-3 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Anterior Wall;SRT;R-4210B |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=M mode;SRT;G-0394
RV IVCT | Right Ventricle IVCT |  | ms | 5 | Isovolumic Contraction Time;LN;59084-4 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016
RV IVRT | Right Ventricle IVRT |  | ms | 5 | Isovolumic Relaxation Time;LN;59083-6 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016
RV Major d | Right Ventricle Major Axis diastole |  | cm mm | 5 | Major Axis;SRT;G-A193 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Major s | Right Ventricle Major Axis systole |  | cm mm | 5 | Major Axis;SRT;G-A193 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Minor d | Right Ventricle Minor Axis diastole |  | cm mm | 5 | Minor Axis;SRT;G-A194 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RV Minor s | Right Ventricle Minor Axis systole |  | cm mm | 5 | Minor Axis;SRT;G-A194 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVID s | Right Ventricle Diameter systole |  | cm mm | 5 | ROI Internal Dimension by US;LN;59090-1 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVID s | Right Ventricle Diameter systole (M-mode) |  | cm mm | 5 | ROI Internal Dimension by US;LN;59090-1 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=M mode;SRT;G-0394
RV PEP | Right Ventricle Pre-ejection Period |  | ms | 5 | Pre-Ejection Period;LN;59085-1 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016
RV PEP | Right Ventricle Pre-ejection Period (M-mode) |  | ms | 5 | Pre-Ejection Period;LN;59085-1 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
RVOT Diam s | Right Ventricular Outflow Tract Diameter systole |  | cm mm | 5 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVOT | Right Ventricular Outflow Tract VTI | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT | Right Ventricular Outflow Tract Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT | Right Ventricular Outflow Tract Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT | Right Ventricular Outflow Tract Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT | Right Ventricular Outflow Tract Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT IVCT | Right Ventricular Outflow Tract IVCT |  | ms | 5 | Isovolumic Contraction Time;LN;59084-4 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT IVRT | Right Ventricular Outflow Tract IVRT |  | ms | 5 | Isovolumic Relaxation Time;LN;59083-6 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT | Right Ventricular Outflow Tract Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT | Right Ventricular Outflow Tract Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
TV C-O Duration | Tricuspid Valve Closure to Opening Time |  | ms | 5 | Closure to Opening Time;LN;59082-8 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016
SVC | Superior Vena Cava Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Superior Vena Cava;SRT;T-48610 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
SVC | Superior Vena Cava Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Superior Vena Cava;SRT;T-48610 |  | Fetal Measurements;DCM;125016
SVC | Superior Vena Cava Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Superior Vena Cava;SRT;T-48610 |  | Fetal Measurements;DCM;125016
Sys Side | Systemic Diameter | Diam | cm mm | 5 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Systemic side;99SIEMENS;SystemicSide |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Sys Side | Systemic Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Systemic side;99SIEMENS;SystemicSide |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
Sys Side | Systemic Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Systemic side;99SIEMENS;SystemicSide |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
Sys Side | Systemic Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Systemic side;99SIEMENS;SystemicSide |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
Sys Side | Systemic Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Systemic side;99SIEMENS;SystemicSide |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
Sys Side | Systemic Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Systemic side;99SIEMENS;SystemicSide |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TAPSE | Tricuspid Annular Plane Systolic Excursion |  | cm mm | 5 | Tricuspid annular plane systolic excursion;99SIEMENS;TAPSE | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=M mode;SRT;G-0394
TV | Tricuspid Valve Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV | Tricuspid Valve Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV | Tricuspid Valve Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV | Tricuspid Valve Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV | Tricuspid Valve Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV A Vmax | Tricuspid Valve A-wave Vmax |  | cm/s m/s | 5 | A-Wave Peak Velocity;LN;59081-0 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016
TV Annulus Diam d | Tricuspid Valve Annulus Diameter diastole |  | cm mm | 5 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
TV Decel | Tricuspid Valve Deceleleration Slope | Slope | cm/s² m/s² | 5 | Deceleration Slope;LN;20216-8 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Decel | Tricuspid Valve Deceleration Time | Time | ms | 5 | Deceleration Time;LN;20217-6 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV Decel | Tricuspid Valve Deceleration Velocity 1 | V1 | cm/s m/s | 5 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016
TV Decel | Tricuspid Valve Deceleration Velocity 2 | V2 | cm/s m/s | 5 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016
TV E Vmax | Tricuspid Valve E-wave Vmax |  | cm/s m/s | 5 | E-Wave Peak Velocity;LN;59080-2 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016
TR | Tricuspid Valve Regurgitation Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR | Tricuspid Valve Regurgitation Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR | Tricuspid Valve Regurgitation Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR | Tricuspid Valve Regurgitation Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR | Tricuspid Valve Regurgitation Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR | Tricuspid Valve Regurgitation Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TR | Tricuspid Valve Regurgitation Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TV | Tricuspid Valve Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV | Tricuspid Valve Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TruncArt V AT | Truncal Valve Acceleration Time |  | ms | 5 | Acceleration Time;LN;20168-1 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TruncArt V Annulus Diam s | Truncal Valve Annulus Diameter systole |  | cm mm | 5 | Cardiovascular Orifice Diameter;SRT;G-038F | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  |  | Finding Site;SRT;G-C0E3=Cardiac valve annulus;SRT;T-3500E |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
TruncArt V ET | Truncal Valve Ejection Time |  | ms | 5 | Ejection Time;LN;20222-6 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TruncArt V Regurg | Truncal Valve Regurgitation Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TruncArt V Regurg | Truncal Valve Regurgitation Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TruncArt V Regurg | Truncal Valve Regurgitation Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TruncArt V Regurg | Truncal Valve Regurgitation Peak Gradent | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TruncArt V Regurg | Truncal Valve Regurgitation Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TruncArt V Regurg Decel | Truncal Valve Regurgitation Deceleration Slope | Slope | cm/s² m/s² | 5 | Deceleration Slope;LN;20216-8 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TruncArt V Regurg Decel | Truncal Valve Regurgitation Deceleration Time | Time | ms | 5 | Deceleration Time;LN;20217-6 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TruncArt V Regurg Decel | Truncal Valve Regurgitation Deceleration Velocity 1 | V1 | cm/s m/s | 5 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016
TruncArt V Regurg Decel | Truncal Valve Regurgitation Deceleration Velocity 2 | V2 | cm/s m/s | 5 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016
TruncArt V Regurg ED | Truncal Valve Regurgitation End Diastolic Velocity | Vel | cm/s m/s | 5 | End Diastolic Velocity;LN;11653-3 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
TruncArt V Regurg ED | Truncal Valve Regurgitation End Diastolic Gradient | Grad | mmHg | 5 | Pressure Gradient;SRT;F-023F7 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367 | Cardiac Cycle Point;SRT;R-4089A=End Diastole;SRT;F-32011
TruncArt V | Truncal Valve Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TruncArt V | Truncal Valve Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TruncArt V | Truncal Valve Vmean | Vmean | cm/s m/s | 5 | Blood velocity, mean;DCM;122205 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TruncArt V | Truncal Valve Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Doppler Trace Method;99SIEMENS;DTraceMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TruncArt V | Truncal Valve Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TruncArt V | Truncal Valve Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TruncArt V | Truncal Valve Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  | Measurement Method;SRT;G-C036=Point Velocity Method;99SIEMENS;PtVelocityMethod |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
Truncus Art Diam s | Truncus Arteriosus Diameter systole |  | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Truncus Arteriosus Communis;SRT;D4-31400 |  | Fetal Measurements;DCM;125016 | [5 empty] | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Umbilical V | Umbilical Vein Velocity Time Integral | VTI | cm | 5 | Velocity Time Integral;LN;20354-7 | NUM | Finding Site;SRT;G-C0E3=Umbilical Vein;SRT;T-D03B4 |  | Fetal Measurements;DCM;125016
Umbilical V | Umbilical Vein Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Umbilical Vein;SRT;T-D03B4 |  | Fetal Measurements;DCM;125016
Umbilical V | Umbilical Vein Vmean | Vmean | cm/s m/s | 5 | Systolic blood velocity, mean;DCM;122203 | NUM | Finding Site;SRT;G-C0E3=Umbilical Vein;SRT;T-D03B4 |  | Fetal Measurements;DCM;125016
Umbilical V | Umbilical Vein Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Umbilical Vein;SRT;T-D03B4 |  | Fetal Measurements;DCM;125016
Umbilical V | Umbilical Vein Mean Gradient | Mean Grad | mmHg | 5 | Mean Gradient;LN;20256-4 | NUM | Finding Site;SRT;G-C0E3=Umbilical Vein;SRT;T-D03B4 |  | Fetal Measurements;DCM;125016
Umbilical V A Vmax | Umbilical Vein Vmax |  | cm/s m/s | 5 | A-Wave Peak Velocity;LN;59081-0 | NUM | Finding Site;SRT;G-C0E3=Umbilical Vein;SRT;T-D03B4 |  | Fetal Measurements;DCM;125016
Umbilical V D Vmax | Umbilical Vein D-wave Vmax |  | cm/s m/s | 5 | Diastolic blood velocity, peak;DCM;122202 | NUM | Finding Site;SRT;G-C0E3=Umbilical Vein;SRT;T-D03B4 |  | Fetal Measurements;DCM;125016
Umbilical V S Vmax | Umbilical Vein S-wave Vmax |  | cm/s m/s | 5 | Systolic blood velocity, peak;DCM;122204 | NUM | Finding Site;SRT;G-C0E3=Umbilical Vein;SRT;T-D03B4 |  | Fetal Measurements;DCM;125016
VSD | Ventricular Septal Defect Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Ventricular Septal Defect;SRT;D4-31150 |  | Fetal Measurements;DCM;125016 | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
VSD | Ventricular Septal Defect Vmax | Vmax | cm/s m/s | 5 | Blood velocity, peak;DCM;122207 | NUM | Finding Site;SRT;G-C0E3=Ventricular Septal Defect;SRT;D4-31150 |  | Fetal Measurements;DCM;125016
VSD | Ventricular Septal Defect Peak Gradient | Peak Grad | mmHg | 5 | Peak Gradient;LN;20247-3 | NUM | Finding Site;SRT;G-C0E3=Ventricular Septal Defect;SRT;D4-31150 |  | Fetal Measurements;DCM;125016
HC | Head Circumference |  | cm mm | 5 | Head Circumference;LN;11984-2 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
HA | Heart Area |  | cm² mm² | 5 | Cardiac Cross-sectional Area, transverse by US;LN;59075-2 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
Heart Circum | Heart Circumference |  | cm mm | 5 | Cardiac Circumference, transverse by US;LN;59073-7 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
HW | Hemispheric Width |  | cm mm | 5 | Width of Hemisphere;LN;12170-7 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
HL | Humerus Length |  | cm mm | 5 | Humerus length;LN;11966-9 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
Lat Vent | Lateral Ventricle width |  | cm mm | 5 | Lateral Ventricle width;LN;12171-5 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
Lt MCA | End Diastolic Velocity | ED | cm/s m/s | 5 | End Diastolic Velocity;LN;11653-3 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Fetal Measurements;DCM;125016
Lt MCA | Peak Systolic Velocity | PS | cm/s m/s | 5 | Peak Systolic Velocity;LN;11726-7 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Fetal Measurements;DCM;125016
Lt MCA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Fetal Measurements;DCM;125016
Lt MCA | Resistivity Index | RI | unitless | 5 | Resistivity Index;LN;12023-8 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Fetal Measurements;DCM;125016
Lt MCA | Pulsatility Index | PI | unitless | 5 | Pulsatility Index;LN;12008-9 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Fetal Measurements;DCM;125016
Lt MCA | Doppler Angle Property of PSV, EDV | Angle | ° | 5 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt MCA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | Time averaged peak velocity;LN;11692-1 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 |  | Fetal Measurements;DCM;125016
Lt MCA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | Time averaged mean velocity;LN;20352-1 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 |  | Fetal Measurements;DCM;125016
Rt MCA | End Diastolic Velocity | ED | cm/s m/s | 5 | End Diastolic Velocity;LN;11653-3 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Fetal Measurements;DCM;125016
Rt MCA | Peak Systolic Velocity | PS | cm/s m/s | 5 | Peak Systolic Velocity;LN;11726-7 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Fetal Measurements;DCM;125016
Rt MCA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Fetal Measurements;DCM;125016
Rt MCA | Resistivity Index | RI | unitless | 5 | Resistivity Index;LN;12023-8 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Fetal Measurements;DCM;125016
Rt MCA | Pulsatility Index | PI | unitless | 5 | Pulsatility Index;LN;12008-9 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Fetal Measurements;DCM;125016
Rt MCA | Doppler Angle Property of PSV, EDV | Angle | ° | 5 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt MCA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | Time averaged peak velocity;LN;11692-1 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 |  | Fetal Measurements;DCM;125016
Rt MCA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | Time averaged mean velocity;LN;20352-1 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 |  | Fetal Measurements;DCM;125016
OFD | Occipital-Frontal Diameter |  | cm mm | 5 | Occipital-Frontal Diameter;LN;11851-3 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
Q1 | Q1 Diameter |  | cm mm | 1 | First Quadrant Diameter;LN;11624-4 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
Q1 | Q1 Diameter |  | cm mm | 1 | First Quadrant Diameter;LN;11624-4 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 | Identifier;DCM;125010=AFI 2
Q1 | Q1 Diameter |  | cm mm | 1 | First Quadrant Diameter;LN;11624-4 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 | Identifier;DCM;125010=AFI 3
Q2 | Q2 Diameter |  | cm mm | 1 | Second Quadrant Diameter;LN;11626-9 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
Q2 | Q2 Diameter |  | cm mm | 1 | Second Quadrant Diameter;LN;11626-9 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 | Identifier;DCM;125010=AFI 2
Q2 | Q2 Diameter |  | cm mm | 1 | Second Quadrant Diameter;LN;11626-9 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 | Identifier;DCM;125010=AFI 3
Q3 | Q3 Diameter |  | cm mm | 1 | Third Quadrant Diameter;LN;11625-1 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
Q3 | Q3 Diameter |  | cm mm | 1 | Third Quadrant Diameter;LN;11625-1 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 | Identifier;DCM;125010=AFI 2
Q3 | Q3 Diameter |  | cm mm | 1 | Third Quadrant Diameter;LN;11625-1 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 | Identifier;DCM;125010=AFI 3
Q4 | Q4 Diameter |  | cm mm | 1 | Fourth Quadrant Diameter;LN;11623-6 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
Q4 | Q4 Diameter |  | cm mm | 1 | Fourth Quadrant Diameter;LN;11623-6 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 | Identifier;DCM;125010=AFI 2
Q4 | Q4 Diameter |  | cm mm | 1 | Fourth Quadrant Diameter;LN;11623-6 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 | Identifier;DCM;125010=AFI 3
TA | Thoracic Area |  | cm² mm² | 5 | Thoracic Area;LN;33068-8 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
TC | Thoracic Circumference |  | cm mm | 5 | Thoracic Circumference;LN;11988-3 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
Cerebellum | Cerebellum |  | cm mm | 5 | Trans Cerebellar Diameter;LN;11863-8 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
TTD | Transverse Thoracic Diameter |  | cm mm | 5 | Transverse Thoracic Diameter;LN;11864-6 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
Umb A | End Diastolic Velocity | ED | cm/s m/s | 5 | End Diastolic Velocity;LN;11653-3 | NUM | Finding Site;SRT;G-C0E3=Umbilical Artery;SRT;T-F1810 |  | Fetal Measurements;DCM;125016
Umb A | Peak Systolic Velocity | PS | cm/s m/s | 5 | Peak Systolic Velocity;LN;11726-7 | NUM | Finding Site;SRT;G-C0E3=Umbilical Artery;SRT;T-F1810 |  | Fetal Measurements;DCM;125016
Umb A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Finding Site;SRT;G-C0E3=Umbilical Artery;SRT;T-F1810 |  | Fetal Measurements;DCM;125016
Umb A | Resistivity Index | RI | unitless | 5 | Resistivity Index;LN;12023-8 | NUM | Finding Site;SRT;G-C0E3=Umbilical Artery;SRT;T-F1810 |  | Fetal Measurements;DCM;125016
Umb A | Pulsatility Index | PI | unitless | 5 | Pulsatility Index;LN;12008-9 | NUM | Finding Site;SRT;G-C0E3=Umbilical Artery;SRT;T-F1810 |  | Fetal Measurements;DCM;125016
Umb A | Doppler Angle Property of PSV, EDV | Angle | ° | 5 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Umb A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | Time averaged peak velocity;LN;11692-1 | NUM | Finding Site;SRT;G-C0E3=Umbilical Artery;SRT;T-F1810 |  | Fetal Measurements;DCM;125016
Umb A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | Time averaged mean velocity;LN;20352-1 | NUM | Finding Site;SRT;G-C0E3=Umbilical Artery;SRT;T-F1810 |  | Fetal Measurements;DCM;125016
Lt Uterine A | Peak Systolic Velocity | PS | cm/s m/s | 5 | Peak Systolic Velocity;LN;11726-7 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Lt Uterine A | End Diastolic Velocity | ED | cm/s m/s | 5 | End Diastolic Velocity;LN;11653-3 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Lt Uterine A | Resistivity Index | RI | unitless | 5 | Resistivity Index;LN;12023-8 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Lt Uterine A | Pulsatility Index | PI | unitless | 5 | Pulsatility Index;LN;12008-9 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Lt Uterine A | Doppler Angle Property of PSV, EDV, TAMx | Angle | ° | 5 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
Lt Uterine A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt Uterine A | Peak Systolic Velocity | PS | cm/s m/s | 5 | Peak Systolic Velocity;LN;11726-7 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt Uterine A | End Diastolic Velocity | ED | cm/s m/s | 5 | End Diastolic Velocity;LN;11653-3 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt Uterine A | Resistivity Index | RI | unitless | 5 | Resistivity Index;LN;12023-8 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt Uterine A | Pulsatility Index | PI | unitless | 5 | Pulsatility Index;LN;12008-9 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt Uterine A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt Uterine A | Doppler Angle Property of PSV, EDV, TAMx | Angle | ° | 5 | Doppler Angle; DCM; 125106 | NUM | Named Attribute
HC | Head Circumference |  | cm mm | 5
HC | Head Circumference |  | cm mm | 5
HC | Head Circumference |  | cm mm | 5
HC | Head Circumference | A | cm² mm² | 5
AC | Head Circumference |  | cm mm | 5
AC | Head Circumference |  | cm mm | 5
AC | Head Circumference |  | cm mm | 5
AC | Abdominal Circumference | A | cm² mm² | 5
Ao Arch | Aortic Arch Vmax | Angle | ° | 5
Ao Desc | Descending Aorta Vmax | Angle | ° | 5
AoV | Aortic Valve Vmax | Angle | ° | 5
AR ED | Aortic Regurgitation End Diastolic Velocity | Angle | ° | 5
LVOT | Left Ventricular Outflow Tract Vmax | Angle | ° | 5
MR | Mitral Valve Regurgitation Vmax | Angle | ° | 5
MV | Mitral Valve Vmax | Angle | ° | 5
LPA | Left Pulmonary Artery Vmax | Angle | ° | 5
MPA | Main Pulmonary Artery Vmax | Angle | ° | 5
RPA | Right Pulmonary Artery Vmax | Angle | ° | 5
PR ED | Pulmonary Regurgitation End Diastolic Velocity | Angle | ° | 5
PR | Pulmonary Regurgitation Vmax | Angle | ° | 5
PV | Pulmonary Valve Vmax | Angle | ° | 5
Ao Coarct | Aortic Coarctation Vmax | Angle | ° | 5
RVOT | Right Ventricular Outflow Tract Vmax | Angle | ° | 5
ASD | Atrial Septal Defect Vmax | Angle | ° | 5
AVSD | Atrioventricular Septal Defect Vmax | Angle | ° | 5
PFO | Patent Foramen Ovale Vmax | Angle | ° | 5
VSD | Ventricular Septal Defect Vmax | Angle | ° | 5
IVC | Inferior Vena Cava Vmax | Angle | ° | 5
SVC | Superior Vena Cava Vmax | Angle | ° | 5
TV | Tricuspid Valve Vmax | Angle | ° | 5
TR | Tricuspid Valve Regurgitation Vmax | Angle | ° | 5
TruncArt V | Truncal Valve Vmax | Angle | ° | 5
TruncArt V Regurg ED | Truncal Valve Regurgitation End Diastolic Velocity | Angle | ° | 5
Ao Asc | Ascending Aorta Peak Gradient | Angle | ° | 5
Lt Uterine A | Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Lt Uterine A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | Time averaged peak velocity;LN;11692-1 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Lt Uterine A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | Time averaged mean velocity;LN;20352-1 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Lt Uterine A | Acceleration Index | Accel | cm/s² m/s² | 5 | Acceleration Index;LN;20167-3 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Lt Uterine A | Acceleration Time | Time | ms | 5 | Acceleration Time;LN;20168-1 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Lt Uterine A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Lt Uterine A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Lt Uterine A | Volume Flow | Vol Flow | ml/min | 5 | Volume Flow;LN;33878-0 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Lt Uterine A | Volume Flow Diameter | D | cm mm | 5 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Lt Uterine A | Volume Flow Area | A | cm² mm² | 5 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Lt Uterine A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt Uterine A | Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts | [6 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Rt Uterine A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | Time averaged peak velocity;LN;11692-1 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt Uterine A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | Time averaged mean velocity;LN;20352-1 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt Uterine A | Acceleration Index | Accel | cm/s² m/s² | 5 | Acceleration Index;LN;20167-3 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt Uterine A | Acceleration Time | Time | ms | 5 | Acceleration Time;LN;20168-1 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt Uterine A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt Uterine A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt Uterine A | Volume Flow | Vol Flow | ml/min | 5 | Volume Flow;LN;33878-0 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt Uterine A | Volume Flow Diameter | D | cm mm | 5 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt Uterine A | Volume Flow Area | A | cm² mm² | 5 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt Uterine A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Umb A | Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Umbilical Artery;SRT;T-F1810 |  | Fetal Measurements;DCM;125016
Lt MCA | Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Fetal Measurements;DCM;125016
Rt MCA | Diameter | Diam | cm mm | 5 | Diameter;SRT;M-02550 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Fetal Measurements;DCM;125016
Ductus A | Ductus Arteriosus Minimum Diastolic Velocity | MD | cm/s m/s | 5 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Finding Site;SRT;G-C0E3=Ductus arteriosus;SRT;T-F6845 |  | Fetal Measurements;DCM;125016
Lt MCA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Fetal Measurements;DCM;125016
Lt Uterine A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Rt MCA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Finding Site;SRT;G-C0E3=Middle Cerebral Artery;SRT;T-45600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Fetal Measurements;DCM;125016
Rt Uterine A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Finding Site;SRT;G-C0E3=Uterine Artery;SRT;T-46820 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Maternal Measurements;99SIEMENS;MaternalMsrmts
Umb A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Finding Site;SRT;G-C0E3=Umbilical Artery;SRT;T-F1810 |  | Fetal Measurements;DCM;125016

## Ped & Fetal Echo Calculations
EnglishLabel | Description | DefaultUnits | DefaultPrecision | MaxInstances | ConceptCoding | ValueType | ParentContainer | ParentLaterality | SRSection | SRSectionLaterality | AdditionalModifiers | AdditionalModifiers2 | MeasurementMethod | FindingSite | TopographicalModifier | FlowDirection | CardiacCyclePoint | ImageMode
AFI | Amniotic Fluid Index | cm | 1 | 1 | Amniotic Fluid Index;LN;11627-7 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
AFI 2 | Amniotic Fluid Index 2 | cm | 1 | 1 | Amniotic Fluid Index;LN;11627-7 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 |  | Identifier;DCM;125010=AFI 2
AFI 3 | Amniotic Fluid Index 3 | cm | 1 | 1 | Amniotic Fluid Index;LN;11627-7 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 |  | Identifier;DCM;125010=AFI 3
APTDxTTD | Fetal Thoracic Cross-sectional Area | cm² | 1 | 1 | APTD x TTD;99SIEMENS;APTDxTTD | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
BPDa | Area-corrected Biparietal Diameter | cm | 1 | 1 | BPD area corrected;LN;11824-0 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
CI | Cephalic Index | unitless | 2 | 1 | Cephalic Index;LN;11823-2 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 |  | Normal Range Lower Limit;SRT;R-10041 | Normal Range Upper Limit;SRT;R-0038B
FL/AC | FL/AC | unitless | 2 | 1 | FL/AC;LN;11871-1 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 |  | Normal Range Lower Limit;SRT;R-10041 | Normal Range Upper Limit;SRT;R-0038B
FL/BPD | FL/BPD | unitless | 2 | 1 | FL/BPD;LN;11872-9 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 |  | Normal Range Lower Limit;SRT;R-10041 | Normal Range Upper Limit;SRT;R-0038B
FL/HC | FL/HC | unitless | 2 | 1 | FL/HC;LN;11873-7 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 |  | Normal Range Lower Limit;SRT;R-10041 | Normal Range Upper Limit;SRT;R-0038B
HC/AC | HC/AC | unitless | 2 | 1 | HC/AC;LN;11947-9 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 |  | Normal Range Lower Limit;SRT;R-10041 | Normal Range Upper Limit;SRT;R-0038B
LVW/HW | LVW/HW | unitless | 2 | 1 | LVW/HW;99SIEMENS;LVWOverHW | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 |  | Normal Range Lower Limit;SRT;R-10041 | Normal Range Upper Limit;SRT;R-0038B
TCD/AC | TCD/AC | unitless | 2 | 1 | TCD/AC;99SIEMENS;TCDOverAC | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 |  | Normal Range Lower Limit;SRT;R-10041 | Normal Range Upper Limit;SRT;R-0038B
Clinical Age | Fetus age based on LMP | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM | Fetus Characteristics;DCM;125015 |  | Fetus Characteristics;DCM;125015
LMP | Last Menstrual Period Date | DATE |  | 1 | LMP;LN;11955-2 | DATE | Summary;DCM;121111 |  | Summary;DCM;121111
DOC | In-vitro Fertilization Date | DATE |  | 1 | Conception date;LN;33067-0 | DATE | Summary;DCM;121111 |  | Summary;DCM;121111
EDD | Estimated Delivery Date | DATE |  | 1 | EDD;LN;11778-8 | DATE | Fetus Characteristics;DCM;125015 |  | Fetus Characteristics;DCM;125015
US EDD | Estimated Delivery Date from average ultrasound age | DATE |  | 1 | EDD from average ultrasound age;LN;11781-2 | DATE | Fetus Characteristics;DCM;125015 |  | Fetus Characteristics;DCM;125015
Num of fetuses | Number of fetuses | unitless | 0 | 1 | Number of Fetuses;LN;11878-6 | NUM | Summary;DCM;121111 |  | Summary;DCM;121111
CTR Circum | Cardiothoracic Circumference Ratio | unitless | 1 | 1 | Cardiothoracic Circumference Ratio;LN;59074-5 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 | [8 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
CTR Area | Cardiothoracic Area Ratio | unitless | 1 | 1 | Cardiothoracic Area Ratio;LN;59076-0 | NUM | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016 | [8 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Ao Root/MPA Diam | Aortic Root Diameter / Main Pulmonary Artery Diameter | unitless | 1 | 1 | Aortic Root to Main Pulmonary Artery Diameter Ratio;99SIEMENS;AoRootOvrMPADiam | NUM | Finding Site;SRT;G-C0E3=Root of Aorta;SRT;T-42110 |  | Findings;DCM;121070 | [8 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
AoV/PV Annulus s | Aortic Valve Annulus s / Pulmonary Valve Annulus s | unitless | 1 | 1 | Aortic Valve to Pulmonary Valve Systolic Annulus Diameter Ratio;99SIEMENS;AoVPVsAnnDRatio | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 | [8 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
AoV AT/ET | Aortic Valve AT/ET | unitless | 1 | 1 | Ratio of Aortic Valve Acceleration Time to Ejection Time;SRT;G-0382 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 | [6 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AoV Annulus Area s | Aortic Valve Annulus Area systole | cm² | 1 | 1 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Aortic Valve Ring;SRT;T-35410 |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
AoV SV | Aortic Valve Stroke Volume | ml | 1 | 1 | Stroke Volume;SRT;F-32120 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 | [8 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
AoV Area (VTI) | Aortic Valve AoV Area (VTI) | cm² | 1 | 1 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AoV Area (Vmax) | Aortic Valve AoV Area (Vmax) | cm² | 1 | 1 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Peak Velocity;DCM;125214 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
AR PHT | Aortic Regurgitation Pressure Half Time | ms | 0 | 1 | Pressure Half-Time;LN;20280-4 | NUM | Finding Site;SRT;G-C0E3=Aortic Valve;SRT;T-35400 |  | Fetal Measurements;DCM;125016 | [6 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
Cor Sinus Length/Diam | Coronary Sinus Length-to-Diameter Ratio | unitless | 1 | 1 | Coronary Sinus Length to Diameter Ratio;99SIEMENS;CorSinusLnDRatio | NUM | Finding Site;SRT;G-C0E3=Coronary sinus;SRT;T-32330 |  | Fetal Measurements;DCM;125016 | [8 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV E/A | Mtral Valve E/A | unitless | 1 | 1 | Peak E wave/Peak A wave by US;LN;59104-0 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 | [6 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV A/E | Mtral Valve A/E | unitless | 1 | 1 | Mitral Valve A to E Ratio;99SIEMENS;MVAtoERatio | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 | [6 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV E/e' (L) | Mitral Valve Lateral E/E' | unitless | 1 | 1 | Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave;SRT;G-037B | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Lateral Mitral Annulus;SRT;G-0392 |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV e'/a' (L) | Mitral Valve Lateral E'/A' | unitless | 1 | 1 | Left Ventricle E to A Tissue Velocity Ratio;LN;59129-7 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Lateral Mitral Annulus;SRT;G-0392 |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV E/e' (M) | Mitral Valve Medial E/E' | unitless | 1 | 1 | Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave;SRT;G-037B | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Medial Mitral Annulus;SRT;G-0391 |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
MV e'/a' (M) | Mitral Valve Medial E'/A' | unitless | 1 | 1 | Left Ventricle E to A Tissue Velocity Ratio;LN;59129-7 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Medial Mitral Annulus;SRT;G-0391 |  |  |  | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
RVFW E/E' | Right Ventricle Free Wall E/E' | unitless | 1 | 1 | Ratio of TV E Vmax to RV Free Wall Ea;99SIEMENS;RatioRVEePrime | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 | [8 empty] | Image Mode;SRT;G-0373=Tissue Doppler Imaging;SRT;P5-B0128
LA A-P Diam s/Ao Root | Left Atrium A-P Diameter s / Aortic Root Diameter | unitless | 1 | 1 | Left Atrium to Aortic Root Ratio;LN;17985-3 | NUM | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 |  | Fetal Measurements;DCM;125016 | [8 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA/RA Major d | Left Atrium Major d / Right Atrium Major d | unitless | 1 | 1 | Left Atrium Major to Right Atrium Major Diameter Ratio;99SIEMENS;LAMajRAMajDRatio | NUM | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 |  | Fetal Measurements;DCM;125016 | [7 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LA %FAC | Left Atrium % Fractional Area Change | % | 1 | 1 | % Fractional Area Change;99SIEMENS;PercentFAC | NUM | Finding Site;SRT;G-C0E3=Left Atrium;SRT;T-32300 |  | Fetal Measurements;DCM;125016 | [8 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Mass | LV Mass (M-mode) | g | 2 | 1 | Wall Mass;DCM;122447 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Left Ventricle Mass by M-mode;DCM;125221 |  |  |  |  | Image Mode;SRT;G-0373=M mode;SRT;G-0394
IVS/LVPW | Interventricular Septum / Posterior Wall Thickness | unitless | 1 | 1 | Interventricular Septum to Posterior Wall Thickness Ratio;LN;18155-2 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 | [7 empty] | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
IVS %Thick | Interventricular Septal % Thickness | % | 1 | 1 | % Thickening;LN;59092-7 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Interventricular Septum;SRT;T-32410 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVPW %Thick | Left Ventricle Posterior Wall % Thickness | % | 1 | 1 | % Thickening;LN;59092-7 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Posterior Wall;SRT;R-42175 |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV %FS | Left Ventricle % Fractional Shortening | % | 1 | 1 | Fractional Shortening;LN;59132-1 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 | [8 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d | Left Ventricle Volume d (Teichholz) | ml | 1 | 1 | Volume;SRT;G-D705 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol s | Left Ventricle Volume s (Teichholz) | ml | 1 | 1 | Volume;SRT;G-D705 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV EF | Left Ventricle Ejection Fraction (Teichholz) | % | 1 | 1 | Cardiac ejection fraction;SRT;F-32070 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV SV | LV Stroke Volume (Teichholz) | ml | 1 | 1 | Stroke Volume;SRT;F-32120 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV CO | LV Cardiac Output (Teichholz) | ml/min | 1 | 1 | Cardiac Output;SRT;F-32100 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV Vol d | Left Ventricle Volume d (Teichholz) [M-mode] | ml | 1 | 1 | Volume;SRT;G-D705 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV Vol s | Left Ventricle Volume s (Teichholz) [M-mode] | ml | 1 | 1 | Volume;SRT;G-D705 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV EF | Left Ventricle Ejection Fraction (Teichholz) [M-mode] | % | 1 | 1 | Cardiac ejection fraction;SRT;F-32070 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  |  |  | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV SV | LV Stroke Volume (Teichholz) [M-mode] | ml | 1 | 1 | Stroke Volume;SRT;F-32120 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  |  |  | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV CO | LV Cardiac Output (Teichholz) [M-mode] | ml/min | 1 | 1 | Cardiac Output;SRT;F-32100 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Teichholz;DCM;125209 |  |  |  |  | Image Mode;SRT;G-0373=M mode;SRT;G-0394
LV/RV Major d | Left Ventricle Major d / Right Ventricle Major d | unitless | 1 | 1 | Left Ventricle to Right Ventricle Diastolic Major Axis Ratio;99SIEMENS;LVRVdMajorRatio | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 | [8 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LVOT Area s | Left Ventricular Outflow Tract Area s (derived) | cm² | 1 | 1 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650 |  |  | Cardiac Cycle Point;SRT;R-4089A=End Systole;SRT;R-FAB5B | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
LV MPI | Left Ventricle Myocardial Performance Index | unitless | 1 | 1 | Myocardial Performance Index (Tei);LN;59099-2 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016
LVOT SV | Left Ventricular Outflow Tract Stroke Volume | ml | 1 | 1 | Stroke Volume;SRT;F-32120 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
LVOT CO | Left Ventricular Outflow Tract Cardiac Output | ml/min | 1 | 1 | Cardiac Output;SRT;F-32100 | NUM | Finding Site;SRT;G-C0E3=Left Ventricle;SRT;T-32600 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Left Ventricle Outflow Tract;SRT;T-32650
MV Annulus Area d | Mitral Valve Annulus Area d (derived) | cm² | 1 | 1 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Mitral Annulus;SRT;T-35313 |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV/TV Annulus d | Mitral Valve Annulus Diameter diastole / Tricuspid Valve Annulus Diameter diastole | unitless | 1 | 1 | Mitral Valve to Tricupid Valve Diastolic Annulus Diameter Ratio;99SIEMENS;MVTVdAnnDRatio | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 | [8 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MV PHT | Mitral Valve Pressure Half Time | ms | 0 | 1 | Pressure Half-Time;LN;20280-4 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 | [6 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Area (PHT) | Mitral Valve Area (PHT) | cm² | 1 | 1 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Area by Pressure Half-Time;DCM;125210 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV Area (VTI) | Mitral Valve Area (VTI) | cm² | 1 | 1 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215 |  |  | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
MV SV | Mitral Valve Stroke Volume | ml | 1 | 1 | Stroke Volume;SRT;F-32120 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 | [8 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
MR RF | Mitral Valve Regurgitant Fraction | % | 1 | 1 | Regurgitant Fraction;SRT;G-0390 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016
MR Vol | Mitral Valve Regurgitant Volume | ml | 1 | 1 | Volume;SRT;G-D705 | NUM | Finding Site;SRT;G-C0E3=Mitral Valve;SRT;T-35300 |  | Fetal Measurements;DCM;125016 | [6 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
PV Annulus Area s | Pulmonary Valve Annulus Area s (derived) | cm² | 1 | 1 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Cardiac valve annulus;SRT;T-3500E |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
PV AT/ET | Pulmonary Valve AT/ET | unitless | 1 | 1 | Ratio of Pulmonic Valve Acceleration Time to Ejection Time;SRT;G-0388 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016
PV Area (Vmax) | Pulmonary Valve Area (Vmax) | cm² | 1 | 1 | Pulmonic Valve Area by continuity;LN;18096-8 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Peak Velocity;DCM;125214
PV Area (VTI) | Pulmonary Valve Area (Velocity Time Integral) | cm² | 1 | 1 | Pulmonic Valve Area by continuity;LN;18096-8 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Continuity Equation by Velocity Time Integral;DCM;125215
PV SV | Pulmonary Valve Stroke Volume | ml | 1 | 1 | Stroke Volume;SRT;F-32120 | NUM | Finding Site;SRT;G-C0E3=Pulmonic Valve;SRT;T-35200 |  | Fetal Measurements;DCM;125016 | [8 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RUPV A Dur - MV A Dur | Right Upper Pulmonary Vein (A Dur - MV A Dur) | ms | 0 | 1 | Pulm Vein A duration to MV A duration difference;LN;59114-9 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 | [5 empty] | Topographical modifier;SRT;G-A1F8=Right Upper Segment;SRT;R-404A0
RUPV S/D Vmax | Right Upper Pulmonary Vein S Vmax / D Vmax | unitless | 1 | 1 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 | [5 empty] | Topographical modifier;SRT;G-A1F8=Right Upper Segment;SRT;R-404A0
RUPV Sys Filling Fraction | Right Upper Pulmonary Vein Systolic Filling Fraction | % | 1 | 1 | Pulmonary Vein Systolic Filling Fraction;99SIEMENS;PVeinSysFillFrac | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 | [5 empty] | Topographical modifier;SRT;G-A1F8=Right Upper Segment;SRT;R-404A0
RLPV A Dur - MV A Dur | Right Lower Pulmonary Vein (A Dur - MV A Dur) | ms | 0 | 1 | Pulm Vein A duration to MV A duration difference;LN;59114-9 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 | [5 empty] | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV S/D Vmax | Right Lower Pulmonary Vein S Vmax / D Vmax | unitless | 1 | 1 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 | [5 empty] | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
RLPV Sys Filling Fraction | Right Lower Pulmonary Vein Systolic Filling Fraction | % | 1 | 1 | Pulmonary Vein Systolic Filling Fraction;99SIEMENS;PVeinSysFillFrac | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 | [5 empty] | Topographical modifier;SRT;G-A1F8=Right Lower Segment;SRT;R-4049E
LUPV A Dur - MV A Dur | Left Upper Pulmonary Vein (A Dur - MV A Dur) | ms | 0 | 1 | Pulm Vein A duration to MV A duration difference;LN;59114-9 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 | [5 empty] | Topographical modifier;SRT;G-A1F8=Left Upper Segment;SRT;R-40491
LUPV S/D Vmax | Left Upper Pulmonary Vein S Vmax / D Vmax | unitless | 1 | 1 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 | [5 empty] | Topographical modifier;SRT;G-A1F8=Left Upper Segment;SRT;R-40491
LUPV Sys Filling Fraction | Left Upper Pulmonary Vein Systolic Filling Fraction | % | 1 | 1 | Pulmonary Vein Systolic Filling Fraction;99SIEMENS;PVeinSysFillFrac | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 | [5 empty] | Topographical modifier;SRT;G-A1F8=Left Upper Segment;SRT;R-40491
LLPV A Dur - MV A Dur | Left Lower Pulmonary Vein (A Dur - MV A Dur) | ms | 0 | 1 | Pulm Vein A duration to MV A duration difference;LN;59114-9 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 | [5 empty] | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV S/D Vmax | Left Lower Pulmonary Vein S Vmax / D Vmax | unitless | 1 | 1 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 | [5 empty] | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
LLPV Sys Filling Fraction | Left Lower Pulmonary Vein Systolic Filling Fraction | % | 1 | 1 | Pulmonary Vein Systolic Filling Fraction;99SIEMENS;PVeinSysFillFrac | NUM | Finding Site;SRT;G-C0E3=Pulmonary Vein;SRT;T-48581 |  | Fetal Measurements;DCM;125016 | [5 empty] | Topographical modifier;SRT;G-A1F8=Left Lower Segment;SRT;R-4214B
RVOT Area s | Right Ventricular Outflow Tract Area systole (derived) | cm² | 1 | 1 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550 |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
RVOT SV | Right Ventricular Outflow Tract Stroke Volume | ml | 1 | 1 | Stroke Volume;SRT;F-32120 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RVOT CO | Right Ventricular Outflow Tract Cardiac Output | ml/min | 1 | 1 | Cardiac Output;SRT;F-32100 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Right Ventricle Outflow Tract;SRT;T-32550
RV MPI | Right Ventricle Myocardial Performance Index (aka TEI Index) | unitless | 1 | 1 | Myocardial Performance Index (Tei);LN;59099-2 | NUM | Finding Site;SRT;G-C0E3=Right Ventricle;SRT;T-32500 |  | Fetal Measurements;DCM;125016
PFO/A Septum Length | Patent Foramen Ovale / Atrial Septum Length Ratio | unitless | 1 | 1 | Patent Foramen Ovale Diameter/Atrial Septum Length;99SIEMENS;PFODiamASLnRatio | NUM | Finding Site;SRT;G-C0E3=Patent Foramen Ovale;SRT;D4-31012 |  | Fetal Measurements;DCM;125016 | [8 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Pulm Side Area | Pulmonic Area (derived) | cm² | 1 | 1 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Pulmonary side;99SIEMENS;PulmonarySide |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Pulm Side SV | Pulmonic Stroke Volume | ml | 1 | 1 | Stroke Volume;SRT;F-32120 | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Pulmonary side;99SIEMENS;PulmonarySide
Pulm Side CO | Pulmonic Cardiac Output | ml/min | 1 | 1 | Cardiac Output;SRT;F-32100 | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Pulmonary side;99SIEMENS;PulmonarySide
Sys Side Area | Systemic Area (derived) | cm² | 1 | 1 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Systemic side;99SIEMENS;SystemicSide |  |  |  | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
Sys Side SV | Systemic Stroke Volume | ml | 1 | 1 | Stroke Volume;SRT;F-32120 | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Systemic side;99SIEMENS;SystemicSide
Sys Side CO | Systemic Cardiac Output | ml/min | 1 | 1 | Cardiac Output;SRT;F-32100 | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Systemic side;99SIEMENS;SystemicSide
Qp/Qs (SV) | Qp / Qs (using stroke volume) | unitless | 1 | 1 | Pulmonary-to-Systemic Shunt Flow Ratio;LN;29462-9 | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Stroke Volume;SRT;F-32120
Qp/Qs (CO) | Qp / Qs (using cardiac output) | unitless | 1 | 1 | Pulmonary-to-Systemic Shunt Flow Ratio;LN;29462-9 | NUM | Finding Site;SRT;G-C0E3=Cardiac Shunt Study;SRT;P5-30031 |  | Fetal Measurements;DCM;125016 |  |  |  | Measurement Method;SRT;G-C036=Cardiac Output;SRT;F-32100
TV Annulus Area d | Tricuspid Valve Annulus Area diastole (derived) | cm² | 1 | 1 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Tricuspid Annulus;SRT;T-35111 |  |  | Cardiac Cycle Point;SRT;R-4089A=Diastole;SRT;F-32010 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
TV E/A | Tricuspid Valve E/A | unitless | 1 | 1 | Peak E wave/Peak A wave by US;LN;59104-0 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016
TV PHT | Tricuspid Valve Pressure Half Time | ms | 0 | 1 | Pressure Half-Time;LN;20280-4 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 | [6 empty] | Flow Direction;SRT;G-C048=Antegrade Flow;SRT;R-42047
TV SV | Tricuspid Valve Stroke Volume | ml | 1 | 1 | Stroke Volume;SRT;F-32120 | NUM | Finding Site;SRT;G-C0E3=Tricuspid Valve;SRT;T-35100 |  | Fetal Measurements;DCM;125016 | [8 empty] | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
TruncArt V Annulus Area s | Truncal Valve Annulus Area systole (derived) | cm² | 1 | 1 | Cardiovascular Orifice Area;SRT;G-038E | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 |  |  |  |  | Finding Site;SRT;G-C0E3=Cardiac valve annulus;SRT;T-3500E |  |  | Cardiac Cycle Point;SRT;R-4089A=Systole;SRT;F-32020 | Image Mode;SRT;G-0373=2D mode;SRT;G-03A2
TruncArt V Regurg PHT | Truncal Valve Regurgitation Pressure Half Time | ms | 0 | 1 | Pressure Half-Time;LN;20280-4 | NUM | Finding Site;SRT;G-C0E3=Truncal Valve Structure;SRT;T-35014 |  | Fetal Measurements;DCM;125016 | [6 empty] | Flow Direction;SRT;G-C048=Regurgitant Flow;SRT;G-0367
Ductus V S/D Vmax | Ductus Venosus S Vmax/D Vmax | unitless | 1 | 1 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Finding Site;SRT;G-C0E3=Ductus venosus;SRT;T-F6806 |  | Fetal Measurements;DCM;125016
CCO | Combined Cardiac Output | ml/min | 1 | 1 | Combined Cardiac Output;99SIEMENS;CombinedCO |  | Fetal Measurements;DCM;125016 |  | Fetal Measurements;DCM;125016
Rt CP Ratio | Rt MCA PI / Umb A PI | unitless | 2 | 1 | Cerebroplacental ratio;99SIEMENS;CPRatio | NUM |  | Laterality;SRT;G-C171=Right;SRT;G-A100 | Fetal Measurements;DCM;125016
Lt CP Ratio | Lt MCA PI / Umb A PI | unitless | 2 | 1 | Cerebroplacental ratio;99SIEMENS;CPRatio | NUM |  | Laterality;SRT;G-C171=Left;SRT;G-A101 | Fetal Measurements;DCM;125016
INTERGROWTH-21st | BPD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD by GA, INTERGROWTH-21st 2014;99SIEMENS;GR107-29-01
INTERGROWTH-21st | HC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC by GA, INTERGROWTH-21st 2014;99SIEMENS;GR116-29-01
INTERGROWTH-21st | AC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=AC by GA, INTERGROWTH-21st 2014;99SIEMENS;GR102-29-01
INTERGROWTH-21st | FL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FL by GA, INTERGROWTH-21st 2014;99SIEMENS;GR112-29-01
INTERGROWTH-21st | OFD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=OFD by GA, INTERGROWTH-21st 2014;99SIEMENS;GR119-29-01
Williams Male | EFW(GC) | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FWP by GA, Williams 1982 (male);LN;33184-3
Williams Female | EFW(GC) | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FWP by GA, Williams 1982 (female);LN;33184-3
Moore | AFI(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=AFI by GA, Moore 1990;99SIEMENS;GR135-27-01
Moore | AFI B(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=AFI by GA, Moore 1990;99SIEMENS;GR135-27-01
Moore | AFI C(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=AFI by GA, Moore 1990;99SIEMENS;GR135-27-01
Williams | HCC(GC) | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FWP by GA, Williams 1982 (average);LN;33184-3
Hadlock | HCC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=HC by GA, Hadlock 1984;LN;33173-6
Hansmann | HCC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC by GA, Hansmann 1986;99SIEMENS;GR116-09-01
Merz | HCC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC by GA, Merz 1996;99SIEMENS;GR116-16-01
Chitty(Derived) | HCC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=HC derived by GA, Chitty 1994;LN;33174-4
Deter | HCC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=HC by GA, Deter 1982;99SIEMENS;GR116-06-01
ASUM 2003 | HCC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC by GA, ASUM 2003;99SIEMENS;GR116-01-01
CFEF | HCC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC by GA, CFEF 2000;99SIEMENS;GR116-03-01
Verburg | HCC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC by GA, Verburg 2008;99SIEMENS;GR116-22-01
Nicolaides | HCC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC by GA, Nicolaides 1994;99SIEMENS;GR116-17-01
INTERGROWTH-21st | HCC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC by GA, INTERGROWTH-21st 2014;99SIEMENS;GR116-29-01
Hadlock | AC (GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=AC by GA, Hadlock 1984;LN;33146-2
Nicolaides | AC (GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=AC by GA, Nicolaides 1994;99SIEMENS;GR102-17-01
Jeanty | AC (GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=AC by GA, Jeanty 1984;99SIEMENS;GR102-11-01
Merz | AC (GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=AC by GA, Merz 1996;99SIEMENS;GR102-16-01
Grumbach (twins) | AC (GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=AC by GA, Grumbach 1986 (twin);99SIEMENS;GR102-07-01
Verburg | AC (GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=AC by GA, Verburg 2008;99SIEMENS;GR102-22-01
Chitty | AC (GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=AC (measured) by GA, Chitty 1994;LN;33147-0
Deter | AC (GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=AC by GA, Deter 1982;99SIEMENS;GR102-06-01
JSUM 2003 | AC (GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=AC by GA, JSUM 2003;99SIEMENS;GR102-12-01
ASUM 2003 | AC (GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=AC by GA, ASUM 2003;99SIEMENS;GR102-01-01
CFEF | AC (GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=AC by GA, CFEF 2000;99SIEMENS;GR102-03-01
Hadlock | BPD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=BPD by GA, Hadlock 1984;LN;33198-3
Verburg | BPD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD by GA, Verburg 2008;99SIEMENS;GR107-22-01
Hansmann | BPD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD by GA, Hansmann 1986;99SIEMENS;GR107-09-01
Lasser | BPD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=BPD by GA, Lasser 1993;99SIEMENS;GR107-14-01
Merz | BPD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD by GA, Merz 1996;99SIEMENS;GR107-16-01
Rempen | BPD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD by GA, Rempen 1991;LN;33155-3
Chitty (outer-inner) | BPD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=BPD outer-inner by GA, Chitty 1994;LN;33556-2
Chitty (outer-outer) | BPD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=BPD outer-outer by GA, Chitty 1994;LN;33152-0
Grumbach (twins) | BPD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD by GA, Grumbach 1986 (twin);99SIEMENS;GR107-07-01
JSUM 2003 | BPD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD by GA, JSUM 2003;99SIEMENS;GR107-12-01
Osaka | BPD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD by GA, Osaka 1983;99SIEMENS;GR107-18-01
ASUM 2003 | BPD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD by GA, ASUM 2003;99SIEMENS;GR107-01-01
CFEF | BPD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD by GA, CFEF 2000;99SIEMENS;GR107-03-01
Nicolaides | BPD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD by GA, Nicolaides 1994;99SIEMENS;GR107-17-01
Hadlock | HC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=HC by GA, Hadlock 1984;LN;33173-6
Hansmann | HC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC by GA, Hansmann 1986;99SIEMENS;GR116-09-01
Merz | HC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC by GA, Merz 1996;99SIEMENS;GR116-16-01
Chitty | HC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=HC by GA, Chitty 1994;99SIEMENS;GR116-05-01
Deter | HC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=HC by GA, Deter 1982;99SIEMENS;GR116-06-01
ASUM 2003 | HC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC by GA, ASUM 2003;99SIEMENS;GR116-01-01
CFEF | HC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC by GA, CFEF 2000;99SIEMENS;GR116-03-01
Verburg | HC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC by GA, Verburg 2008;99SIEMENS;GR116-22-01
Nicolaides | HC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC by GA, Nicolaides 1994;99SIEMENS;GR116-17-01
Hadlock | FL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=FL by GA, Hadlock 1984;LN;33166-0
Jeanty | FL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FL by GA, Jeanty 1984;99SIEMENS;GR112-11-01
Merz | FL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FL by GA, Merz 1991;99SIEMENS;GR112-16-01
Chitty | FL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=FL by GA, Chitty 1994;LN;33167-8
Warda | FL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=FL by GA, Warda 1985;99SIEMENS;GR112-23-01
JSUM 2003 | FL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FL by GA, JSUM 2003;99SIEMENS;GR112-12-01
Tokyo | FL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FL by GA, Tokyo;99SIEMENS;GR112-20-01
Osaka | FL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=FL by GA, Osaka 1983;99SIEMENS;GR112-18-01
ASUM 2003 | FL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FL by GA, ASUM 2003;99SIEMENS;GR112-01-01
CFEF | FL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FL by GA, CFEF 2000;99SIEMENS;GR112-03-01
Hansmann | FL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FL by GA, Hansmann 1986;99SIEMENS;GR112-09-01
Nicolaides | FL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FL by GA, Nicolaides 1994;99SIEMENS;GR112-17-01
Verburg | FL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FL by GA, Verburg 2008;99SIEMENS;GR112-22-01
ASUM 2003 | OFD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=OFD by GA, ASUM 2003;99SIEMENS;GR119-01-01
Hansmann | OFD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=OFD by GA, Hansmann 1986;99SIEMENS;GR119-09-01
Nicolaides | OFD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=OFD by GA, Nicolaides 1994;99SIEMENS;GR119-17-01
Chitty | OFD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=OFD by GA, Chitty 1994;LN;33179-3
Jeanty | HL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HL by GA, Jeanty 1984;99SIEMENS;GR117-11-01
Chitty | HL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HL by GA, Chitty 2002;99SIEMENS;GR117-05-01
Merz | HL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HL by GA, Merz 1996;99SIEMENS;GR117-16-01
Osaka | HL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HL by GA, Osaka 1983;99SIEMENS;GR117-18-01
ASUM 2003 | HL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HL by GA, ASUM 2003;99SIEMENS;GR117-01-01
Hansmann | HL(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HL by GA, Hansmann 1986;99SIEMENS;GR117-09-01
JSUM 2003 RI | UMA PI(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Umbilical Artery RI by GA JSUM 2003;99SIEMENS;GR134-12-01
JSUM 2003 PI | UMA PI(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Umbilical Artery PI by GA JSUM 2003;99SIEMENS;GR133-12-01
Ebbing | UMA PI(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=Umbilical Artery PI by GA Ebbing;99SIEMENS;UmbArtPIEbbing
Ebbing | Rt CPRatio(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=Right Cerebroplacental Ratio by GA Ebbing;99SIEMENS;RtCPRatioEbbing
Ebbing | Lt CPRatio(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=Left Cerebroplacental Ratio by GA Ebbing;99SIEMENS;LtCPRatioEbbing
Ebbing | Rt MCA(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=Right MCA PI by GA Ebbing;99SIEMENS;RightMCAPIEbbing
Ebbing | Lt MCA(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=Left MCA PI by GA Ebbing;99SIEMENS;LeftMCAPIEbbing
JSUM 2003 RI | Rt MCA(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Right MCA RI by GA JSUM 2003;99SIEMENS;GR132-12-01
JSUM 2003 PI | Rt MCA(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Right MCA PI by GA JSUM 2003;99SIEMENS;GR131-12-01
JSUM 2003 RI | Lt MCA(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Left MCA RI by GA JSUM 2003;99SIEMENS;GR130-12-01
JSUM 2003 PI | Lt MCA(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Left MCA PI by GA JSUM 2003;99SIEMENS;GR129-12-01
Hadlock | EFW(GC) | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=EFW by GA, Hadlock 1991;LN;33162-9
Jeanty | EFW(GC) | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FWP by GA, Jeanty 1994;99SIEMENS;GR101-11-01
Yarkoni (twins) | EFW(GC) | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FWP by GA, Yarkoni 1987;99SIEMENS;GR101-24-01
JSUM 2003 | EFW(GC) | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FWP by GA, JSUM 2003;99SIEMENS;GR101-12-01
Tokyo* | EFW(GC) | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FWP by GA, Tokyo;99SIEMENS;GR101-20-01
Osaka | EFW(GC) | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FWP by GA, Osaka 1983;99SIEMENS;GR101-18-01
Williams Female | EFW(GC) | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FWP by GA, Williams 1982 (female);LN;33184-3
Brenner | EFW(GC) | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FWP by GA, Brenner 1976;LN;33189-2
Doubilet | EFW(GC) | % | 1 | 1 | EFW percentile rank;LN;11767-1 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FWP by GA, Doubilet 1996;99SIEMENS;GR101-26-01
Nicolaides | TCD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=TCD by GA, Nicolaides 1994;99SIEMENS;GR108-17-01
Verburg | TCD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=TCD by GA, Verburg 2008;99SIEMENS;GR108-22-01
Chitkara | TC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=TC by GA, Chitkara 1987;99SIEMENS;GR122-04-01
Campbell | HC/AC(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC/AC by GA, Campbell 1977;LN;33182-7
Tokyo | APTDxTTD(GC) | % | 1 | 1 | Growth Percentile Rank;DCM;125012 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=APTDxTTD by GA, Tokyo;99SIEMENS;GR105-20-01
Average | Composite Ultrasound Age | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Average Ultrasound Age;LN;11884-4 | 2 Sigma deviation of population ;DCM;121417
Hadlock (BPD, HC, AC, FL) | Composite Ultrasound Age | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Composite US Age, Hadlock (BPD, HC, AC, FL);99SIEMENS;CA100-08-07 | 2 Sigma deviation of population ;DCM;121417
Hadlock (BPD, HC, AC) | Composite Ultrasound Age | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Composite US Age, Hadlock (BPD, HC, AC);99SIEMENS;CA100-08-06 | 2 Sigma deviation of population ;DCM;121417
Hadlock (HC, AC, FL) | Composite Ultrasound Age | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Composite US Age, Hadlock (HC, AC, FL);99SIEMENS;CA100-08-10 | 2 Sigma deviation of population ;DCM;121417
Hadlock (BPD, HC, FL) | Composite Ultrasound Age | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Composite US Age, Hadlock (BPD, HC, FL);99SIEMENS;CA100-08-08 | 2 Sigma deviation of population ;DCM;121417
Hadlock (BPD, AC, FL) | Composite Ultrasound Age | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Composite US Age, Hadlock (BPD, AC, FL);99SIEMENS;CA100-08-03 | 2 Sigma deviation of population ;DCM;121417
Hadlock (AC, FL) | Composite Ultrasound Age | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Composite US Age, Hadlock (AC, FL);99SIEMENS;CA100-08-01 | 2 Sigma deviation of population ;DCM;121417
Hadlock (HC, FL) | Composite Ultrasound Age | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Composite US Age, Hadlock (HC, FL);99SIEMENS;CA100-08-11 | 2 Sigma deviation of population ;DCM;121417
Hadlock (HC, AC) | Composite Ultrasound Age | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Composite US Age, Hadlock (HC, AC);99SIEMENS;CA100-08-09 | 2 Sigma deviation of population ;DCM;121417
Hadlock (BPD, FL) | Composite Ultrasound Age | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Composite US Age, Hadlock (BPD, FL);99SIEMENS;CA100-08-04 | 2 Sigma deviation of population ;DCM;121417
Hadlock (BPD, HC) | Composite Ultrasound Age | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Composite US Age, Hadlock (BPD, HC);99SIEMENS;CA100-08-05 | 2 Sigma deviation of population ;DCM;121417
Hadlock (BPD, AC) | Composite Ultrasound Age | days | 0 | 1 | Composite Ultrasound Age;LN;11888-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Composite US Age, Hadlock (BPD, AC);99SIEMENS;CA100-08-02 | 2 Sigma deviation of population ;DCM;121417
Hadlock (BPD, HC, AC, FL) | Estimated Fetal Weight | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=EFW by AC, BPD, FL, HC, Hadlock 1985;LN;11732-5 | 2 Sigma deviation of population ;DCM;121417
Hadlock (BPD, AC, FL) | Estimated Fetal Weight | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=EFW by AC, BPD, FL, Hadlock 1985;LN;11735-8 | 2 Sigma deviation of population ;DCM;121417
Hadlock (HC, AC, FL) | Estimated Fetal Weight | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=EFW by AC, FL, HC, Hadlock 1985;LN;11746-5 | 2 Sigma deviation of population ;DCM;121417
Hadlock (AC, FL) | Estimated Fetal Weight | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=EFW by AC, FL, Hadlock 1985;LN;11751-5 | 2 Sigma deviation of population ;DCM;121417
Shepard (BPD, AC) | Estimated Fetal Weight | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=EFW by AC and BPD, Shepard 1982;LN;11739-0 | 2 Sigma deviation of population ;DCM;121417
Merz (BPD, AC) | Estimated Fetal Weight | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=EFW by BPD, AC, Merz 2005;99SIEMENS;EFW101-16-01 | 2 Sigma deviation of population ;DCM;121417
Tokyo (BPD, APTD, TTD, FL) | Estimated Fetal Weight | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=EFW by BPD, APTD, TTD, FL, Tokyo 1987;99SIEMENS;EFW101-20-01
JSUM (BPD, AC, FL) | Estimated Fetal Weight | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=EFW2 by Shinozuka 1996;LN;33142-1
JSUM (BPD, APTD, TTD, FL) | Estimated Fetal Weight | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=EFW by BPD, APTD, TTD, FL, JSUM 2003;99SIEMENS;EFW101-12-01
Hansmann (BPD, TTD) | Estimated Fetal Weight | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=EFW by BPD, TTD, Hansmann 1986;LN;33139-7
Osaka (BPD, FTA, FL) | Estimated Fetal Weight | g | 1 | 1 | Estimated Weight;LN;11727-5 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=EFW by BPD, FTA, FL, Osaka 1990;LN;33140-5
Hadlock | BPD (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=BPD, Hadlock 1984;LN;11902-4 | 2 Sigma deviation of population ;DCM;121417
Verburg (outer-outer) | BPD (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD (outer-outer), Verburg 2008;99SIEMENS;GA107-22-01 | 2 Sigma deviation of population ;DCM;121417
Jeanty | BPD (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD, Jeanty 1984;LN;11905-7 | 2 Sigma deviation of population ;DCM;121417
Lasser (outer-inner) | BPD (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD (outer-inner), Lasser 1993;99SIEMENS;GA107-14-01 | 2 Sigma deviation of population ;DCM;121417
Rempen | BPD (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD, Rempen 2001;99SIEMENS;GA107-25-01 | 2 Sigma deviation of population ;DCM;121417
Merz | BPD (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD, Merz 1991;99SIEMENS;GA107-16-01 | 2 Sigma deviation of population ;DCM;121417
Hansmann | BPD (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD, Hansmann 1986;LN;33538-0 | 2 Sigma deviation of population ;DCM;121417
JSUM 2003 | BPD (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD, JSUM 2003;99SIEMENS;GA107-12-01 | 2 Sigma deviation of population ;DCM;121417
Chitty (outer-inner) | BPD (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD-oi, Chitty 1997;LN;33086-0 | 2 Sigma deviation of population ;DCM;121417
Chitty (outer-outer) | BPD (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD-oo, Chitty 1997;LN;33087-8 | 2 Sigma deviation of population ;DCM;121417
ASUM 2003 | BPD (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=BPD, ASUM 2003;99SIEMENS;GA107-01-01
Osaka* | BPD (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=BPD, Osaka 1983;99SIEMENS;GA107-18-01
Hadlock | AC (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=AC, Hadlock 1984;LN;11892-7 | 2 Sigma deviation of population ;DCM;121417
Lasser | AC (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=AC, Lasser 1993;99SIEMENS;GA102-14-01 | 2 Sigma deviation of population ;DCM;121417
Merz | AC (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=AC, Merz 1991;99SIEMENS;GA102-16-01 | 2 Sigma deviation of population ;DCM;121417
JSUM 2003 | AC (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=AC, JSUM 2003;99SIEMENS;GA102-12-01 | 2 Sigma deviation of population ;DCM;121417
ASUM 2003 | AC (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=AC, ASUM 2003;99SIEMENS;GA102-01-01
Chitty | AC (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=AC, Chitty Plotted 1994;99SIEMENS;GA102-17-01 | 2 Sigma deviation of population ;DCM;121417
Hadlock | FL (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=FL, Hadlock 1984;LN;11920-6 | 2 Sigma deviation of population ;DCM;121417
Jeanty | FL (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=FL, Jeanty 1984;LN;11923-0 | 2 Sigma deviation of population ;DCM;121417
Merz | FL (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FL, Merz 1991;99SIEMENS;GA112-16-01 | 2 Sigma deviation of population ;DCM;121417
Warda | FL (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=FL, Warda;99SIEMENS;GA112-23-01 | 2 Sigma deviation of population ;DCM;121417
JSUM 2003 | FL (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FL, JSUM 2003;99SIEMENS;GA112-12-01 | 2 Sigma deviation of population ;DCM;121417
Tokyo* | FL (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FL, Tokyo 1986;LN;33103-3
Osaka* | FL (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FL, Osaka 1983;99SIEMENS;GA112-18-01
ASUM 2003 | FL (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=FL, ASUM 2003;99SIEMENS;GA112-01-01
Chitty | FL (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=FL, Chitty 1997;LN;33098-5 | 2 Sigma deviation of population ;DCM;121417
Jeanty | HL(GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=Humerus, Jeanty 1984;LN;11936-2 | 2 Sigma deviation of population ;DCM;121417
Osaka* | HL(GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HL, Osaka 1983;99SIEMENS;GA117-18-01
ASUM 2003 | HL(GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=HL, ASUM 2003;99SIEMENS;GA117-01-01
Hadlock | HC (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=HC, Hadlock 1984;LN;11932-1 | 2 Sigma deviation of population ;DCM;121417
Merz | HC (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC, Merz 1991;99SIEMENS;GA116-16-01 | 2 Sigma deviation of population ;DCM;121417
Lasser | HC (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC, Lasser 1993;99SIEMENS;GA116-14-01 | 2 Sigma deviation of population ;DCM;121417
Hansmann | HC (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC, Hansmann 1986;LN;33543-0 | 2 Sigma deviation of population ;DCM;121417
ASUM 2003 | HC (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=HC, ASUM 2003;99SIEMENS;GA116-01-01 | 2 Sigma deviation of population ;DCM;121417
Chitty | HC (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC measured, Chitty 1997;LN;33110-8 | 2 Sigma deviation of population ;DCM;121417
Jeanty | HC (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC, Jeanty 2015;99SIEMENS;GA102-19-01 | 2 Sigma deviation of population ;DCM;121417
ASUM 2003 | HC(c) (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=HC, ASUM 2003;99SIEMENS;GA116-01-01
Hadlock | HC(c) (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=HC, Hadlock 1984;LN;11932-1 | 2 Sigma deviation of population ;DCM;121417
Merz | HC(c) (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC, Merz 1991;99SIEMENS;GA116-16-01 | 2 Sigma deviation of population ;DCM;121417
Hansmann | HC(c) (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC, Hansmann 1986;LN;33543-0 | 2 Sigma deviation of population ;DCM;121417
Chitty(Derived) | HC(c) (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC derived, Chitty 1997;LN;33111-6 | 2 Sigma deviation of population ;DCM;121417
Jeanty | HC(c) (GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=HC, Jeanty 2015;99SIEMENS;GA102-19-01 | 2 Sigma deviation of population ;DCM;121417
Hill | TCD(GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=TCD, Hill 1990;LN;33134-8 | 2 Sigma deviation of population ;DCM;121417
ASUM 2003 | OFD(GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Equation;DCM;121420=OFD, ASUM 2003;99SIEMENS;GA119-01-01
Tokyo* | APTDxTTD(GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=APTDxTTD, Tokyo;99SIEMENS;GA105-20-01
Tokyo* | EFW(GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=Gestational Age by EFW, Tokyo 1987;99SIEMENS;GA101-20-01
Osaka* | EFW(GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=Gestational Age by EFW, Osaka 1983;99SIEMENS;GA101-18-01
JSUM 2003* | EFW(GA) | days | 0 | 1 | Gestational Age;LN;18185-9 | NUM |  |  | Fetus Characteristics;DCM;125015 |  | Table of Values;DCM;121424=Gestational Age by EFW, JSUM 2003;99SIEMENS;GA101-12-01
HC(c) | Head Circumference from BPD and OFD | mm | 3 | 1 | Head Circumference;LN;11984-2 | NUM |  |  | Fetal Measurements;DCM;125016-7 |  | Derivation;DCM;121401=Calculated;DCM;121428

## Vascular & GI Measurements
EnglishLabel | Description | DetailEnglishLabel | PossibleUnits | MaxInstances | DefaultPrecision | ConceptCoding | ValueType | ParentContainer | ParentLaterality | SRSection | SRSectionLaterality | VesselBranchModifier | TopographicalModifier | MeasurementOrientationModifier | AdditionalModifiers | AdditionalModifiers2
AAA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
AAA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
AAA (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA (Post Repair) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA Sac Length (Post Repair) | Abdominal Aortic Aneurysm Sac Length |  | cm mm | 5 | 2 | Abdominal Aortic Aneurysm Sac Length;99SIEMENS;AAASacLength | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA (Post Repair) | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
AAA dist from SMA | Abdominal Aortic Aneurysm distance from SMA |  | cm mm | 5 | 2 | Abdominal Aortic Aneurysm distance from SMA;99SIEMENS;AAADistFromSMA | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA Sac Length | Abdominal Aortic Aneurysm Sac Length |  | cm mm | 5 | 2 | Abdominal Aortic Aneurysm Sac Length;99SIEMENS;AAASacLength | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Rt ARA Origin | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt ARA Origin | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Dist ARA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ARA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ARA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ARA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ARA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ARA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ARA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ARA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ARA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ARA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ARA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist ARA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ARA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ARA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ARA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid ARA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ARA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ARA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ARA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ARA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ARA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ARA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ARA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ARA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ARA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ARA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid ARA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ARA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ARA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ARA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox ARA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ARA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ARA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ARA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ARA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ARA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ARA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ARA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ARA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ARA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ARA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox ARA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ARA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ARA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ARA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist ARA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ARA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ARA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ARA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ARA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ARA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ARA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ARA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ARA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ARA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ARA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist ARA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ARA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ARA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ARA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid ARA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ARA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ARA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ARA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ARA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ARA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ARA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ARA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ARA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ARA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ARA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid ARA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ARA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ARA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ARA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox ARA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ARA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ARA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ARA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ARA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ARA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ARA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ARA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ARA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ARA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ARA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox ARA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ARA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ARA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ARA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt ARA Origin | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ARA Origin | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ARA Origin | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ARA Origin | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ARA Origin | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ARA Origin | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ARA Origin | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ARA Origin | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ARA Origin | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ARA Origin | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ARA Origin | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt ARA Origin | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ARA Origin | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ARA Origin | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ARA Origin | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt ARA Origin | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt ARA Origin | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt ARA Origin | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt ARA Origin | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt ARA Origin | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt ARA Origin | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt ARA Origin | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt ARA Origin | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt ARA Origin | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt ARA Origin | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt ARA Origin | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt ARA Origin | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt ARA Origin | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ACA A1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A1 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt ACA A1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A1 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt ACA A2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A2 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt ACA A2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A2 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
ACoA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Anterior Communicating Artery;SRT;T-45530 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
ACoA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Anterior Communicating Artery;SRT;T-45530 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
ACoA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Anterior Communicating Artery;SRT;T-45530 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
ACoA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Anterior Communicating Artery;SRT;T-45530 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
ACoA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Anterior Communicating Artery;SRT;T-45530 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
ACoA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Anterior Communicating Artery;SRT;T-45530 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
ACoA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Anterior Communicating Artery;SRT;T-45530 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
ACoA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Anterior Communicating Artery;SRT;T-45530 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
ACoA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Anterior Communicating Artery;SRT;T-45530 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
ACoA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist ATA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist ATA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid ATA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid ATA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox ATA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox ATA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist ATA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist ATA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid ATA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid ATA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox ATA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox ATA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist ATV 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATV 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATV 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATV 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid ATV 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATV 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATV 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATV 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox ATV 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATV 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATV 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATV 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist ATV 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATV 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATV 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATV 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid ATV 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATV 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATV 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATV 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox ATV 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATV 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATV 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATV 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist ATV 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATV 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATV 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ATV 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid ATV 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATV 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATV 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATV 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox ATV 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATV 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATV 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATV 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist ATV 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATV 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATV 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATV 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid ATV 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATV 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATV 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ATV 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox ATV 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATV 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATV 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Anterior Tibial Vein;SRT;T-49630 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ATV 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Aorta | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Aorta | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Aorta | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Aorta | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Aorta | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Aorta | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Aorta | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Aorta | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Aorta | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Aorta | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Aorta | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Aorta | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Aorta | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Aorta | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Aorta | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Aorta | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Aorta | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Aorta | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Aorta | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Aorta | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Aorta | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Aorta | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Aorta | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Aorta | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Aorta | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Aorta | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid Aorta | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Aorta | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Aorta | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Aorta | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Aorta | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Aorta | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Aorta | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Aorta | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Aorta | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Aorta | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Aorta | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Aorta | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Aorta | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Aorta | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Aorta | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Aorta | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Aorta | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Aorta | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Aorta | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Aorta (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Aorta (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid Aorta (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Aorta (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta (Post Repair) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta (Post Repair) | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Mid Aorta | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Prox Aorta | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Ao Bif | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Ao Bif | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Ao Bif (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif (Post Repair) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif (Post Repair) | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Ao Bif | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Dist Aorta | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Mid Aorta | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Prox Aorta | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Dist Ao Stent Graft | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Ao Stent Graft | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Ao Stent Graft | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Ao Stent Graft | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Ao Stent Graft | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Ao Stent Graft | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Ao Stent Graft | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Ao Stent Graft | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Ao Stent Graft | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Ao Stent Graft | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Ao Stent Graft | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Ao Stent Graft | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Ao Stent Graft | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Ao Stent Graft | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Ao Stent Graft | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Ao Stent Graft | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Ao Stent Graft | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Ao Stent Graft | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Ao Stent Graft | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Ao Stent Graft | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Ao Stent Graft | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Ao Stent Graft | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Ao Stent Graft | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Ao Stent Graft | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Ao Stent Graft | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Ao Stent Graft | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid Ao Stent Graft | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Ao Stent Graft | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Ao Stent Graft | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Ao Stent Graft | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Ao Stent Graft | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Ao Stent Graft | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Ao Stent Graft | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Ao Stent Graft | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Ao Stent Graft | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Ao Stent Graft | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Ao Stent Graft | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Ao Stent Graft | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Ao Stent Graft | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Ao Stent Graft | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Ao Stent Graft | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Ao Stent Graft | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Ao Stent Graft | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Ao Stent Graft | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Ao Stent Graft | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Ao Stent Graft | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Mid Ao Stent Graft | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Prox Ao Stent Graft | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Dist Ao Stent Graft | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Mid Ao Stent Graft | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Prox Ao Stent Graft | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Lt Mid Arcuate A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Inf Arcuate A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Mid Arcuate A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Arcuate A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Sup Arcuate A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Vert Atlas Loop | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert Atlas Loop | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert Atlas Loop | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert Atlas Loop | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert Atlas Loop | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert Atlas Loop | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert Atlas Loop | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert Atlas Loop | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert Atlas Loop | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert Atlas Loop | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Vert Atlas Loop | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert Atlas Loop | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert Atlas Loop | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert Atlas Loop | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert Atlas Loop | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert Atlas Loop | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert Atlas Loop | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert Atlas Loop | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert Atlas Loop | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert Atlas Loop | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Axillary A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Axillary A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Axillary A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Axillary A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Axillary A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Axillary A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Axillary A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Axillary A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Axillary A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Axillary A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Axillary A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Axillary A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Axillary A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Axillary A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Axillary A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Axillary A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Axillary A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Axillary A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Axillary A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Axillary A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Axillary A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Axillary A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Axillary A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Axillary A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Axillary A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Axillary A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Axillary A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Axillary A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Axillary A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Axillary A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Axillary A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Axillary A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Axillary A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Axillary A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Axillary A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Axillary A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Axillary A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Axillary A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Axillary A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Axillary A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Axillary A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Axillary A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Axillary A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Axillary A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Axillary A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Axillary A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Axillary A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Axillary A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Axillary A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Axillary A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Axillary A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Axillary A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Axillary A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Axillary A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Axillary A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Axillary A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Axillary A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Axillary A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Axillary A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Axillary A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Axillary A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Axillary A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Axillary A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Axillary A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Axillary A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Axillary A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Axillary A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Axillary A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Axillary A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Axillary A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Axillary A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Axillary A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Axillary A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Axillary A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Axillary A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Axillary A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Axillary A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Axillary A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Axillary A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Axillary A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Axillary A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Axillary A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Axillary A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Axillary A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Axillary A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Axillary A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Axillary A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Axillary A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Axillary A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Axillary A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Axillary V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Axillary Vein;SRT;T-49110 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Axillary V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Axillary Vein;SRT;T-49110 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Axillary V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Axillary V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Axillary Vein;SRT;T-49110 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Axillary V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Axillary Vein;SRT;T-49110 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Axillary V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Axillary V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Axillary Vein;SRT;T-49110 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Axillary V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Axillary Vein;SRT;T-49110 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Axillary V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Axillary V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Axillary Vein;SRT;T-49110 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Axillary V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Axillary Vein;SRT;T-49110 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Axillary V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Axillary V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Axillary Vein;SRT;T-49110 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Axillary V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Axillary Vein;SRT;T-49110 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Axillary V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Axillary V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Axillary Vein;SRT;T-49110 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Axillary V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Axillary Vein;SRT;T-49110 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Axillary V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Basilar A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Basilar A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Basilar A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Basilar A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Basilar A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Basilar A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Basilar A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Basilar A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Basilar A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Basilar A | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid Basilar A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Basilar A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Basilar A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Basilar A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Basilar A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Basilar A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Basilar A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Basilar A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Basilar A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Basilar A | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Basilar A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Basilar A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Basilar A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Basilar A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Basilar A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Basilar A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Basilar A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Basilar A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Basilar A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Basilar A | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Basilic V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Basilic Vein;SRT;T-48052 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Basilic V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Basilic Vein;SRT;T-48052 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Basilic V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Basilic V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Basilic Vein;SRT;T-48052 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Basilic V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Basilic Vein;SRT;T-48052 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Basilic V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Basilic V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Basilic Vein;SRT;T-48052 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Basilic V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Basilic Vein;SRT;T-48052 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Basilic V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Basilic V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Basilic Vein;SRT;T-48052 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Basilic V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Basilic Vein;SRT;T-48052 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Basilic V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Basilic V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Basilic Vein;SRT;T-48052 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Basilic V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Basilic Vein;SRT;T-48052 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Basilic V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Basilic V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Basilic Vein;SRT;T-48052 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Basilic V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Basilic Vein;SRT;T-48052 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Basilic V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Boyd's Perf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Boyd's Perforating Vein;SRT;T-49424 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Boyd's Perf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Boyd's Perforating Vein;SRT;T-49424 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Boyd's Perf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Boyd's Perforating Vein;SRT;T-49424 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Boyd's Perf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Boyd's Perf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Boyd's Perforating Vein;SRT;T-49424 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Boyd's Perf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Boyd's Perforating Vein;SRT;T-49424 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Boyd's Perf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Boyd's Perforating Vein;SRT;T-49424 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Boyd's Perf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Brach A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Brach A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Brach A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Brach A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Brach A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Brach A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Brach A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Brach A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Brach A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Brach A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Brach A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Brach A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Brach V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Brach V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Brach V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Brach V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Brach V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Brach V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Brach V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Brach V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Brach V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Brach V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Brach V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Brach V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Brach V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Brach V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Brachial Vein;SRT;T-49350 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Brach V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Bulb | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Bulb | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Bulb | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Bulb | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Celiac Origin | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Celiac Origin | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin, Postprandial | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Celiac Origin, Postprandial | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Celiac Origin, Postprandial | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Celiac Origin, Postprandial | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Celiac Origin, Postprandial | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Celiac Origin, Postprandial | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Celiac Origin, Postprandial | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Celiac Origin, Postprandial | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Celiac Origin, Postprandial | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Celiac Origin, Postprandial | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Celiac Origin, Postprandial | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Celiac Origin, Postprandial | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Celiac Origin, Postprandial | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Celiac Origin, Postprandial | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Celiac Origin, Postprandial | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Lt Dist Cephalic V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Cephalic Vein;SRT;T-49240 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Cephalic V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Cephalic Vein;SRT;T-49240 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Cephalic V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Cephalic V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Cephalic Vein;SRT;T-49240 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Cephalic V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Cephalic Vein;SRT;T-49240 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Cephalic V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Cephalic V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Cephalic Vein;SRT;T-49240 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Cephalic V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Cephalic Vein;SRT;T-49240 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Cephalic V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Cephalic V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Cephalic Vein;SRT;T-49240 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Cephalic V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Cephalic Vein;SRT;T-49240 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Cephalic V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Cephalic V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Cephalic Vein;SRT;T-49240 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Cephalic V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Cephalic Vein;SRT;T-49240 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Cephalic V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Cephalic V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Cephalic Vein;SRT;T-49240 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Cephalic V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Cephalic Vein;SRT;T-49240 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Cephalic V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist MCA M1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M1 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid MCA M1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M1 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox MCA M1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M1 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist MCA M1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M1 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid MCA M1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M1 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox MCA M1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M1 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist MCA M2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M2 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid MCA M2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M2 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox MCA M2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M2 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist MCA M2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M2 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid MCA M2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid MCA M2 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox MCA M2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox MCA M2 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Cockett's Perf 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Cockett's Perf 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Cockett's Perf 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Cockett's Perf 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Cockett's Perf 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Cockett's Perf 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Cockett's Perf 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Cockett's Perf 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Cockett's Perf 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Cockett's Perf 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Cockett's Perf 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Cockett's Perf 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Cockett's Perf 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Cockett's Perf 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Cockett's Perf 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Cockett's Perf 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Cockett's Perf 3 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Cockett's Perf 3 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Cockett's Perf 3 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Cockett's Perf 3 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Cockett's Perf 3 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Cockett's Perf 3 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Cockett's Perf 3 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Cockett's Perforating Vein;SRT;T-49426 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Cockett's Perf 3 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
CBD | Common Bile Duct Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Common Bile Duct;SRT;T-64500 |  | Finding Site;SRT;G-C0E3=Gallbladder;SRT;T-63000
Lt Dist CCA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist CCA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid CCA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid CCA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox CCA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox CCA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist CCA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist CCA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid CCA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid CCA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox CCA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox CCA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt CFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt CFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist CFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist CFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid CFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid CFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox CFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox CFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt CFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt CFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Dist CFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist CFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid CFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid CFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox CFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox CFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt CFA Pseudoaneurysm | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt CFA Pseudoaneurysm | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt CFA Pseudoaneurysm | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt CFA Pseudoaneurysm | Doppler Angle Property of PS | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt CFA Pseudoaneurysm | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt CFA Pseudoaneurysm | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt CFA Pseudoaneurysm | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt CFA Pseudoaneurysm | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt CFA Pseudoaneurysm | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt CFA Pseudoaneurysm | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt CFA Pseudoaneurysm | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt CFA Pseudoaneurysm | Doppler Angle Property of PS | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt CFA Pseudoaneurysm | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt CFA Pseudoaneurysm | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt CFA Pseudoaneurysm | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt CFA Pseudoaneurysm | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt CFA Pseudoaneurysm D1 | Distance 1 |  | cm mm | 5 | 2 | Distance 1;99SIEMENS;Distance1 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt CFA Pseudoaneurysm D1 | Distance 1 |  | cm mm | 5 | 2 | Distance 1;99SIEMENS;Distance1 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt CFA Pseudoaneurysm D2 | Distance 2 |  | cm mm | 5 | 2 | Distance 2;99SIEMENS;Distance2 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt CFA Pseudoaneurysm D2 | Distance 2 |  | cm mm | 5 | 2 | Distance 2;99SIEMENS;Distance2 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt CFA Pseudoaneurysm D3 | Distance 3 |  | cm mm | 5 | 2 | Distance 3;99SIEMENS;Distance3 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt CFA Pseudoaneurysm D3 | Distance 3 |  | cm mm | 5 | 2 | Distance 3;99SIEMENS;Distance3 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Dist CFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Dist CFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Dist CFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Dist CFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Dist CFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Dist CFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Dist CFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist CFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Dist CFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Dist CFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Dist CFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Dist CFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Dist CFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Dist CFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Dist CFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid CFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox CFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Dist CFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Dist CFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Dist CFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Dist CFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Dist CFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist CFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid CFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Mid CFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Mid CFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Mid CFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Mid CFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid CFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox CFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Prox CFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Prox CFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Prox CFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Prox CFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox CFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt CFV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist CFV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CFV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CFV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CFV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid CFV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox CFV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt CFV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist CFV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid CFV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CFV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox CFV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Common Femoral Vein;SRT;G-035B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CFV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
CHA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
CHA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
CHA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
CHA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
CHA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
CHA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
CHA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
CHA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
CHA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
CHA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
CHA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
CHA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
CHA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
CHA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
CHA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Dist CHA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist CHA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist CHA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist CHA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist CHA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist CHA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist CHA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist CHA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist CHA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist CHA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist CHA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist CHA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist CHA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist CHA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist CHA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid CHA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid CHA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid CHA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid CHA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid CHA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid CHA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid CHA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid CHA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid CHA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid CHA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid CHA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid CHA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid CHA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid CHA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid CHA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox CHA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox CHA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox CHA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox CHA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox CHA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox CHA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox CHA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox CHA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox CHA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox CHA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox CHA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox CHA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox CHA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox CHA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox CHA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
CHD | Common Hepatic Duct Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Common Hepatic Duct;SRT;T-64100 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000
Lt CIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CIA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt CIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist CIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CIA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist CIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid CIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CIA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid CIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox CIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CIA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox CIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt CIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CIA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt CIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Dist CIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CIA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist CIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid CIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CIA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid CIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox CIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CIA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox CIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt CIA (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt CIA (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist CIA (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid CIA (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox CIA (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt CIA (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist CIA (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid CIA (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox CIA (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA (Post Repair) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA (Post Repair) | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt CIA Aneurysm | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt CIA Aneurysm | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt CIA Aneurysm (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt CIA Aneurysm (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm (Post Repair) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm (Post Repair) | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Lt Dist CIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Lt Mid CIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Lt Prox CIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Rt CIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Rt Dist CIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Rt Mid CIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Rt Prox CIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Lt Dist CIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist CIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist CIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist CIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist CIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist CIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist CIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist CIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist CIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist CIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist CIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist CIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist CIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist CIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist CIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid CIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox CIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist CIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid CIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox CIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist CIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist CIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Lt Dist CIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Lt Mid CIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Lt Prox CIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Rt CIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Rt Dist CIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Rt Mid CIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Rt Prox CIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Lt CIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist CIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid CIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox CIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt CIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Dist CIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid CIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox CIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt CIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist CIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid CIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox CIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt CIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Dist CIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid CIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox CIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt CIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist CIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid CIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox CIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt CIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist CIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid CIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox CIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt CIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist CIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid CIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox CIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt CIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist CIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid CIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox CIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Common Iliac Vein;SRT;T-48920 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Deep Brach A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Deep Brach A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Deep Brach A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Deep Brach A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Deep Brach A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Deep Brach A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Deep Brach A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Deep Brach A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Deep Brach A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Deep Brach A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Deep Brach A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Deep Brach A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Deep Brach A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Deep Brach A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Deep Brach A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Deep Brach A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Deep Brach A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Deep Brach A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Deep Brach A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Deep Brach A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Deep Brach A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Deep Brach A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Deep Brach A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Deep Brach A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Deep Brach A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Deep Brach A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Deep Brach A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Deep Brach A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Deep Brach A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Deep Brach A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Deep Brach A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Deep Brach A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Deep Brach A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Deep Brach A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Deep Brach A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Deep Brach A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Deep Brach A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Deep Brach A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Deep Brach A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Deep Brach A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Deep Brach A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Deep Brach A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Deep Brach A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Deep Brach A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Deep Brach A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Deep Brach A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Deep Brach A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Deep Brach A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Deep Brach A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Deep Brach A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Deep Brach A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Deep Brach A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Deep Brach A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Deep Brach A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Deep Brach A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Deep Brach A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Deep Brach A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Deep Brach A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Deep Brach A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Deep Brach A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Deep Brach A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Deep Brach A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Deep Brach A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Deep Brach A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Deep Brach A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Deep Brach A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Deep Brach A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Deep Brach A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Deep Brach A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Deep Brach A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Deep Brach A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Deep Brach A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Deep Brach A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Deep Brach A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Deep Brach A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Deep Brach A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Deep Brach A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Deep Brach A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Deep Brach A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Deep Brach A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Deep Brach A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Deep Brach A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Deep Brach A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Deep Brach A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Deep Brach A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Deep Brach A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Deep Brach A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Deep Brach A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Deep Brach A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Deep Brach A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Deep Fem V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Deep Fem V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Deep Fem V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Deep Fem V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Deep Fem V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Deep Fem V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Deep Fem V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Deep Fem V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Deep Fem V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Deep Fem V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Deep Fem V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Deep Fem V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Deep Fem V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Deep Fem V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Deep Fem V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Deep Fem V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Deep Fem V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Deep Fem V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Deep Fem V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Deep Fem V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Deep Fem V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Deep Fem V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Deep Fem V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Profunda Femoris Vein;SRT;T-49660 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Deep Fem V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dodd Perf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Dodd Perforator;99SIEMENS;DoddPerforator |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dodd Perf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Dodd Perforator;99SIEMENS;DoddPerforator |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dodd Perf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Dodd Perforator;99SIEMENS;DoddPerforator |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dodd Perf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dodd Perf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Dodd Perforator;99SIEMENS;DoddPerforator |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Dodd Perf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Dodd Perforator;99SIEMENS;DoddPerforator |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Dodd Perf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Dodd Perforator;99SIEMENS;DoddPerforator |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Dodd Perf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt DPA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DPA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DPA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DPA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DPA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DPA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DPA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt DPA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DPA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DPA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DPA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DPA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DPA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DPA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DPA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt DPA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DPA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DPA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DPA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DPA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DPA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DPA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt DPA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DPA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DPA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DPA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DPA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DPA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DPA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DPA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Ejaculatory Duct | Ejaculatory Duct Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Ejaculatory Duct;SRT;T-93070 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000
Rt Ejaculatory Duct | Ejaculatory Duct Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Ejaculatory Duct;SRT;T-93070 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000
Lt Epididymal A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Epididymal A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Epididymal A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Epididymal A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Epididymal A | End Diastoic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Epididymal A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Epididymal A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Epididymal A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Epididymal A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Epididymal A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Epididymal A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Epididymal A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Epididymal A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Epididymal A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Epididymal A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Epididymal A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Epididymal A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Epididymal A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Epididymal A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Epididymal A | End Diastoic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Epididymal A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Epididymal A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Epididymal A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Epididymal A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Epididymal A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Epididymal A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Epididymal A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Epididymal A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Epididymal A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Epididymal A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Epididymal V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Epididymal Vein;99SIEMENS;EpididymalV |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Epididymal V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Epididymal Vein;99SIEMENS;EpididymalV |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Epididymal V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Epididymal Vein;99SIEMENS;EpididymalV |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Epididymal V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Epididymal V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Epididymal Vein;99SIEMENS;EpididymalV |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Epididymal V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Epididymal V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Epididymal Vein;99SIEMENS;EpididymalV |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Epididymal V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Epididymal V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Epididymal Vein;99SIEMENS;EpididymalV |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Epididymal V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Epididymis | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Epididymis;SRT;T-95000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Epididymis | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Epididymis;SRT;T-95000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Epididymis | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Epididymis;SRT;T-95000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Epididymis | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Epididymis;SRT;T-95000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Epididymis | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Epididymis;SRT;T-95000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Epididymis | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Epididymis;SRT;T-95000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist ECA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist ECA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt ECA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt ECA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Mid ECA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid ECA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox ECA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox ECA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist ECA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist ECA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt ECA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt ECA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Mid ECA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid ECA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox ECA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox ECA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist ECA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ECA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt ECA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Mid ECA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox ECA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist ECA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt ECA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ECA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Mid ECA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ECA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox ECA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ECA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist EIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist EIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt EIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt EIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Mid EIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid EIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox EIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox EIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist EIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist EIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt EIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt EIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Mid EIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid EIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox EIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox EIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist EIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Lt EIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Lt Mid EIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Lt Prox EIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Rt Dist EIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Rt EIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Rt Mid EIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Rt Prox EIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Lt Dist EIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Dist EIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Dist EIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Dist EIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Dist EIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Dist EIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Dist EIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Dist EIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Dist EIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Dist EIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Dist EIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist EIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Dist EIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Dist EIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Dist EIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid EIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox EIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist EIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid EIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox EIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Dist EIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Dist EIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Dist EIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Lt EIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Lt Mid EIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Lt Prox EIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Rt Dist EIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Rt EIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Rt Mid EIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Rt Prox EIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Lt Dist EIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt EIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid EIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox EIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist EIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt EIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid EIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox EIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist EIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt EIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid EIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox EIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist EIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt EIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt EIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid EIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid EIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox EIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | External Iliac Vein;SRT;T-48930 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox EIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Flexed %FHC | Femoral Head Coverage, Flexed |  | % | 1 | 0 | Percent Femoral Head Coverage Flexed;99SIEMENS;FHCFlexed | NUM | Hip;SRT;T-D2500 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Hip;SRT;T-D2500 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Flexed %FHC | Femoral Head Coverage, Flexed |  | % | 1 | 0 | Percent Femoral Head Coverage Flexed;99SIEMENS;FHCFlexed | NUM | Hip;SRT;T-D2500 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Hip;SRT;T-D2500 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Stressed %FHC | Femoral Head Coverage, Stressed |  | % | 1 | 0 | Percent Femoral Head Coverage Stressed;99SIEMENS;FHCStressed | NUM | Hip;SRT;T-D2500 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Hip;SRT;T-D2500 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Stressed %FHC | Femoral Head Coverage, Stressed |  | % | 1 | 0 | Percent Femoral Head Coverage Stressed;99SIEMENS;FHCStressed | NUM | Hip;SRT;T-D2500 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Hip;SRT;T-D2500 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist Fem V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Fem V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Fem V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Fem V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Fem V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Fem V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Fem V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Fem V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Fem V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Fem V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Fem V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Fem V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Fem V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Fem V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Fem V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Fem V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Fem V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Fem V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Fem V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Fem V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Fem V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Fem V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Fem V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Femoral Vein;SRT;T-49410 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Fem V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
GB | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Gallbladder;SRT;T-63000 |  | Finding Site;SRT;G-C0E3=GallBladder;SRT;T-63000
GB | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Gallbladder;SRT;T-63000 |  | Finding Site;SRT;G-C0E3=GallBladder;SRT;T-63000
GB Wall | Wall Thickness |  | cm mm | 5 | 2 | Gallbladder Wall Thickness;99SIEMENS;GBWallThk | NUM | Gallbladder;SRT;T-63000 |  | Finding Site;SRT;G-C0E3=GallBladder;SRT;T-63000
GB | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Gallbladder;SRT;T-63000 |  | Finding Site;SRT;G-C0E3=GallBladder;SRT;T-63000
Lt Dist Med Gastroc V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Med Gastroc V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Med Gastroc V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Med Gastroc V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Med Gastroc V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Med Gastroc V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Med Gastroc V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Med Gastroc V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Med Gastroc V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Med Gastroc V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Med Gastroc V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Med Gastroc V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Med Gastroc V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Med Gastroc V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Med Gastroc V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Med Gastroc V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Med Gastroc V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Med Gastroc V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Med Gastroc V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Med Gastroc V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Med Gastroc V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Med Gastroc V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Med Gastroc V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Med Gastroc V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Med Gastroc V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Med Gastroc V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Med Gastroc V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Med Gastroc V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Med Gastroc V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Med Gastroc V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Med Gastroc V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Med Gastroc V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Med Gastroc V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Med Gastroc V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Med Gastroc V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Med Gastroc V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Med Gastroc V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Med Gastroc V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Med Gastroc V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Med Gastroc V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Med Gastroc V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Med Gastroc V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Med Gastroc V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Med Gastroc V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Med Gastroc V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Med Gastroc V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Med Gastroc V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Med Gastroc V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Anast 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 1 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Anast 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Prox Anast 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 1 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Anast 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Dist Graft 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 1 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Graft 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Mid Graft 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 1 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid Graft 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Prox Graft 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 1 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Graft 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Dist Inflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 1 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Inflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Inflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 1 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid Inflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Inflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 1 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Inflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Outflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 1 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Outflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Outflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 1 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid Outflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Outflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 1 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Outflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Anast 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 2 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Anast 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Dist Anast 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Prox Anast 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 2 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Anast 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Prox Anast 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Dist Graft 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 2 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Graft 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Dist Graft 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Mid Graft 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 2 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid Graft 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Mid Graft 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Prox Graft 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 2 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Graft 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Prox Graft 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Dist Inflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 2 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Inflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Inflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Inflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 2 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid Inflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Inflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Inflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 2 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Inflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Inflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Outflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 2 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Outflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Outflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Outflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 2 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid Outflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Outflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Outflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 2 | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Outflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Outflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt GSV: Ankle | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Great Saphenous Vein;SRT;T-49530 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Ankle;SRT;T-15750
Lt GSV: Ankle | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Great Saphenous Vein;SRT;T-49530 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Ankle;SRT;T-15750
Lt GSV: Ankle | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Great Saphenous Vein;SRT;T-49530 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Ankle;SRT;T-15750
Lt GSV: Ankle | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt GSV: Ankle | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Great Saphenous Vein;SRT;T-49530 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Ankle;SRT;T-15750
Rt GSV: Ankle | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Great Saphenous Vein;SRT;T-49530 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Ankle;SRT;T-15750
Rt GSV: Ankle | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Great Saphenous Vein;SRT;T-49530 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Ankle;SRT;T-15750
Rt GSV: Ankle | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt GSV: Mid Calf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109
Lt GSV: Mid Calf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109
Lt GSV: Mid Calf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109
Lt GSV: Mid Calf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt GSV: Mid Calf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109
Rt GSV: Mid Calf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109
Rt GSV: Mid Calf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109
Rt GSV: Mid Calf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt GSV: Knee | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Great Saphenous Vein;SRT;T-49530 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Knee;SRT;T-D9200
Lt GSV: Knee | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Great Saphenous Vein;SRT;T-49530 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Knee;SRT;T-D9200
Lt GSV: Knee | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Great Saphenous Vein;SRT;T-49530 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Knee;SRT;T-D9200
Lt GSV: Knee | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt GSV: Knee | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Great Saphenous Vein;SRT;T-49530 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Knee;SRT;T-D9200
Rt GSV: Knee | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Great Saphenous Vein;SRT;T-49530 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Knee;SRT;T-D9200
Rt GSV: Knee | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Great Saphenous Vein;SRT;T-49530 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Knee;SRT;T-D9200
Rt GSV: Knee | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt GSV: Lower Calf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115
Lt GSV: Lower Calf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115
Lt GSV: Lower Calf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115
Lt GSV: Lower Calf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt GSV: Lower Calf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115
Rt GSV: Lower Calf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115
Rt GSV: Lower Calf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115
Rt GSV: Lower Calf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt GSV: Lower Thigh | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115
Lt GSV: Lower Thigh | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115
Lt GSV: Lower Thigh | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115
Lt GSV: Lower Thigh | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt GSV: Lower Thigh | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115
Rt GSV: Lower Thigh | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115
Rt GSV: Lower Thigh | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115
Rt GSV: Lower Thigh | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt GSV: Mid Thigh | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109
Lt GSV: Mid Thigh | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109
Lt GSV: Mid Thigh | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109
Lt GSV: Mid Thigh | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt GSV: Mid Thigh | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109
Rt GSV: Mid Thigh | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109
Rt GSV: Mid Thigh | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109
Rt GSV: Mid Thigh | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt GSV: Upper Calf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116
Lt GSV: Upper Calf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116
Lt GSV: Upper Calf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116
Lt GSV: Upper Calf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt GSV: Upper Calf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116
Rt GSV: Upper Calf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116
Rt GSV: Upper Calf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Great Saphenous Vein of Calf;SRT;R-1025A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116
Rt GSV: Upper Calf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt GSV: Upper Thigh | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116
Lt GSV: Upper Thigh | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116
Lt GSV: Upper Thigh | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116
Lt GSV: Upper Thigh | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt GSV: Upper Thigh | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116
Rt GSV: Upper Thigh | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116
Rt GSV: Upper Thigh | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Great Saphenous Vein of Thigh;SRT;R-10259 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116
Rt GSV: Upper Thigh | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist LHA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LHA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LHA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LHA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LHA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LHA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LHA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LHA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LHA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LHA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LHA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist LHA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LHA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LHA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LHA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid LHA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LHA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LHA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LHA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LHA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LHA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LHA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LHA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LHA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LHA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LHA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid LHA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LHA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LHA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LHA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox LHA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LHA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LHA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LHA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LHA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LHA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LHA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LHA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LHA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LHA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LHA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox LHA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LHA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LHA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LHA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist RHA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RHA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RHA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RHA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RHA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RHA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RHA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RHA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RHA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RHA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RHA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist RHA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RHA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RHA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RHA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid RHA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RHA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RHA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RHA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RHA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RHA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RHA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RHA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RHA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RHA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RHA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid RHA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RHA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RHA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RHA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox RHA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RHA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RHA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RHA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RHA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RHA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RHA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RHA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RHA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RHA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RHA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox RHA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RHA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RHA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RHA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist LHV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Left Hepatic Vein;SRT;T-48727 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
LHV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Left Hepatic Vein;SRT;T-48727 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Mid LHV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Left Hepatic Vein;SRT;T-48727 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox LHV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Left Hepatic Vein;SRT;T-48727 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist MHV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Middle Hepatic Vein;SRT;T-48726 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
MHV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Middle Hepatic Vein;SRT;T-48726 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Mid MHV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Middle Hepatic Vein;SRT;T-48726 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox MHV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Middle Hepatic Vein;SRT;T-48726 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist RHV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Right Hepatic Vein;SRT;T-48725 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
RHV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Right Hepatic Vein;SRT;T-48725 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Mid RHV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Right Hepatic Vein;SRT;T-48725 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox RHV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Right Hepatic Vein;SRT;T-48725 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist LHV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Left Hepatic Vein;SRT;T-48727 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LHV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
LHV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Left Hepatic Vein;SRT;T-48727 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
LHV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid LHV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Left Hepatic Vein;SRT;T-48727 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LHV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox LHV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Left Hepatic Vein;SRT;T-48727 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LHV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist MHV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Middle Hepatic Vein;SRT;T-48726 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist MHV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
MHV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Middle Hepatic Vein;SRT;T-48726 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
MHV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid MHV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Middle Hepatic Vein;SRT;T-48726 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid MHV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox MHV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Middle Hepatic Vein;SRT;T-48726 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox MHV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist RHV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Right Hepatic Vein;SRT;T-48725 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RHV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
RHV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Right Hepatic Vein;SRT;T-48725 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
RHV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid RHV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Right Hepatic Vein;SRT;T-48725 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RHV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox RHV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Right Hepatic Vein;SRT;T-48725 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RHV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist LHV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Left Hepatic Vein;SRT;T-48727 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LHV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
LHV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Left Hepatic Vein;SRT;T-48727 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
LHV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid LHV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Left Hepatic Vein;SRT;T-48727 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LHV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox LHV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Left Hepatic Vein;SRT;T-48727 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LHV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist MHV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Middle Hepatic Vein;SRT;T-48726 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist MHV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
MHV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Middle Hepatic Vein;SRT;T-48726 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
MHV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid MHV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Middle Hepatic Vein;SRT;T-48726 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid MHV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox MHV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Middle Hepatic Vein;SRT;T-48726 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox MHV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist RHV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Right Hepatic Vein;SRT;T-48725 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RHV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
RHV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Right Hepatic Vein;SRT;T-48725 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
RHV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid RHV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Right Hepatic Vein;SRT;T-48725 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RHV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox RHV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Right Hepatic Vein;SRT;T-48725 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RHV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Hunterian Perf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Hunterian Perforating Vein;SRT;T-4942A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Hunterian Perf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Hunterian Perforating Vein;SRT;T-4942A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Hunterian Perf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Hunterian Perforating Vein;SRT;T-4942A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Hunterian Perf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Hunterian Perf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Hunterian Perforating Vein;SRT;T-4942A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Hunterian Perf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Hunterian Perforating Vein;SRT;T-4942A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Hunterian Perf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Hunterian Perforating Vein;SRT;T-4942A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Hunterian Perf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist IMA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist IMA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist IMA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist IMA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist IMA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist IMA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist IMA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist IMA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist IMA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist IMA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist IMA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist IMA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist IMA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist IMA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist IMA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid IMA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid IMA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid IMA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid IMA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid IMA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid IMA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid IMA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid IMA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid IMA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid IMA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid IMA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid IMA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid IMA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid IMA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid IMA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox IMA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox IMA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox IMA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox IMA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox IMA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox IMA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox IMA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox IMA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox IMA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox IMA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox IMA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox IMA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox IMA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox IMA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox IMA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
IMA Origin | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
IMA Origin | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
IMA Origin | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
IMA Origin | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
IMA Origin | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
IMA Origin | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
IMA Origin | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
IMA Origin | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
IMA Origin | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
IMA Origin | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
IMA Origin | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
IMA Origin | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
IMA Origin | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
IMA Origin | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
IMA Origin | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
IMA Origin (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
IMA Origin (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
IMA Origin (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
IMA Origin (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
IMA Origin (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
IMA Origin (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
IMA Origin (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
IMA Origin (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
IMA Origin (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
IMA Origin (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
IMA Origin (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
IMA Origin (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
IMA Origin (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
IMA Origin (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
IMA Origin (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Inf Thyroid A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Inf Thyroid A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Inf Thyroid A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Inf Thyroid A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Inf Thyroid A | End Diastoic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Inf Thyroid A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Inf Thyroid A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Inf Thyroid A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Inf Thyroid A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Inf Thyroid A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Inf Thyroid A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Inf Thyroid A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Inf Thyroid A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Inf Thyroid A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Inf Thyroid A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Thyroid A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Thyroid A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Thyroid A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Thyroid A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Thyroid A | End Diastoic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Thyroid A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Thyroid A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Inf Thyroid A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Thyroid A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Thyroid A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Thyroid A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Thyroid A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Thyroid A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Thyroid A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Thyroid A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Dist IVC | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Inferior Vena Cava;SRT;T-48710 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
IVC | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Inferior Vena Cava;SRT;T-48710 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Mid IVC | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Inferior Vena Cava;SRT;T-48710 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox IVC | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Inferior Vena Cava;SRT;T-48710 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist IVC | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Inferior Vena Cava;SRT;T-48710 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist IVC | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
IVC | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Inferior Vena Cava;SRT;T-48710 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
IVC | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid IVC | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Inferior Vena Cava;SRT;T-48710 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid IVC | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox IVC | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Inferior Vena Cava;SRT;T-48710 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox IVC | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist IVC | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Inferior Vena Cava;SRT;T-48710 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist IVC | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
IVC | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Inferior Vena Cava;SRT;T-48710 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
IVC | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid IVC | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Inferior Vena Cava;SRT;T-48710 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid IVC | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox IVC | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Inferior Vena Cava;SRT;T-48710 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox IVC | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Neck Diam | Infra-renal Neck Diameter |  | cm mm | 5 | 2 | Neck Diameter;99SIEMENS;NeckDiam | NUM | Infra-renal Aorta;SRT;T-42520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Neck Diam (Post Repair) | Infra-renal Neck Diameter |  | cm mm | 5 | 2 | Neck Diameter;99SIEMENS;NeckDiam | NUM | Infra-renal Aorta;SRT;T-42520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Neck Length (Post Repair) | Infra-renal Neck Length |  | cm mm | 5 | 2 | Neck Length;99SIEMENS;NeckLength | NUM | Infra-renal Aorta;SRT;T-42520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Neck Length | Infra-renal Neck Length |  | cm mm | 5 | 2 | Neck Length;99SIEMENS;NeckLength | NUM | Infra-renal Aorta;SRT;T-42520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Dist Innom A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Innom A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Innom A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Innom A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Innom A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Innom A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Innom A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Innom A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Innom A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Innom A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Innom A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Innom A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Innom A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Innom A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Innom A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Innom A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Innom A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Innom A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Innom A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Innom A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Innom A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Innom A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Innom A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Innom A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Innom A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Innom A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Innom A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Innom A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Innom A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Innom A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Mid Innom A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Innom A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Innom A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Innom A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Innom A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Innom A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Innom A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Innom A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Innom A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Innom A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid Innom A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Innom A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Innom A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Innom A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Innom A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Innom A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Innom A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Innom A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Innom A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Innom A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Innom A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Innom A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Innom A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Innom A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Innom A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Innom A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Innom A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Innom A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Innom A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Innom A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Innom V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Innominate Vein;SRT;T-48620 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Innom V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Innominate Vein;SRT;T-48620 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Innom V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Innom V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Innominate Vein;SRT;T-48620 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Innom V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Innominate Vein;SRT;T-48620 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Innom V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Innom V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Innominate Vein;SRT;T-48620 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Innom V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Innominate Vein;SRT;T-48620 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Innom V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Innom V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Innominate Vein;SRT;T-48620 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Innom V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Innominate Vein;SRT;T-48620 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Innom V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Innom V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Innominate Vein;SRT;T-48620 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Innom V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Innominate Vein;SRT;T-48620 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Innom V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Innom V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Innominate Vein;SRT;T-48620 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Innom V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Innominate Vein;SRT;T-48620 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Innom V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Interlobar A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Interlobar A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Interlobar A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Inf Interlobar A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Inf Interlobar A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Inf Interlobar A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Sup Interlobar A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Sup Interlobar A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Sup Interlobar A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Dist ICA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist ICA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid ICA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid ICA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox ICA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox ICA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist ICA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist ICA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid ICA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid ICA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox ICA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox ICA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist ICA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid ICA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox ICA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist ICA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid ICA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox ICA | Diameter % Stenosis | %Stenosis | % | 5 | 2 | Lumen Diameter Stenosis;SRT;R-101BB | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Area % Stenosis | %Stenosis | % | 5 | 2 | Lumen Area Stenosis;SRT;R-101BA | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Stenosis Distance 1 | D1 | cm mm | 5 | 2 | Vessel Intimal Diameter;SRT;R-1025C | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Stenosis Distance 2 | D2 | cm mm | 5 | 2 | Vessel lumen diameter;SRT;G-0364 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Stenosis Area 1 | A1 | cm² mm² | 5 | 2 | Vessel Intimal Cross-Sectional Area;SRT;R-1025D | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA | Stenosis Area 2 | A2 | cm² mm² | 5 | 2 | Vessel lumen cross-sectional area;SRT;G-0366 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt ICA Siphon | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ICA Siphon | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ICA Siphon | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ICA Siphon | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ICA Siphon | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ICA Siphon | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ICA Siphon | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ICA Siphon | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ICA Siphon | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ICA Siphon | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt ICA Siphon | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ICA Siphon | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ICA Siphon | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ICA Siphon | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ICA Siphon | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ICA Siphon | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ICA Siphon | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ICA Siphon | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ICA Siphon | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ICA Siphon | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist ICA-skull entrance | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA-skull entrance | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA-skull entrance | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA-skull entrance | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA-skull entrance | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA-skull entrance | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA-skull entrance | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA-skull entrance | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA-skull entrance | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA-skull entrance | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist ICA-skull entrance | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA-skull entrance | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA-skull entrance | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA-skull entrance | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA-skull entrance | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA-skull entrance | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA-skull entrance | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA-skull entrance | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA-skull entrance | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA-skull entrance | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist IIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist IIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt IIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IIA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt IIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Mid IIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid IIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox IIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox IIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist IIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist IIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt IIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IIA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt IIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Mid IIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid IIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox IIA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox IIA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist IIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Lt IIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Lt Mid IIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Lt Prox IIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Rt Dist IIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Rt IIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Rt Mid IIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Rt Prox IIA | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Lt Dist IIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist IIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist IIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist IIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist IIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist IIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist IIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist IIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist IIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist IIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist IIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist IIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist IIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist IIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist IIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid IIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox IIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist IIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid IIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox IIA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist IIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Lt Mid IIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Lt Prox IIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Rt Dist IIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist IIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist IIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Lt IIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Lt Mid IIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Lt Prox IIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Rt Dist IIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Rt IIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Rt Mid IIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Rt Prox IIA | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Lt Dist IIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt IIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Mid IIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox IIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist IIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt IIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Mid IIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox IIV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist IIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt IIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Mid IIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox IIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist IIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt IIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Mid IIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox IIV | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist IIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt IIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid IIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox IIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist IIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt IIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid IIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox IIV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist IIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt IIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid IIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox IIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist IIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt IIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid IIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox IIV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Internal Iliac Vein;SRT;T-48940 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IIV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist IJV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Jugular Vein;SRT;T-48170 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IJV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Internal Jugular Vein;SRT;T-48170 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IJV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt IJV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Jugular Vein;SRT;T-48170 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IJV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Internal Jugular Vein;SRT;T-48170 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IJV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid IJV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Jugular Vein;SRT;T-48170 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IJV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Internal Jugular Vein;SRT;T-48170 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IJV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox IJV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Jugular Vein;SRT;T-48170 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IJV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Internal Jugular Vein;SRT;T-48170 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IJV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist IJV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Jugular Vein;SRT;T-48170 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IJV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Internal Jugular Vein;SRT;T-48170 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IJV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt IJV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Jugular Vein;SRT;T-48170 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IJV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Internal Jugular Vein;SRT;T-48170 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt IJV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid IJV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Jugular Vein;SRT;T-48170 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IJV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Internal Jugular Vein;SRT;T-48170 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid IJV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox IJV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Jugular Vein;SRT;T-48170 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IJV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Internal Jugular Vein;SRT;T-48170 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox IJV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Isthmus | Isthmus Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Thyroid Isthmus;SRT;T-B6300 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Kidney | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Kidney | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Kidney | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Kidney | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Kidney | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Kidney | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist Lat Gastroc V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Lat Gastroc V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Lat Gastroc V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Lat Gastroc V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Lat Gastroc V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Lat Gastroc V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Lat Gastroc V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Lat Gastroc V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Lat Gastroc V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Lat Gastroc V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Lat Gastroc V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Lat Gastroc V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute |  |  |  | Vessel Branch;DCM;125101=Lateral;SRT;G-A104
Rt Dist Lat Gastroc V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Lat Gastroc V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Lat Gastroc V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Lat Gastroc V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Lat Gastroc V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Lat Gastroc V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Lat Gastroc V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Lat Gastroc V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Lat Gastroc V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Lat Gastroc V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Lat Gastroc V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Lat Gastroc V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Lat Gastroc V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Lat Gastroc V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Lat Gastroc V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Lat Gastroc V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Lat Gastroc V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Lat Gastroc V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Lat Gastroc V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Lat Gastroc V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Lat Gastroc V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Lat Gastroc V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Lat Gastroc V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Lat Gastroc V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Lat Gastroc V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Lat Gastroc V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Lat Gastroc V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Lat Gastroc V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Lat Gastroc V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Lat Gastroc V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Lat Gastroc V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Lat Gastroc V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Lat Gastroc V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Lat Gastroc V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Lat Gastroc V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Gastrocnemius Vein;SRT;T-4942D |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Lateral;SRT;G-A104 | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Lat Gastroc V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Liver | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000
Liver | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000
Liver | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Liver;SRT;T-62000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000
Lt LE Dist Anast 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Dist Anast 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Prox Anast 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Prox Anast 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Dist Anast 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Dist Anast 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Prox Anast 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Prox Anast 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Dist Graft 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Dist Graft 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Mid Graft 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Mid Graft 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Prox Graft 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Prox Graft 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Dist Graft 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Dist Graft 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Mid Graft 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Mid Graft 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Prox Graft 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Prox Graft 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Dist Inflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Dist Inflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Mid Inflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Mid Inflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Prox Inflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Prox Inflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Dist Inflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Dist Inflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Mid Inflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Mid Inflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Prox Inflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Prox Inflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Dist Outflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Dist Outflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Mid Outflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Mid Outflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Prox Outflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Prox Outflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Dist Outflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Dist Outflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Mid Outflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Mid Outflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Prox Outflow Vessel 1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 1 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Prox Outflow Vessel 1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 1 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 1 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 1 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 1 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Dist Anast 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Dist Anast 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Dist Anast 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Prox Anast 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Prox Anast 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Prox Anast 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Dist Anast 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Dist Anast 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Dist Anast 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Prox Anast 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Prox Anast 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Prox Anast 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Dist Graft 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Dist Graft 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Dist Graft 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Mid Graft 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Mid Graft 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Mid Graft 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Prox Graft 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Prox Graft 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Prox Graft 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Dist Graft 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Dist Graft 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Dist Graft 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Mid Graft 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Mid Graft 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Mid Graft 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Prox Graft 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Prox Graft 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Prox Graft 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Dist Inflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Dist Inflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Inflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Mid Inflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Mid Inflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Inflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Prox Inflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Prox Inflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Inflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Dist Inflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Dist Inflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Inflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Mid Inflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Mid Inflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Inflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Prox Inflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Prox Inflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Inflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Dist Outflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Dist Outflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Dist Outflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Mid Outflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Mid Outflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Mid Outflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Prox Outflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt LE Prox Outflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Prox Outflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Dist Outflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Dist Outflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Dist Outflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Mid Outflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Mid Outflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Mid Outflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Prox Outflow Vessel 2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 2 | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt LE Prox Outflow Vessel 2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 2 | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 2 | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 2 | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Prox Outflow Vessel 2 | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Isthmus Mass 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Thyroid Isthmus;SRT;T-B6300 | Identifier;DCM;125010=1
Liver Mass 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1
Lt Neck Mass 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=1
Rt Neck Mass 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=1
Prostate Lt Mass 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1
Prostate Rt Mass 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1
Lt Testis Mass 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Identifier;DCM;125010=1
Rt Testis Mass 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Identifier;DCM;125010=1
Lt Thyroid Mass 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1
Rt Thyroid Mass 1 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1
Isthmus Mass 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Thyroid Isthmus;SRT;T-B6300 | Identifier;DCM;125010=1
Liver Mass 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1
Lt Neck Mass 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=1
Rt Neck Mass 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=1
Prostate Lt Mass 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1
Prostate Rt Mass 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1
Lt Testis Mass 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Identifier;DCM;125010=1
Rt Testis Mass 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Identifier;DCM;125010=1
Lt Thyroid Mass 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1
Rt Thyroid Mass 1 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1
Lt Mass 1 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1
Lt Mass 1 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1
Lt Mass 1 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1
Rt Mass 1 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1
Rt Mass 1 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1
Rt Mass 1 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1
Isthmus Mass 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Thyroid Isthmus;SRT;T-B6300 | Identifier;DCM;125010=1
Liver Mass 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1
Lt Neck Mass 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=1
Rt Neck Mass 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=1
Prostate Lt Mass 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1
Prostate Rt Mass 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1
Lt Testis Mass 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Identifier;DCM;125010=1
Rt Testis Mass 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Identifier;DCM;125010=1
Lt Thyroid Mass 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1
Rt Thyroid Mass 1 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1
Liver Mass 10 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10
Liver Mass 10 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10
Liver Mass 10 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10
Isthmus Mass 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Thyroid Isthmus;SRT;T-B6300 | Identifier;DCM;125010=2
Liver Mass 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2
Lt Neck Mass 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=2
Rt Neck Mass 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=2
Prostate Lt Mass 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2
Prostate Rt Mass 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2
Lt Testis Mass 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Identifier;DCM;125010=2
Rt Testis Mass 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Identifier;DCM;125010=2
Lt Thyroid Mass 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2
Rt Thyroid Mass 2 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2
Isthmus Mass 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Thyroid Isthmus;SRT;T-B6300 | Identifier;DCM;125010=2
Liver Mass 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2
Lt Neck Mass 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=2
Rt Neck Mass 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=2
Prostate Lt Mass 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2
Prostate Rt Mass 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2
Lt Testis Mass 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Identifier;DCM;125010=2
Rt Testis Mass 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Identifier;DCM;125010=2
Lt Thyroid Mass 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2
Rt Thyroid Mass 2 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2
Lt Mass 2 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2
Lt Mass 2 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2
Lt Mass 2 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2
Rt Mass 2 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2
Rt Mass 2 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2
Rt Mass 2 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2
Isthmus Mass 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Thyroid Isthmus;SRT;T-B6300 | Identifier;DCM;125010=2
Liver Mass 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2
Lt Neck Mass 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=2
Rt Neck Mass 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=2
Prostate Lt Mass 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2
Prostate Rt Mass 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2
Lt Testis Mass 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Identifier;DCM;125010=2
Rt Testis Mass 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Identifier;DCM;125010=2
Lt Thyroid Mass 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2
Rt Thyroid Mass 2 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2
Isthmus Mass 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Thyroid Isthmus;SRT;T-B6300 | Identifier;DCM;125010=3
Liver Mass 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3
Lt Neck Mass 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=3
Rt Neck Mass 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=3
Prostate Lt Mass 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3
Prostate Rt Mass 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3
Lt Testis Mass 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Identifier;DCM;125010=3
Rt Testis Mass 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Identifier;DCM;125010=3
Lt Thyroid Mass 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3
Rt Thyroid Mass 3 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3
Isthmus Mass 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Thyroid Isthmus;SRT;T-B6300 | Identifier;DCM;125010=3
Liver Mass 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3
Lt Neck Mass 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=3
Rt Neck Mass 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=3
Prostate Lt Mass 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3
Prostate Rt Mass 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3
Lt Testis Mass 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Identifier;DCM;125010=3
Rt Testis Mass 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Identifier;DCM;125010=3
Lt Thyroid Mass 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3
Rt Thyroid Mass 3 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3
Lt Mass 3 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3
Lt Mass 3 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3
Lt Mass 3 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3
Rt Mass 3 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3
Rt Mass 3 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3
Rt Mass 3 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3
Isthmus Mass 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Thyroid Isthmus;SRT;T-B6300 | Identifier;DCM;125010=3
Liver Mass 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3
Lt Neck Mass 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=3
Rt Neck Mass 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=3
Prostate Lt Mass 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3
Prostate Rt Mass 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3
Lt Testis Mass 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Identifier;DCM;125010=3
Rt Testis Mass 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Identifier;DCM;125010=3
Lt Thyroid Mass 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3
Rt Thyroid Mass 3 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3
Isthmus Mass 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Thyroid Isthmus;SRT;T-B6300 | Identifier;DCM;125010=4
Liver Mass 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4
Lt Neck Mass 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=4
Rt Neck Mass 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=4
Prostate Lt Mass 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4
Prostate Rt Mass 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4
Lt Thyroid Mass 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4
Rt Thyroid Mass 4 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4
Isthmus Mass 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Thyroid Isthmus;SRT;T-B6300 | Identifier;DCM;125010=4
Liver Mass 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4
Lt Neck Mass 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=4
Rt Neck Mass 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=4
Prostate Lt Mass 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4
Prostate Rt Mass 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4
Lt Thyroid Mass 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4
Rt Thyroid Mass 4 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4
Lt Mass 4 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4
Lt Mass 4 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4
Lt Mass 4 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4
Rt Mass 4 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4
Rt Mass 4 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4
Rt Mass 4 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4
Isthmus Mass 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Thyroid Isthmus;SRT;T-B6300 | Identifier;DCM;125010=4
Liver Mass 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4
Lt Neck Mass 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=4
Rt Neck Mass 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=4
Prostate Lt Mass 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4
Prostate Rt Mass 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4
Lt Thyroid Mass 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4
Rt Thyroid Mass 4 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4
Isthmus Mass 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Thyroid Isthmus;SRT;T-B6300 | Identifier;DCM;125010=5
Liver Mass 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5
Lt Neck Mass 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=5
Rt Neck Mass 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=5
Prostate Lt Mass 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5
Prostate Rt Mass 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5
Lt Thyroid Mass 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5
Rt Thyroid Mass 5 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5
Isthmus Mass 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Thyroid Isthmus;SRT;T-B6300 | Identifier;DCM;125010=5
Liver Mass 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5
Lt Neck Mass 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=5
Rt Neck Mass 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=5
Prostate Lt Mass 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5
Prostate Rt Mass 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5
Lt Thyroid Mass 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5
Rt Thyroid Mass 5 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5
Lt Mass 5 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5
Lt Mass 5 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5
Lt Mass 5 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5
Rt Mass 5 | Strain Ratio | Strain Ratio | unitless | 5 | 2 | Strain Ratio;99SIEMENS;StrainRatio | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5
Rt Mass 5 | Strain Ratio Reference ROI | ROI 1 (Reference) | % | 5 | 3 | Reference ROI of Strain Ratio;99SIEMENS;StrainRatioRROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5
Rt Mass 5 | Strain Ratio Mass ROI | ROI 2 (Mass) | % | 5 | 3 | Mass ROI of Strain Ratio;99SIEMENS;StrainRatioMROI | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5
Isthmus Mass 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Thyroid Isthmus;SRT;T-B6300 | Identifier;DCM;125010=5
Liver Mass 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5
Lt Neck Mass 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=5
Rt Neck Mass 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Finding Site;SRT;G-C0E3=Neck;SRT;T-D1600 | Identifier;DCM;125010=5
Prostate Lt Mass 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5
Prostate Rt Mass 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5
Lt Thyroid Mass 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5
Rt Thyroid Mass 5 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5
Liver Mass 6 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6
Liver Mass 6 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6
Liver Mass 6 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6
Liver Mass 7 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7
Liver Mass 7 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7
Liver Mass 7 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7
Liver Mass 8 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8
Liver Mass 8 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8
Liver Mass 8 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8
Liver Mass 9 | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9
Liver Mass 9 | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9
Liver Mass 9 | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9
Lt Dist Ophthalmic A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ophthalmic A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ophthalmic A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ophthalmic A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ophthalmic A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ophthalmic A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ophthalmic A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ophthalmic A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ophthalmic A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ophthalmic A | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Ophthalmic A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ophthalmic A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ophthalmic A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ophthalmic A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ophthalmic A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ophthalmic A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ophthalmic A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ophthalmic A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ophthalmic A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ophthalmic A | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Ophthalmic A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ophthalmic A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ophthalmic A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ophthalmic A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ophthalmic A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ophthalmic A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ophthalmic A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ophthalmic A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ophthalmic A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ophthalmic A | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Ophthalmic A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ophthalmic A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ophthalmic A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ophthalmic A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ophthalmic A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ophthalmic A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ophthalmic A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ophthalmic A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ophthalmic A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ophthalmic A | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Ophthalmic A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ophthalmic A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ophthalmic A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ophthalmic A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ophthalmic A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ophthalmic A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ophthalmic A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ophthalmic A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ophthalmic A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ophthalmic A | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Ophthalmic A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ophthalmic A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ophthalmic A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ophthalmic A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ophthalmic A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ophthalmic A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ophthalmic A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ophthalmic A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ophthalmic A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ophthalmic A | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Pampiniform Plexus | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Pampiniform Plexus;SRT;T-48784 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Pampiniform Plexus | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Pampiniform Plexus;SRT;T-48784 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Pancreas Body | Pancreas Body Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Body of Pancreas;SRT;T-65200 |  | Finding Site;SRT;G-C0E3=Pancreas;SRT;T-65000
Pancreas Head | Pancreas Head Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Head of Pancreas;SRT;T-65100 |  | Finding Site;SRT;G-C0E3=Pancreas;SRT;T-65000
Pancreas Tail | Pancreas Tail Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Tail of Pancreas;SRT;T-65300 |  | Finding Site;SRT;G-C0E3=Pancreas;SRT;T-65000
Pancreatic Duct | Pancreatic Duct Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Pancreatic Duct;SRT;T-65010 |  | Finding Site;SRT;G-C0E3=Pancreas;SRT;T-65000
Lt Parathyroid | Lt Parathyroid Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Parathyroid;SRT;T-B7000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Parathyroid | Rt Parathyroid Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Parathyroid;SRT;T-B7000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Hip Angle | Hip alpha angle | α | ° | 1 | 0 | Hip alpha angle;99SIEMENS;HipAlphaAngle | NUM | Hip;SRT;T-D2500 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Hip;SRT;T-D2500 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Hip Angle | Hip beta angle | β | ° | 1 | 0 | Hip beta angle;99SIEMENS;HipBetaAngle | NUM | Hip;SRT;T-D2500 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Hip;SRT;T-D2500 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Hip Angle | Hip alpha angle | α | ° | 1 | 0 | Hip alpha angle;99SIEMENS;HipAlphaAngle | NUM | Hip;SRT;T-D2500 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Hip;SRT;T-D2500 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Hip Angle | Hip beta angle | β | ° | 1 | 0 | Hip beta angle;99SIEMENS;HipBetaAngle | NUM | Hip;SRT;T-D2500 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Hip;SRT;T-D2500 | Laterality;SRT;G-C171=Right;SRT;G-A100
Penile V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Penile Vein;SRT;T-48991 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000
Penile V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Penile Vein;SRT;T-48991 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000
Penile V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Penile V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Penile Vein;SRT;T-48991 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000
Penile V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Peroneal A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Peroneal A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Peroneal A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Peroneal A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Peroneal A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Peroneal A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Peroneal A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Peroneal A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Peroneal A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Peroneal A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Peroneal A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Peroneal A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Peroneal V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Peroneal V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Peroneal V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Peroneal V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Peroneal V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Peroneal V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal V 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Peroneal V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Peroneal V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Peroneal V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Peroneal V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Peroneal V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Peroneal V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Peroneal V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal V 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Peroneal Vein;SRT;T-49650 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Peroneal V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Pop A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Pop A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Pop A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Pop A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Pop A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Pop A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Pop A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Pop A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Pop A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Pop A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Pop A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Pop A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Pop A Aneurysm | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Pop A Aneurysm | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Pop A Aneurysm | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Pop A Aneurysm | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Pop A Aneurysm | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Pop A Aneurysm | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Pop A Aneurysm | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Pop A Aneurysm | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Pop A Aneurysm | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Pop A Aneurysm | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Pop A Aneurysm | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Pop A Aneurysm | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Pop A Aneurysm | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Pop A Aneurysm | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Pop A Aneurysm | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Pop A Aneurysm | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Pop A Aneurysm | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Pop A Aneurysm | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop A Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop A Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop A Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop A Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop A Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop A Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop A Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Pop A Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop A Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop A Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop A Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop A Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop A Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop A Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop A Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid Pop A Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid Pop A Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid Pop A Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid Pop A Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid Pop A Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid Pop A Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid Pop A Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Pop A Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid Pop A Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid Pop A Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid Pop A Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid Pop A Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid Pop A Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid Pop A Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid Pop A Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox Pop A Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox Pop A Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox Pop A Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox Pop A Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox Pop A Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox Pop A Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox Pop A Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Pop A Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox Pop A Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox Pop A Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox Pop A Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox Pop A Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox Pop A Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox Pop A Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox Pop A Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist Pop A Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist Pop A Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist Pop A Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist Pop A Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist Pop A Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist Pop A Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist Pop A Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Pop A Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist Pop A Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist Pop A Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist Pop A Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist Pop A Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist Pop A Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist Pop A Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist Pop A Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid Pop A Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid Pop A Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid Pop A Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid Pop A Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid Pop A Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid Pop A Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid Pop A Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Pop A Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid Pop A Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid Pop A Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid Pop A Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid Pop A Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid Pop A Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid Pop A Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid Pop A Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox Pop A Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox Pop A Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox Pop A Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox Pop A Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox Pop A Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox Pop A Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox Pop A Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Pop A Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox Pop A Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox Pop A Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox Pop A Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox Pop A Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox Pop A Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox Pop A Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox Pop A Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Pop V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Pop V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Pop V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Pop V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Pop V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Pop V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Popliteal Vein;SRT;T-49640 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Pop V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist MPV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Portal Vein;SRT;T-48810 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LPV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Left Main Branch of Portal Vein;SRT;T-4881F |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid LPV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Left Main Branch of Portal Vein;SRT;T-4881F |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox LPV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Left Main Branch of Portal Vein;SRT;T-4881F |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Mid MPV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Portal Vein;SRT;T-48810 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Portal V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Portal Vein;SRT;T-48810 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Prox MPV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Portal Vein;SRT;T-48810 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist RPV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Right Main Branch of Portal Vein;SRT;T-4882A |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid RPV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Right Main Branch of Portal Vein;SRT;T-4882A |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox RPV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Right Main Branch of Portal Vein;SRT;T-4882A |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist MPV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Portal Vein;SRT;T-48810 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist MPV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist LPV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Left Main Branch of Portal Vein;SRT;T-4881F |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LPV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid LPV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Left Main Branch of Portal Vein;SRT;T-4881F |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LPV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox LPV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Left Main Branch of Portal Vein;SRT;T-4881F |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LPV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid MPV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Portal Vein;SRT;T-48810 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid MPV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Portal V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Portal Vein;SRT;T-48810 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Portal V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox MPV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Portal Vein;SRT;T-48810 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox MPV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist RPV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Right Main Branch of Portal Vein;SRT;T-4882A |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RPV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid RPV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Right Main Branch of Portal Vein;SRT;T-4882A |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RPV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox RPV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Right Main Branch of Portal Vein;SRT;T-4882A |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RPV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist MPV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Portal Vein;SRT;T-48810 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist MPV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist LPV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Left Main Branch of Portal Vein;SRT;T-4881F |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist LPV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid LPV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Left Main Branch of Portal Vein;SRT;T-4881F |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid LPV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox LPV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Left Main Branch of Portal Vein;SRT;T-4881F |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox LPV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid MPV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Portal Vein;SRT;T-48810 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid MPV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Portal V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Portal Vein;SRT;T-48810 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Portal V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox MPV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Portal Vein;SRT;T-48810 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox MPV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist RPV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Right Main Branch of Portal Vein;SRT;T-4882A |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist RPV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid RPV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Right Main Branch of Portal Vein;SRT;T-4882A |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid RPV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox RPV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Right Main Branch of Portal Vein;SRT;T-4882A |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox RPV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Post-inj Cavernosal A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt Post-inj Cavernosal A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt Post-inj Cavernosal A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt Post-inj Cavernosal A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt Post-inj Cavernosal A | End Diastoic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt Post-inj Cavernosal A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt Post-inj Cavernosal A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Post-inj Cavernosal A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt Post-inj Cavernosal A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt Post-inj Cavernosal A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt Post-inj Cavernosal A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt Post-inj Cavernosal A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt Post-inj Cavernosal A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt Post-inj Cavernosal A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt Post-inj Cavernosal A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt Post-inj Cavernosal A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt Post-inj Cavernosal A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt Post-inj Cavernosal A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt Post-inj Cavernosal A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt Post-inj Cavernosal A | End Diastoic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt Post-inj Cavernosal A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt Post-inj Cavernosal A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Post-inj Cavernosal A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt Post-inj Cavernosal A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt Post-inj Cavernosal A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt Post-inj Cavernosal A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt Post-inj Cavernosal A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt Post-inj Cavernosal A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt Post-inj Cavernosal A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt Post-inj Cavernosal A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Post-inj Dorsal A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Post-inj Dorsal A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Post-inj Dorsal A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Post-inj Dorsal A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Post-inj Dorsal A | End Diastoic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Post-inj Dorsal A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Post-inj Dorsal A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Post-inj Dorsal A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Post-inj Dorsal A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Post-inj Dorsal A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Post-inj Dorsal A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Post-inj Dorsal A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Post-inj Dorsal A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Post-inj Dorsal A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Post-inj Dorsal A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt PCA P1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P1 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt PCA P1 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P1 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P1 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P1 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P1 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P1 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P1 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P1 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P1 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P1 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt PCA P2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P2 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt PCA P2 | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P2 | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P2 | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P2 | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P2 | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P2 | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P2 | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P2 | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P2 | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCA P2 | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt PCoA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCoA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCoA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCoA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCoA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCoA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCoA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCoA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCoA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCoA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt PCoA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCoA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCoA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCoA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCoA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCoA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCoA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCoA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCoA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt PCoA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist PTA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist PTA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid PTA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid PTA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox PTA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox PTA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist PTA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist PTA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid PTA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid PTA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox PTA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox PTA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist PTV 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTV 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTV 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTV 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid PTV 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTV 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTV 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTV 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox PTV 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTV 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTV 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTV 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist PTV 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTV 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTV 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTV 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid PTV 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTV 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTV 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTV 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox PTV 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTV 1 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTV 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTV 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist PTV 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTV 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTV 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist PTV 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid PTV 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTV 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTV 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid PTV 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox PTV 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTV 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTV 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox PTV 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist PTV 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTV 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTV 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist PTV 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid PTV 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTV 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTV 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid PTV 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox PTV 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTV 2 | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTV 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Posterior Tibial Vein;SRT;T-49620 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox PTV 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Pre-inj Cavernosal A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt Pre-inj Cavernosal A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt Pre-inj Cavernosal A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt Pre-inj Cavernosal A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt Pre-inj Cavernosal A | End Diastoic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt Pre-inj Cavernosal A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt Pre-inj Cavernosal A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Pre-inj Cavernosal A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt Pre-inj Cavernosal A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt Pre-inj Cavernosal A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt Pre-inj Cavernosal A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt Pre-inj Cavernosal A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt Pre-inj Cavernosal A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt Pre-inj Cavernosal A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt Pre-inj Cavernosal A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt Pre-inj Cavernosal A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt Pre-inj Cavernosal A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt Pre-inj Cavernosal A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt Pre-inj Cavernosal A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt Pre-inj Cavernosal A | End Diastoic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt Pre-inj Cavernosal A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt Pre-inj Cavernosal A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Pre-inj Cavernosal A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt Pre-inj Cavernosal A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt Pre-inj Cavernosal A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt Pre-inj Cavernosal A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt Pre-inj Cavernosal A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt Pre-inj Cavernosal A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt Pre-inj Cavernosal A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt Pre-inj Cavernosal A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Pre-inj Dorsal A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Pre-inj Dorsal A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Pre-inj Dorsal A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Pre-inj Dorsal A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Pre-inj Dorsal A | End Diastoic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Pre-inj Dorsal A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Pre-inj Dorsal A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Pre-inj Dorsal A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Pre-inj Dorsal A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Pre-inj Dorsal A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Pre-inj Dorsal A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Pre-inj Dorsal A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Pre-inj Dorsal A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Pre-inj Dorsal A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Pre-inj Dorsal A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt Dist DFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist DFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist DFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist DFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist DFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist DFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist DFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist DFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist DFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist DFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist DFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist DFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist DFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist DFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist DFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid DFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid DFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid DFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid DFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid DFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid DFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid DFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid DFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid DFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid DFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid DFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid DFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid DFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid DFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid DFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt DFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt DFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt DFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Prox DFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox DFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox DFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox DFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox DFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox DFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox DFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox DFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox DFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox DFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox DFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox DFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox DFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox DFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox DFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist DFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist DFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist DFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist DFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist DFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist DFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist DFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist DFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist DFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist DFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist DFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist DFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist DFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist DFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist DFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid DFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid DFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid DFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid DFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid DFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid DFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid DFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid DFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid DFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid DFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid DFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid DFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid DFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid DFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid DFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt DFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt DFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt DFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Prox DFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox DFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox DFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox DFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox DFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox DFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox DFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox DFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox DFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox DFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox DFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox DFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox DFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox DFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox DFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist DFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Dist DFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Dist DFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Dist DFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Dist DFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Dist DFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Dist DFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist DFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Dist DFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Dist DFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Dist DFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Dist DFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Dist DFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Dist DFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Dist DFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid DFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid DFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid DFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid DFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid DFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid DFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid DFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid DFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid DFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid DFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid DFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid DFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid DFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid DFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid DFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox DFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox DFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox DFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox DFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox DFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox DFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox DFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox DFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox DFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox DFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox DFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox DFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox DFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox DFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox DFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist DFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist DFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist DFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist DFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist DFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist DFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist DFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist DFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist DFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist DFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist DFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist DFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist DFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist DFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist DFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid DFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid DFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid DFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid DFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid DFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid DFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid DFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid DFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid DFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid DFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid DFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid DFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid DFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid DFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid DFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox DFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox DFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox DFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox DFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox DFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox DFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox DFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox DFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox DFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox DFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox DFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox DFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox DFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox DFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox DFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Prostate | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Prostate Height;LN;15301-5 | NUM | Prostate;SRT;T-92000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000
Prostate | Length | Length | cm mm | 5 | 2 | Prostate Length;LN;15303-1 | NUM | Prostate;SRT;T-92000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000
Prostate | Width | Width | cm mm | 5 | 2 | Prostate Width;LN;15302-3 | NUM | Prostate;SRT;T-92000 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000
Prostatic Urethra | Prostatic Urethra Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Prostatic Urethra;SRT;T-75110 |  | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000
Lt Dist Radial A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Radial A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Radial A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Radial A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Radial A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Radial A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Radial A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Radial A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Radial A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Radial A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Radial A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Radial A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Radial V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Radial V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Radial V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Radial V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Radial V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Radial V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Radial V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Radial V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Radial V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Radial V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Radial V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Radial V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Radial V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Radial V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Radial V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Radial V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Radial V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Radial Vein;SRT;T-49340 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Radial V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Inf Arcuate A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Inf Arcuate A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Sup Arcuate A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Sup Arcuate A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Main RA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Main RA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Main RA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Main RA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Main RA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Main RA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Main RA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Main RA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Main RA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Main RA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Main RA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Main RA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Main RA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Main RA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Main RA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Main RA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Main RA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Main RA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Main RA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Main RA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Main RA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Main RA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Main RA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Main RA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Main RA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Main RA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Main RA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Main RA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Main RA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Main RA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Main RA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Main RA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Main RA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Main RA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Main RA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Main RA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Main RA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Main RA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Main RA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Main RA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Main RA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Main RA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Main RA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Main RA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Main RA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Main RA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Main RA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Main RA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Main RA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Main RA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Main RA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Main RA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Main RA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Main RA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Main RA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Main RA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Main RA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Main RA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Main RA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Main RA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Main RA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Main RA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Main RA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Main RA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Main RA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Main RA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Main RA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Main RA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Main RA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Main RA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Main RA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Main RA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Main RA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Main RA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Main RA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Main RA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Main RA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Main RA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Main RA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Main RA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Main RA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Main RA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Main RA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Main RA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Main RA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Main RA | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Main RA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Main RA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Main RA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Main RA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Main RA Origin | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Main RA Origin | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Main RA Origin | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Main RA Origin | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Main RA Origin | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Main RA Origin | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Main RA Origin | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Main RA Origin | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Main RA Origin | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Main RA Origin | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Main RA Origin | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Main RA Origin | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Main RA Origin | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Main RA Origin | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Main RA Origin | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt Main RA Origin | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt Main RA Origin | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt Main RA Origin | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt Main RA Origin | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt Main RA Origin | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt Main RA Origin | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt Main RA Origin | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt Main RA Origin | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt Main RA Origin | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt Main RA Origin | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt Main RA Origin | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Main RA Origin | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt Main RA Origin | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt Main RA Origin | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt Main RA Origin | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Main RA Origin (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Main RA Origin (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Main RA Origin (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Main RA Origin (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Main RA Origin (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Main RA Origin (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Main RA Origin (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Main RA Origin (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Main RA Origin (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Main RA Origin (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Main RA Origin (Post Repair) | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Main RA Origin (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Main RA Origin (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Main RA Origin (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Main RA Origin (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin (Post Repair) | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Main RA Origin (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt RA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt RA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt RA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt RA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt RA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt RA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt RA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt RA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt RA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt RA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt RA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt RA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt RA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt RA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt RA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt RA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt RA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt RA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt RA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt RA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt RA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt RA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt RA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt RA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt RA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt RA Stent | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt RA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt RA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt RA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt RA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt Kidney Cortical Thickness | Lt Kidney Cortical Thickness |  | cm mm | 5 | 2 | Renal Cortical Thickness;99SIEMENS;RenalCorticalThk | NUM | Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Kidney Cortical Thickness | Rt Kidney Cortical Thickness |  | cm mm | 5 | 2 | Renal Cortical Thickness;99SIEMENS;RenalCorticalThk | NUM | Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Kidney;SRT;T-71000 | Laterality;SRT;G-C171=Right;SRT;G-A100
PostVoid Bladder | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Post voiding;DCM;109135 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
PostVoid Bladder | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Post voiding;DCM;109135 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
PostVoid Bladder | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Post voiding;DCM;109135 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
PreVoid Bladder | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Prior to voiding;DCM;109134 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
PreVoid Bladder | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Prior to voiding;DCM;109134 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
PreVoid Bladder | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Prior to voiding;DCM;109134 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
Lt Inf Segmental A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Inf Segmental A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Sup Segmental A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Sup Segmental A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Renal V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Renal V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Renal V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Renal V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Renal V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Renal V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Renal V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Renal V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Renal V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Renal V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Renal V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Renal V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Renal V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Renal V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Renal V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Renal V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Renal V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Renal V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Renal V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Renal V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Renal V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Renal V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Renal V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Renal V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Renal V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Renal V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Renal V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Renal V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Renal V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Renal Vein;SRT;T-48740 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Renal V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt SFJ | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Saphenofemoral Junction;SRT;T-D930A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt SFJ | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Saphenofemoral Junction;SRT;T-D930A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt SFJ | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Saphenofemoral Junction;SRT;T-D930A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt SFJ | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt SFJ | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Saphenofemoral Junction;SRT;T-D930A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt SFJ | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Saphenofemoral Junction;SRT;T-D930A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt SFJ | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Saphenofemoral Junction;SRT;T-D930A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt SFJ | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt SPJ | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Saphenopopliteal Junction;SRT;T-4941A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt SPJ | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Saphenopopliteal Junction;SRT;T-4941A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt SPJ | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Saphenopopliteal Junction;SRT;T-4941A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt SPJ | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt SPJ | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Saphenopopliteal Junction;SRT;T-4941A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt SPJ | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Saphenopopliteal Junction;SRT;T-4941A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt SPJ | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Saphenopopliteal Junction;SRT;T-4941A |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt SPJ | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Scrotal Wall | Left Scrotal Wall Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Scrotal Wall;SRT;T-98009 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Scrotal Wall | Right Scrotal Wall Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Scrotal Wall;SRT;T-98009 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Mid Segmental A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Inf Segmental A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Mid Segmental A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Segmental A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Sup Segmental A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Seminal Ves | Lt Seminal Vesicle Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Seminal Vesicle;SRT;T-93000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000
Rt Seminal Ves | Rt Seminal Vesicle Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Seminal Vesicle;SRT;T-93000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000
Lt SSV: Ankle | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Ankle;SRT;T-15750
Lt SSV: Ankle | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Ankle;SRT;T-15750
Lt SSV: Ankle | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Ankle;SRT;T-15750
Lt SSV: Ankle | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt SSV: Ankle | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Ankle;SRT;T-15750
Rt SSV: Ankle | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Ankle;SRT;T-15750
Rt SSV: Ankle | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Ankle;SRT;T-15750
Rt SSV: Ankle | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt SSV: Mid Calf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Lt SSV: Mid Calf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Lt SSV: Mid Calf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Lt SSV: Mid Calf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt SSV: Mid Calf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Rt SSV: Mid Calf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Rt SSV: Mid Calf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Middle;SRT;G-A109 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Rt SSV: Mid Calf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt SSV: Lower Calf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Lt SSV: Lower Calf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Lt SSV: Lower Calf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Lt SSV: Lower Calf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt SSV: Lower Calf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Rt SSV: Lower Calf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Rt SSV: Lower Calf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Lower;SRT;G-A115 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Rt SSV: Lower Calf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt SSV: Upper Calf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Lt SSV: Upper Calf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Lt SSV: Upper Calf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Lt SSV: Upper Calf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt SSV: Upper Calf | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Rt SSV: Upper Calf | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Rt SSV: Upper Calf | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Lesser Saphenous Vein;SRT;T-49550 |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Upper;SRT;G-A116 |  | Finding Site;SRT;G-C0E3=Calf of leg;SRT;T-D9440
Rt SSV: Upper Calf | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Soleus V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Soleal Vein;SRT;G-036B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Soleus V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Soleal Vein;SRT;G-036B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Soleus V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Soleal Vein;SRT;G-036B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Soleus V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Soleus V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Soleal Vein;SRT;G-036B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Soleus V | Reflux Time | Reflux Time | ms | 5 | 2 | Reflux Time;99SIEMENS;RefluxTime | NUM | Soleal Vein;SRT;G-036B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Soleus V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Soleal Vein;SRT;G-036B |  | Finding Site;SRT;G-C0E3=Vein of Lower Extremity;SRT;T-49403 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Soleus V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Spleen | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Spleen;SRT;T-C3000 |  | Finding Site;SRT;G-C0E3=Spleen;SRT;T-C3000
Spleen | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Spleen;SRT;T-C3000 |  | Finding Site;SRT;G-C0E3=Spleen;SRT;T-C3000
Spleen | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Spleen;SRT;T-C3000 |  | Finding Site;SRT;G-C0E3=Spleen;SRT;T-C3000
Dist Splenic A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Splenic A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Splenic A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Splenic A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Splenic A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Splenic A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Splenic A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Splenic A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Splenic A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Splenic A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Splenic A | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Splenic A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Splenic A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Splenic A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Splenic A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Splenic A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Splenic A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Splenic A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Splenic A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Splenic A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Splenic A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Splenic A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Splenic A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Splenic A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Splenic A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Splenic A | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid Splenic A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Splenic A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Splenic A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Splenic A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Splenic A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Splenic A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Splenic A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Splenic A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Splenic A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Splenic A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Splenic A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Splenic A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Splenic A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Splenic A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Splenic A | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Splenic A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Splenic A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Splenic A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Splenic A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Splenic A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Splenic A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Splenic A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Splenic A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Splenic A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Splenic A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Splenic A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Splenic A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Splenic A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Splenic A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Splenic A | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Splenic A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Splenic A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Splenic A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Splenic A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Dist Splenic V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Splenic Vein;SRT;T-48890 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Splenic V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Splenic Vein;SRT;T-48890 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Splenic V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Splenic Vein;SRT;T-48890 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Splenic V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Splenic Vein;SRT;T-48890 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Dist Splenic V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Splenic Vein;SRT;T-48890 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Splenic V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid Splenic V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Splenic Vein;SRT;T-48890 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Splenic V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Splenic V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Splenic Vein;SRT;T-48890 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Splenic V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Splenic V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Splenic Vein;SRT;T-48890 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Splenic V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Splenic V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Splenic Vein;SRT;T-48890 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Splenic V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid Splenic V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Splenic Vein;SRT;T-48890 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Splenic V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Splenic V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Splenic Vein;SRT;T-48890 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Splenic V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Splenic V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Splenic Vein;SRT;T-48890 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Splenic V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Landing Site | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Landing Site | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Landing Site | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Landing Site | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Landing Site | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Landing Site | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Landing Site | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Landing Site | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Landing Site | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Landing Site | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Landing Site | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Landing Site | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Landing Site | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Landing Site | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Landing Site | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Prox Landing Site | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Landing Site | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Landing Site | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Landing Site | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Landing Site | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Landing Site | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Landing Site | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Landing Site | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Landing Site | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Landing Site | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Landing Site | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Landing Site | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Landing Site | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Landing Site | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Landing Site | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Landing Site | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Prox Landing Site | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Dist Landing Site | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Prox Landing Site | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Lt Dist Subclav A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Subclav A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Subclav A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Subclav A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Subclav A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Subclav A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Subclav A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Subclav A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Subclav A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Subclav A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Subclav A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Subclav A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Subclav A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Subclav A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Subclav A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Subclavian Artery;SRT;T-46101 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45006 | Left;SRT;G-A102
Lt Subclav A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Subclav A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Subclav A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Subclav A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Subclav A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Subclav A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Dist Subclav A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Right;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Subclav A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Subclav A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Right;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Subclav A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Subclav A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Subclav A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Subclav A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Subclav A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Subclav A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Subclav A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Subclav A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Subclav A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Subclav A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Subclav A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Subclav A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Subclav A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Subclav A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Subclav A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Subclav A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Subclav A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Subclav A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist Subclav A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Subclav A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Subclav A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Subclav A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Subclav A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Subclav A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Subclav A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Subclav A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Subclav A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Subclav A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Subclav A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Subclav A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Subclav V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Vein;SRT;T-48330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Subclavian Vein;SRT;T-48330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Subclav V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Vein;SRT;T-48330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Subclavian Vein;SRT;T-48330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Subclav V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Vein;SRT;T-48330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Subclavian Vein;SRT;T-48330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Subclav V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Vein;SRT;T-48330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Subclavian Vein;SRT;T-48330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Subclav V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Subclav V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Vein;SRT;T-48330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Subclavian Vein;SRT;T-48330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Subclav V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Subclav V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Subclavian Vein;SRT;T-48330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Subclavian Vein;SRT;T-48330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Subclav V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist SFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist SFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist SFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist SFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist SFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist SFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist SFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist SFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist SFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist SFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist SFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist SFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist SFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist SFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist SFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid SFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid SFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid SFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid SFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid SFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid SFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid SFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid SFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid SFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid SFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid SFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid SFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid SFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid SFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid SFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox SFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox SFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox SFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox SFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox SFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox SFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox SFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox SFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox SFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox SFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox SFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox SFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox SFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox SFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox SFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist SFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist SFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist SFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist SFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist SFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist SFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist SFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist SFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist SFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist SFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist SFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist SFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist SFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist SFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist SFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid SFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid SFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid SFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid SFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid SFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid SFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid SFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid SFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid SFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid SFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid SFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid SFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid SFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid SFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid SFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox SFA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox SFA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox SFA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox SFA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox SFA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox SFA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox SFA | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox SFA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox SFA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox SFA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox SFA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox SFA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox SFA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox SFA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox SFA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist SFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Dist SFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Dist SFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Dist SFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Dist SFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Dist SFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Dist SFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist SFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Dist SFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Dist SFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Dist SFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Dist SFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Dist SFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Dist SFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Dist SFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid SFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid SFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid SFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid SFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid SFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid SFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid SFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid SFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid SFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid SFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid SFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid SFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid SFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid SFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid SFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox SFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox SFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox SFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox SFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox SFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox SFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox SFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox SFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox SFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox SFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox SFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox SFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox SFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox SFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox SFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist SFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist SFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist SFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist SFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist SFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist SFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist SFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist SFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist SFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist SFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist SFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist SFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist SFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist SFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist SFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid SFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid SFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid SFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid SFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid SFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid SFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid SFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid SFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid SFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid SFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid SFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid SFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid SFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid SFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid SFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox SFA Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox SFA Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox SFA Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox SFA Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox SFA Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox SFA Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox SFA Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox SFA Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox SFA Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox SFA Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox SFA Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox SFA Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox SFA Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox SFA Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox SFA Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Dist SMA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist SMA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist SMA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist SMA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist SMA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist SMA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist SMA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist SMA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist SMA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist SMA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist SMA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist SMA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist SMA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist SMA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist SMA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid SMA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SMA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SMA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SMA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SMA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SMA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SMA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SMA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SMA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SMA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SMA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid SMA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SMA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SMA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SMA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox SMA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SMA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SMA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SMA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SMA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SMA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SMA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SMA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SMA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SMA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SMA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox SMA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SMA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SMA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SMA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
SMA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
SMA | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA Origin | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
SMA Origin | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
SMA Origin (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin, Postprandial | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
SMA Origin, Postprandial | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
SMA Origin, Postprandial | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
SMA Origin, Postprandial | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
SMA Origin, Postprandial | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
SMA Origin, Postprandial | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
SMA Origin, Postprandial | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
SMA Origin, Postprandial | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
SMA Origin, Postprandial | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
SMA Origin, Postprandial | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
SMA Origin, Postprandial | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
SMA Origin, Postprandial | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
SMA Origin, Postprandial | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
SMA Origin, Postprandial | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
SMA Origin, Postprandial | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist SMA, Postprandial | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist SMA, Postprandial | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist SMA, Postprandial | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist SMA, Postprandial | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist SMA, Postprandial | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist SMA, Postprandial | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist SMA, Postprandial | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist SMA, Postprandial | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist SMA, Postprandial | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist SMA, Postprandial | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist SMA, Postprandial | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist SMA, Postprandial | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist SMA, Postprandial | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist SMA, Postprandial | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist SMA, Postprandial | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Mid SMA, Postprandial | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Mid SMA, Postprandial | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Mid SMA, Postprandial | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Mid SMA, Postprandial | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Mid SMA, Postprandial | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Mid SMA, Postprandial | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Mid SMA, Postprandial | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Mid SMA, Postprandial | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Mid SMA, Postprandial | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Mid SMA, Postprandial | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Mid SMA, Postprandial | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid SMA, Postprandial | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Mid SMA, Postprandial | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Mid SMA, Postprandial | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Mid SMA, Postprandial | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Prox SMA, Postprandial | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Prox SMA, Postprandial | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Prox SMA, Postprandial | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Prox SMA, Postprandial | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Prox SMA, Postprandial | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Prox SMA, Postprandial | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Prox SMA, Postprandial | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Prox SMA, Postprandial | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Prox SMA, Postprandial | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Prox SMA, Postprandial | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Prox SMA, Postprandial | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox SMA, Postprandial | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Prox SMA, Postprandial | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Prox SMA, Postprandial | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Prox SMA, Postprandial | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
SMV | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superior Mesenteric Vein;SRT;T-48840 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMV | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Superior Mesenteric Vein;SRT;T-48840 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMV | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
SMV | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Superior Mesenteric Vein;SRT;T-48840 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMV | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Sup Thyroid A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Sup Thyroid A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Sup Thyroid A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Sup Thyroid A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Sup Thyroid A | End Diastoic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Sup Thyroid A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Sup Thyroid A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Sup Thyroid A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Sup Thyroid A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Sup Thyroid A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Sup Thyroid A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Sup Thyroid A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Sup Thyroid A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Sup Thyroid A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Sup Thyroid A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Sup Thyroid A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Sup Thyroid A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Sup Thyroid A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Sup Thyroid A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Sup Thyroid A | End Diastoic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Sup Thyroid A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Sup Thyroid A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Sup Thyroid A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Sup Thyroid A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Sup Thyroid A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Sup Thyroid A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Sup Thyroid A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Sup Thyroid A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Sup Thyroid A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Sup Thyroid A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Dist SVC | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superior Vena Cava;SRT;T-48610 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid SVC | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superior Vena Cava;SRT;T-48610 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox SVC | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Superior Vena Cava;SRT;T-48610 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist SVC | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Superior Vena Cava;SRT;T-48610 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist SVC | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid SVC | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Superior Vena Cava;SRT;T-48610 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SVC | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox SVC | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Superior Vena Cava;SRT;T-48610 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SVC | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist SVC | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Superior Vena Cava;SRT;T-48610 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist SVC | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid SVC | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Superior Vena Cava;SRT;T-48610 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SVC | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox SVC | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Superior Vena Cava;SRT;T-48610 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SVC | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Suprarenal Ao | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Suprarenal Ao | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Suprarenal Ao | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Suprarenal Ao | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Suprarenal Ao | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Suprarenal Ao | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Suprarenal Ao | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Suprarenal Ao | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Suprarenal Ao | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Suprarenal Ao | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Suprarenal Ao | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Suprarenal Ao | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Suprarenal Ao | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Suprarenal Ao | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Suprarenal Ao | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Suprarenal Ao (Post Repair) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao (Post Repair) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao (Post Repair) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao (Post Repair) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao (Post Repair) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao (Post Repair) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao (Post Repair) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao (Post Repair) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao (Post Repair) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao (Post Repair) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao (Post Repair) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Suprarenal Ao (Post Repair) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao (Post Repair) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao (Post Repair) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao (Post Repair) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao (Post Repair) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao (Post Repair) | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Suprarenal Ao | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  | Measurement Orientation;DCM;125105=Anterior-Posterior;DCM;122675
Suprarenal Ao | Transverse Diameter | Trans Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  | Measurement Orientation;DCM;125105=Transverse;SRT;G-A117
Lt Terminal ICA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal ICA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal ICA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal ICA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal ICA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal ICA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal ICA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal ICA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal ICA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal ICA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Terminal ICA | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal ICA | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal ICA | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal ICA | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal ICA | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal ICA | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal ICA | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal ICA | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal ICA | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal ICA | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Terminal Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal Vert A | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Terminal Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Terminal Vert A | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Testicle | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Testis;SRT;T-94000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Testicle | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Testis;SRT;T-94000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Testicle | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Testis;SRT;T-94000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Testicle | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Testis;SRT;T-94000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Testicle | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Testis;SRT;T-94000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Testicle | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Testis;SRT;T-94000 |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Testicular A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Testicular A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Testicular A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Testicular A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Testicular A | End Diastoic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Testicular A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Testicular A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Testicular A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Testicular A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Testicular A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Testicular A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Testicular A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Testicular A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Testicular A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Testicular A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Testicular A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Testicular A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Testicular A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Testicular A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Testicular A | End Diastoic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Testicular A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Testicular A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Testicular A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Testicular A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Testicular A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Testicular A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Testicular A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Testicular A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Testicular A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Testicular A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Testicular V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Testicular Vein;SRT;T-48770 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Testicular V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Testicular Vein;SRT;T-48770 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Testicular V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Testicular Vein;SRT;T-48770 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Testicular V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Testicular V | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Testicular Vein;SRT;T-48770 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Testicular V | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Testicular V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Testicular Vein;SRT;T-48770 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Testicular V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Testicular V | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Testicular Vein;SRT;T-48770 |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Testicular V | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Thyroid | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Lobe of Thyroid Gland;SRT;T-B6070 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Thyroid | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Lobe of Thyroid Gland;SRT;T-B6070 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Thyroid | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Lobe of Thyroid Gland;SRT;T-B6070 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Thyroid | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Lobe of Thyroid Gland;SRT;T-B6070 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Thyroid | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Lobe of Thyroid Gland;SRT;T-B6070 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Thyroid | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Lobe of Thyroid Gland;SRT;T-B6070 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Lt Dist T-P Trunk | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist T-P Trunk | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist T-P Trunk | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist T-P Trunk | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist T-P Trunk | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist T-P Trunk | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist T-P Trunk | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist T-P Trunk | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist T-P Trunk | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist T-P Trunk | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist T-P Trunk | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist T-P Trunk | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist T-P Trunk | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist T-P Trunk | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist T-P Trunk | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid T-P Trunk | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid T-P Trunk | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid T-P Trunk | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid T-P Trunk | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid T-P Trunk | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid T-P Trunk | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid T-P Trunk | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid T-P Trunk | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid T-P Trunk | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid T-P Trunk | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid T-P Trunk | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid T-P Trunk | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid T-P Trunk | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid T-P Trunk | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid T-P Trunk | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox T-P Trunk | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox T-P Trunk | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox T-P Trunk | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox T-P Trunk | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox T-P Trunk | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox T-P Trunk | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox T-P Trunk | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox T-P Trunk | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox T-P Trunk | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox T-P Trunk | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox T-P Trunk | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox T-P Trunk | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox T-P Trunk | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox T-P Trunk | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox T-P Trunk | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist T-P Trunk | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist T-P Trunk | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist T-P Trunk | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist T-P Trunk | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist T-P Trunk | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist T-P Trunk | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist T-P Trunk | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist T-P Trunk | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist T-P Trunk | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist T-P Trunk | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist T-P Trunk | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist T-P Trunk | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist T-P Trunk | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist T-P Trunk | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist T-P Trunk | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid T-P Trunk | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid T-P Trunk | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid T-P Trunk | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid T-P Trunk | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid T-P Trunk | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid T-P Trunk | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid T-P Trunk | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid T-P Trunk | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid T-P Trunk | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid T-P Trunk | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid T-P Trunk | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid T-P Trunk | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid T-P Trunk | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid T-P Trunk | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid T-P Trunk | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox T-P Trunk | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox T-P Trunk | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox T-P Trunk | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox T-P Trunk | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox T-P Trunk | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox T-P Trunk | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox T-P Trunk | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox T-P Trunk | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox T-P Trunk | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox T-P Trunk | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox T-P Trunk | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox T-P Trunk | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox T-P Trunk | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox T-P Trunk | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox T-P Trunk | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Stent-Mid | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Transjugular Intrahepatic Portosystemic Shunt;SRT;G-036C |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Stent-Hepatic End | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Transjugular Intrahepatic Portosystemic Shunt;SRT;G-036C |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Stent-Hepatic End | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Transjugular Intrahepatic Portosystemic Shunt;SRT;G-036C |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Stent-Hepatic End | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Stent-Portal End | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Transjugular Intrahepatic Portosystemic Shunt;SRT;G-036C |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Stent-Portal End | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Transjugular Intrahepatic Portosystemic Shunt;SRT;G-036C |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Stent-Portal End | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Stent-Mid | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Transjugular Intrahepatic Portosystemic Shunt;SRT;G-036C |  | Finding Site;SRT;G-C0E3=Vein of Abdomen;SRT;T-487A0 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Stent-Mid | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Donor A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Donor A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Donor A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Donor A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Donor A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Donor A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Donor A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Donor A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Donor A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Donor A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Donor A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Dist Donor A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Donor A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Donor A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Donor A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Donor A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Donor A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Donor A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Donor A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Donor A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Donor A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Donor A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Donor A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Donor A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Donor A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Donor A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Mid Donor A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Donor A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Donor A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Donor A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Donor A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Donor A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Donor A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Donor A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Donor A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Donor A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Donor A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Donor A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Donor A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Donor A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Donor A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Prox Donor A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Donor A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Donor A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Donor A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Portal V Anast | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Liver Transplant Anastomosis Site Portal Vein from graft;LN;44775-5 |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10
Portal V Anast | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Liver Transplant Anastomosis Site Portal Vein from graft;LN;44775-5 |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10
Portal V Anast | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Portal V Anast | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Liver Transplant Anastomosis Site Portal Vein from graft;LN;44775-5 |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10
Portal V Anast | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Portal V Pre-anast | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Portal Vein;SRT;T-48810 |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10
Portal V Pre-anast | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Portal Vein;SRT;T-48810 |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10
Portal V Pre-anast | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Portal V Pre-anast | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Portal Vein;SRT;T-48810 |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10
Portal V Pre-anast | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Arcuate A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Arcuate A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Arcuate A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Arcuate A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Inf Arcuate A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Inf Arcuate A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Arcuate A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Inf Arcuate A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Sup Arcuate A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Sup Arcuate A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Arcuate A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Sup Arcuate A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Art Anast | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Art Anast | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Art Anast | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Art Anast | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Art Anast | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Art Anast | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Art Anast | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Art Anast | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Art Anast | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Art Anast | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Art Anast | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Art Anast | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Art Anast | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Art Anast | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Art Anast | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Art Anast | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Dist RA (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist RA (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist RA (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist RA (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist RA (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist RA (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist RA (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist RA (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist RA (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist RA (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist RA (Tx) | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist RA (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist RA (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist RA (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist RA (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid RA (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid RA (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid RA (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid RA (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid RA (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid RA (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid RA (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid RA (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid RA (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid RA (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid RA (Tx) | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid RA (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid RA (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid RA (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid RA (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox RA (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox RA (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox RA (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox RA (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox RA (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox RA (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox RA (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox RA (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox RA (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox RA (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox RA (Tx) | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox RA (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox RA (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox RA (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox RA (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist RA (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist RA (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist RA (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist RA (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist RA (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist RA (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist RA (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist RA (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist RA (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist RA (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist RA (Tx) | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist RA (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist RA (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist RA (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist RA (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid RA (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid RA (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid RA (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid RA (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid RA (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid RA (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid RA (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid RA (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid RA (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid RA (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid RA (Tx) | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid RA (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid RA (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid RA (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid RA (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox RA (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox RA (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox RA (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox RA (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox RA (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox RA (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox RA (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox RA (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox RA (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox RA (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox RA (Tx) | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox RA (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox RA (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox RA (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox RA (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Donor A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Donor A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Donor A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Donor A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Donor A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Donor A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Donor A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Donor A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Donor A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Donor A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Donor A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Donor A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Donor A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Donor A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Donor A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Donor A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Donor A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Donor A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Donor A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Donor A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Donor A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Donor A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Donor A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Donor A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Donor A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Donor A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Donor A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Donor A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Donor A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Donor A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Donor A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Donor A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Donor A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Donor A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Donor A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Donor A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Donor A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Donor A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Donor A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Donor A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Donor A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Donor A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Donor A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Donor A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Donor A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Donor A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Donor A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Donor A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Donor A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Donor A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Donor A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Donor A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Donor A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Donor A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Donor A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Donor A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Donor A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Donor A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Donor A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Donor A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Donor A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Donor A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Donor A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Donor A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Donor A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Donor A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Donor A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Donor A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Donor A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Donor A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Donor A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Donor A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Donor A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Donor A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Donor A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Donor A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Donor A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Donor A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Donor A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Donor A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Donor A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Donor A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Donor A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Donor A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Donor A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Donor A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Donor A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Donor A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Donor A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Donor A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid Interlobar A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Interlobar A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Interlobar A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Interlobar A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Inf Interlobar A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Inf Interlobar A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Interlobar A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Inf Interlobar A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Sup Interlobar A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Sup Interlobar A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Interlobar A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Sup Interlobar A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Mid Segmental A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Segmental A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Mid Segmental A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Segmental A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Inf Segmental A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Inf Segmental A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Inf Segmental A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Inf Segmental A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Sup Segmental A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Sup Segmental A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Sup Segmental A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A (Tx) | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A (Tx) | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A (Tx) | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A (Tx) | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A (Tx) | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A (Tx) | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A (Tx) | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A (Tx) | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A (Tx) | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A (Tx) | Doppler Angle Property of PS, ED, TAMx, TAMn | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Sup Segmental A (Tx) | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A (Tx) | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A (Tx) | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A (Tx) | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Renal V (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Vein;SRT;T-48740 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Renal V (Tx) | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Renal Vein;SRT;T-48740 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Renal V Anast | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Vein;SRT;T-48930 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Renal V Anast | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Iliac Vein;SRT;T-48930 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Renal V Anast | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | External Iliac Vein;SRT;T-48930 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Renal V Anast | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Renal V Anast | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | External Iliac Vein;SRT;T-48930 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Renal V Anast | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Renal V Anast | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | External Iliac Vein;SRT;T-48930 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Renal V Anast | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Renal V Anast | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | External Iliac Vein;SRT;T-48930 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Renal V Anast | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Renal V (Tx) | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Renal Vein;SRT;T-48740 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Renal V (Tx) | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Renal V (Tx) | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Renal Vein;SRT;T-48740 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Renal V (Tx) | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Renal V (Tx) | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Renal Vein;SRT;T-48740 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Renal V (Tx) | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Renal V (Tx) | Minimum Velocity | Vmin | cm/s m/s | 5 | 2 | Blood velocity, minimum;DCM;122206 | NUM | Renal Vein;SRT;T-48740 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Renal V (Tx) | Doppler Angle Property of Vmin | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Kidney (Tx) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Transplanted Kidney;SRT;T-7100A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Kidney (Tx) | Anterior-Posterior Diameter | AP | cm mm | 5 | 2 | Height;DCM;121207 | NUM | Transplanted Kidney;SRT;T-7100A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Kidney (Tx) | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Transplanted Kidney;SRT;T-7100A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Kidney (Tx) | Length | Length | cm mm | 5 | 2 | Length;SRT;G-A22A | NUM | Transplanted Kidney;SRT;T-7100A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Kidney (Tx) | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Transplanted Kidney;SRT;T-7100A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Kidney (Tx) | Width | Width | cm mm | 5 | 2 | Width;SRT;G-A220 | NUM | Transplanted Kidney;SRT;T-7100A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Dist Ulnar A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Ulnar A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Ulnar A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Ulnar A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Ulnar A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Ulnar A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Ulnar A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Ulnar A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Ulnar A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Ulnar A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Ulnar A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar A | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Ulnar A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Ulnar V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Ulnar V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Ulnar V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Ulnar V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Ulnar V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Ulnar V 1 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar V 1 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar V 1 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Ulnar V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Ulnar V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Ulnar V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Ulnar V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ulnar V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Ulnar V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Ulnar V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Ulnar V 2 | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar V 2 | Maximum Velocity | Vmax | cm/s m/s | 5 | 2 | Blood velocity, peak;DCM;122207 | NUM | Ulnar Vein;SRT;T-49330 |  | Finding Site;SRT;G-C0E3=Vein of Upper Extremity;SRT;T-49103 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Ulnar V 2 | Doppler Angle Property of Vmax | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt AVF Art Anast | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Art Anast | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Art Anast | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Art Anast | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Art Anast | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Art Anast | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Art Anast | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Art Anast | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Art Anast | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt AVF Art Anast | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Art Anast | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Art Anast | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Art Anast | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Art Anast | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Art Anast | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt AVF Art Anast | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Art Anast | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Art Anast | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Art Anast | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Art Anast | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Art Anast | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Art Anast | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Art Anast | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Art Anast | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt AVF Art Anast | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Art Anast | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Art Anast | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Art Anast | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Art Anast | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Art Anast | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist AVF | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist AVF | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid AVF | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid AVF | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox AVF | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox AVF | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist AVF | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist AVF | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid AVF | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid AVF | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox AVF | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox AVF | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist AVF Inflow A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Inflow A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Inflow A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Inflow A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Inflow A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Inflow A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Inflow A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Inflow A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Inflow A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist AVF Inflow A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Inflow A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Inflow A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Inflow A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Inflow A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Inflow A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid AVF Inflow A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Inflow A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Inflow A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Inflow A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Inflow A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Inflow A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Inflow A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Inflow A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Inflow A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid AVF Inflow A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Inflow A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Inflow A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Inflow A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Inflow A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Inflow A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox AVF Inflow A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Inflow A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Inflow A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Inflow A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Inflow A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Inflow A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Inflow A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Inflow A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Inflow A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox AVF Inflow A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Inflow A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Inflow A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Inflow A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Inflow A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Inflow A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist AVF Inflow A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Inflow A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Inflow A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Inflow A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Inflow A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Inflow A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Inflow A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Inflow A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Inflow A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist AVF Inflow A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Inflow A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Inflow A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Inflow A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Inflow A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Inflow A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid AVF Inflow A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Inflow A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Inflow A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Inflow A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Inflow A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Inflow A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Inflow A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Inflow A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Inflow A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid AVF Inflow A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Inflow A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Inflow A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Inflow A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Inflow A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Inflow A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox AVF Inflow A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Inflow A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Inflow A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Inflow A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Inflow A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Inflow A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Inflow A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Inflow A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Inflow A | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox AVF Inflow A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Inflow A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Inflow A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Inflow A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Inflow A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Inflow A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist AVF Outflow V | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Outflow V | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Outflow V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Outflow V | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Outflow V | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Outflow V | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Outflow V | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Outflow V | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Outflow V | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist AVF Outflow V | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Outflow V | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Outflow V | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Outflow V | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Outflow V | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist AVF Outflow V | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid AVF Outflow V | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Outflow V | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Outflow V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Outflow V | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Outflow V | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Outflow V | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Outflow V | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Outflow V | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Outflow V | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid AVF Outflow V | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Outflow V | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Outflow V | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Outflow V | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Outflow V | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid AVF Outflow V | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox AVF Outflow V | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Outflow V | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Outflow V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Outflow V | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Outflow V | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Outflow V | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Outflow V | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Outflow V | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Outflow V | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox AVF Outflow V | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Outflow V | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Outflow V | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Outflow V | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Outflow V | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox AVF Outflow V | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist AVF Outflow V | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Outflow V | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Outflow V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Outflow V | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Outflow V | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Outflow V | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Outflow V | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Outflow V | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Outflow V | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist AVF Outflow V | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Outflow V | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Outflow V | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Outflow V | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Outflow V | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist AVF Outflow V | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid AVF Outflow V | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Outflow V | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Outflow V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Outflow V | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Outflow V | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Outflow V | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Outflow V | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Outflow V | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Outflow V | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid AVF Outflow V | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Outflow V | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Outflow V | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Outflow V | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Outflow V | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid AVF Outflow V | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox AVF Outflow V | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Outflow V | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Outflow V | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Outflow V | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Outflow V | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Outflow V | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Outflow V | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Outflow V | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Outflow V | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox AVF Outflow V | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Outflow V | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Outflow V | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Outflow V | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Outflow V | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox AVF Outflow V | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt AVF Venous Anast | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Venous Anast | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Venous Anast | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Venous Anast | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Venous Anast | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Venous Anast | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Venous Anast | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Venous Anast | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Venous Anast | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt AVF Venous Anast | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Venous Anast | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Venous Anast | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Venous Anast | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Venous Anast | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Venous Anast | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt AVF Venous Anast | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Venous Anast | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Venous Anast | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Venous Anast | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Venous Anast | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Venous Anast | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Venous Anast | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Venous Anast | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Venous Anast | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt AVF Venous Anast | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Venous Anast | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Venous Anast | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Venous Anast | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Venous Anast | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt AVF Venous Anast | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt UE Dist Anast | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Dist Anast | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Dist Anast | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Dist Anast | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Dist Anast | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Dist Anast | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Dist Anast | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Dist Anast | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Dist Anast | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt UE Dist Anast | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Dist Anast | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Dist Anast | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Dist Anast | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Dist Anast | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Dist Anast | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Prox Anast | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Prox Anast | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Prox Anast | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Prox Anast | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Prox Anast | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Prox Anast | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Prox Anast | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Prox Anast | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Prox Anast | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt UE Prox Anast | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Prox Anast | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Prox Anast | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Prox Anast | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Prox Anast | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Prox Anast | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Dist Anast | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Dist Anast | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Dist Anast | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Dist Anast | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Dist Anast | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Dist Anast | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Dist Anast | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Dist Anast | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Dist Anast | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt UE Dist Anast | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Dist Anast | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Dist Anast | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Dist Anast | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Dist Anast | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Dist Anast | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Prox Anast | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Prox Anast | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Prox Anast | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Prox Anast | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Prox Anast | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Prox Anast | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Prox Anast | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Prox Anast | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Prox Anast | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt UE Prox Anast | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Prox Anast | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Prox Anast | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Prox Anast | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Prox Anast | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Prox Anast | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Dist Graft | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Dist Graft | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Dist Graft | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Dist Graft | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Dist Graft | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Dist Graft | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Dist Graft | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Dist Graft | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Dist Graft | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt UE Dist Graft | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Dist Graft | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Dist Graft | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Dist Graft | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Dist Graft | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Dist Graft | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Mid Graft | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Mid Graft | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Mid Graft | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Mid Graft | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Mid Graft | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Mid Graft | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Mid Graft | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Mid Graft | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Mid Graft | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt UE Mid Graft | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Mid Graft | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Mid Graft | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Mid Graft | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Mid Graft | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Mid Graft | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Prox Graft | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Prox Graft | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Prox Graft | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Prox Graft | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Prox Graft | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Prox Graft | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Prox Graft | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Prox Graft | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Prox Graft | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt UE Prox Graft | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Prox Graft | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Prox Graft | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Prox Graft | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Prox Graft | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Prox Graft | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Dist Graft | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Dist Graft | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Dist Graft | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Dist Graft | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Dist Graft | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Dist Graft | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Dist Graft | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Dist Graft | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Dist Graft | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt UE Dist Graft | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Dist Graft | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Dist Graft | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Dist Graft | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Dist Graft | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Dist Graft | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Mid Graft | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Mid Graft | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Mid Graft | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Mid Graft | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Mid Graft | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Mid Graft | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Mid Graft | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Mid Graft | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Mid Graft | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt UE Mid Graft | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Mid Graft | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Mid Graft | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Mid Graft | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Mid Graft | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Mid Graft | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Prox Graft | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Prox Graft | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Prox Graft | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Prox Graft | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Prox Graft | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Prox Graft | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Prox Graft | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Prox Graft | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Prox Graft | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt UE Prox Graft | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Prox Graft | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Prox Graft | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Prox Graft | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Prox Graft | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Prox Graft | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Dist Inflow Vessel | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Inflow Vessel | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Inflow Vessel | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Inflow Vessel | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Inflow Vessel | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Inflow Vessel | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Inflow Vessel | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Inflow Vessel | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Inflow Vessel | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt UE Dist Inflow Vessel | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Inflow Vessel | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Inflow Vessel | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Inflow Vessel | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Inflow Vessel | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Inflow Vessel | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Mid Inflow Vessel | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Inflow Vessel | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Inflow Vessel | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Inflow Vessel | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Inflow Vessel | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Inflow Vessel | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Inflow Vessel | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Inflow Vessel | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Inflow Vessel | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt UE Mid Inflow Vessel | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Inflow Vessel | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Inflow Vessel | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Inflow Vessel | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Inflow Vessel | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Inflow Vessel | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Prox Inflow Vessel | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Inflow Vessel | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Inflow Vessel | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Inflow Vessel | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Inflow Vessel | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Inflow Vessel | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Inflow Vessel | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Inflow Vessel | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Inflow Vessel | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt UE Prox Inflow Vessel | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Inflow Vessel | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Inflow Vessel | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Inflow Vessel | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Inflow Vessel | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Inflow Vessel | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Dist Inflow Vessel | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Inflow Vessel | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Inflow Vessel | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Inflow Vessel | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Inflow Vessel | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Inflow Vessel | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Inflow Vessel | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Inflow Vessel | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Inflow Vessel | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt UE Dist Inflow Vessel | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Inflow Vessel | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Inflow Vessel | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Inflow Vessel | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Inflow Vessel | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Inflow Vessel | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Mid Inflow Vessel | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Inflow Vessel | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Inflow Vessel | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Inflow Vessel | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Inflow Vessel | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Inflow Vessel | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Inflow Vessel | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Inflow Vessel | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Inflow Vessel | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt UE Mid Inflow Vessel | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Inflow Vessel | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Inflow Vessel | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Inflow Vessel | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Inflow Vessel | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Inflow Vessel | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Prox Inflow Vessel | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Inflow Vessel | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Inflow Vessel | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Inflow Vessel | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Inflow Vessel | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Inflow Vessel | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Inflow Vessel | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Inflow Vessel | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Inflow Vessel | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt UE Prox Inflow Vessel | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Inflow Vessel | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Inflow Vessel | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Inflow Vessel | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Inflow Vessel | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Inflow Vessel | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Dist Outflow Vessel | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Outflow Vessel | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Outflow Vessel | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Outflow Vessel | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Outflow Vessel | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Outflow Vessel | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Outflow Vessel | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Outflow Vessel | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Outflow Vessel | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt UE Dist Outflow Vessel | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Outflow Vessel | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Outflow Vessel | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Outflow Vessel | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Outflow Vessel | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Dist Outflow Vessel | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Mid Outflow Vessel | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Outflow Vessel | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Outflow Vessel | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Outflow Vessel | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Outflow Vessel | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Outflow Vessel | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Outflow Vessel | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Outflow Vessel | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Outflow Vessel | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt UE Mid Outflow Vessel | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Outflow Vessel | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Outflow Vessel | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Outflow Vessel | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Outflow Vessel | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Mid Outflow Vessel | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Prox Outflow Vessel | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Outflow Vessel | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Outflow Vessel | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Outflow Vessel | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Outflow Vessel | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Outflow Vessel | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Outflow Vessel | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Outflow Vessel | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Outflow Vessel | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt UE Prox Outflow Vessel | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Outflow Vessel | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Outflow Vessel | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Outflow Vessel | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Outflow Vessel | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Prox Outflow Vessel | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Dist Outflow Vessel | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Outflow Vessel | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Outflow Vessel | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Outflow Vessel | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Outflow Vessel | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Outflow Vessel | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Outflow Vessel | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Outflow Vessel | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Outflow Vessel | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt UE Dist Outflow Vessel | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Outflow Vessel | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Outflow Vessel | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Outflow Vessel | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Outflow Vessel | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Dist Outflow Vessel | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Mid Outflow Vessel | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Outflow Vessel | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Outflow Vessel | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Outflow Vessel | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Outflow Vessel | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Outflow Vessel | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Outflow Vessel | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Outflow Vessel | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Outflow Vessel | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt UE Mid Outflow Vessel | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Outflow Vessel | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Outflow Vessel | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Outflow Vessel | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Outflow Vessel | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Mid Outflow Vessel | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Prox Outflow Vessel | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Outflow Vessel | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Outflow Vessel | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Outflow Vessel | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Outflow Vessel | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Outflow Vessel | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Outflow Vessel | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Outflow Vessel | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Outflow Vessel | Doppler Angle Property of PS, ED, TAMn, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt UE Prox Outflow Vessel | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Outflow Vessel | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Outflow Vessel | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Outflow Vessel | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Outflow Vessel | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Prox Outflow Vessel | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist UE Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist UE Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist UE Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist UE Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist UE Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist UE Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist UE Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist UE Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist UE Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist UE Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist UE Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist UE Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist UE Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist UE Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist UE Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid UE Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid UE Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid UE Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid UE Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid UE Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid UE Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid UE Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid UE Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid UE Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid UE Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid UE Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid UE Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid UE Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid UE Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid UE Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox UE Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox UE Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox UE Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox UE Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox UE Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox UE Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox UE Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox UE Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox UE Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox UE Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox UE Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox UE Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox UE Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox UE Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox UE Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist UE Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist UE Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist UE Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist UE Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist UE Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist UE Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist UE Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist UE Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist UE Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist UE Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist UE Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist UE Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist UE Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist UE Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist UE Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid UE Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid UE Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid UE Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid UE Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid UE Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid UE Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid UE Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid UE Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid UE Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid UE Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid UE Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid UE Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid UE Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid UE Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid UE Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox UE Stent | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox UE Stent | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox UE Stent | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox UE Stent | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox UE Stent | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox UE Stent | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox UE Stent | Doppler Angle Property of PS, ED | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox UE Stent | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox UE Stent | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox UE Stent | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox UE Stent | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox UE Stent | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox UE Stent | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox UE Stent | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox UE Stent | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Ureter | Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Ureter;SRT;T-73000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
Rt Ureter | Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Ureter;SRT;T-73000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
Urethra | Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Urethra;SRT;T-75000 |  | Finding Site;SRT;G-C0E3=Bladder;SRT;T-74000
Lt Vas Deferens | Vas Deferens Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Vas Deferens;SRT;T-96000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000
Rt Vas Deferens | Vas Deferens Distance |  | cm mm | 5 | 2 | Distance;DCM;121206 | NUM | Vas Deferens;SRT;T-96000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000
Lt Dist Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Vert A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Vert A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Vert A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Vert A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Dist Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Volume Flow | Vol Flow | ml/min | 5 | 2 | Volume Flow;LN;33878-0 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Volume Flow Diameter | D | cm mm | 5 | 2 | Diameter of Volume Flow;99SIEMENS;VolFDiam | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Volume Flow Area | A | cm² mm² | 5 | 2 | Area of Volume Flow;99SIEMENS;VolFArea | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Volume Flow Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity of Volume Flow;99SIEMENS;VolFTAMn | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Diameter | Diam | cm mm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Liver Mass 1 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Elasticity | E | kPa | 20 | 1 | Elasticity;99SIEMENS;Elasticity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 1 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 1 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 1 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 1 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 1 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 2 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 2 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 2 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 2 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 2 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 2 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 3 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 3 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 3 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 3 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 3 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 3 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 4 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 4 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 4 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 4 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 4 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 4 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 5 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 5 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 5 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 5 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 5 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 5 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 6 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 6 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 6 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 6 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 6 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 6 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 7 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 7 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 7 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 7 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 7 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 7 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 8 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 8 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 8 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 8 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 8 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 8 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 9 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 9 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 9 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 9 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 9 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 9 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 10 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 10 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 10 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 10 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 10 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 10 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 1 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 1 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 1 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 1 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 1 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 1 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 2 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 2 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 2 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 2 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 2 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 2 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 3 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 3 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 3 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 3 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 3 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 3 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 4 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 4 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 4 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 4 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 4 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 4 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 5 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 5 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 5 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 5 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 5 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 5 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 6 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 6 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 6 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 6 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 6 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 6 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 7 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 7 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 7 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 7 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 7 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 7 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 8 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 8 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 8 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 8 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 8 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 8 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 9 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 9 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 9 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 9 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 9 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 9 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 10 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 10 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 10 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 10 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 10 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 10 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 1 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 1 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 1 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 1 | Shear Wave Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 1 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 1 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 2 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 2 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 2 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 2 | Shear Wave Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 2 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 2 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 3 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 3 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 3 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 3 | Shear Wave Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 3 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 3 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 4a | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 4a | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 4a | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 4a | Shear Wave Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 4a | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 4a | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 4b | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 4b | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 4b | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 4b | Shear Wave Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 4b | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 4b | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 5 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 5 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 5 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 5 | Shear Wave Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 5 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 5 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 6 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 6 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 6 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 6 | Shear Wave Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 6 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 6 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 7 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 7 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 7 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 7 | Shear Wave Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 7 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 7 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 8 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 8 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 8 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 8 | Shear Wave Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 8 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 8 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 1 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 1 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 1 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 1 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 1 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 1 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 2 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 2 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 2 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 2 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 2 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 2 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 3 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 3 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 3 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 3 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 3 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 3 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 4 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 4 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 4 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 4 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 4 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 4 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 5 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 5 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 5 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 5 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 5 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 5 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 1 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 1 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 1 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 1 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 1 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 1 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 2 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 2 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 2 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 2 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 2 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 2 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 3 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 3 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 3 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 3 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 3 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 3 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 4 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 4 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 4 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 4 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 4 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 4 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 5 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 5 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 5 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 5 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 5 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 5 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 1 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 1 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 1 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 1 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 1 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 1 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 2 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 2 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 2 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 2 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 2 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 2 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 3 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 3 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 3 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 3 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 3 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 3 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 4 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 4 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 4 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 4 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 4 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 4 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 5 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 5 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 5 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 5 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 5 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 5 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 1 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 1 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 1 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 1 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 1 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 1 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 2 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 2 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 2 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 2 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 2 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 2 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 3 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 3 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 3 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 3 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 3 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 3 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 4 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 4 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 4 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 4 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 4 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 4 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 5 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 5 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 5 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 5 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 5 | ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 5 | ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 1 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 1 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 1 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 1 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 1 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 1 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 2 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 2 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 2 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 2 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 2 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 2 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 3 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 3 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 3 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 3 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 3 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 3 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 4 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 4 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 4 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 4 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 4 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 4 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 5 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 5 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 5 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 5 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 5 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 5 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 1 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 1 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 1 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 1 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 1 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 1 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 2 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 2 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 2 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 2 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 2 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 2 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 3 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 3 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 3 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 3 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 3 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 3 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 4 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 4 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 4 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 4 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 4 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 4 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 5 | Elasticity Median | Median | kPa | 20 | 1 | Elasticity Median;99SIEMENS;ElasticityMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 5 | Elasticity Mean | Mean | kPa | 20 | 1 | Elasticity Median: Mean;99SIEMENS;EMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 5 | Elasticity Maximum | Max | kPa | 20 | 1 | Elasticity Median: Maximum;99SIEMENS;EMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 5 | Shear Velocity Median | Median | m/s | 20 | 2 | Shear Wave Velocity Median;99SIEMENS;VsMedian | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 5 | Elasticity ROI Depth | Depth | cm | 20 | 2 | Elasticity ROI Depth;99SIEMENS;ElastROIDepth | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 5 | Elasticity ROI Diameter | Diam | cm | 20 | 2 | Elasticity ROI Diameter;99SIEMENS;ElastROIDiam | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Prox CCA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Mid CCA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Dist CCA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Prox CCA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid CCA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Dist CCA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Bulb | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Bulb | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt ECA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Prox ECA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Mid ECA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Dist ECA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt ECA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Prox ECA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid ECA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Dist ECA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Prox ICA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Mid ICA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Dist ICA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Prox ICA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid ICA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Dist ICA | Diameter | Diam | cm | 5 | 2 | Diameter;SRT;M-02550 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
AAA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt ARA Origin | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Dist ARA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid ARA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox ARA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist ARA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid ARA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox ARA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt ARA Origin | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ACA A1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt ACA A1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt ACA A2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt ACA A2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
ACoA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Anterior Communicating Artery;SRT;T-45530 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Lt Dist ATA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid ATA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox ATA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist ATA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid ATA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox ATA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Aorta | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Aorta | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Aorta | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Aorta (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Ao Stent Graft | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Ao Stent Graft | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Ao Stent Graft | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid Arcuate A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Sup Arcuate A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Vert Atlas Loop | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Vert Atlas Loop | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist Axillary A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Axillary A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Axillary A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Axillary A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Axillary A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Axillary A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Basilar A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Basilar A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Basilar A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Brach A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Brach A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Brach A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Brach A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Brach A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Brach A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Bulb | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Bulb | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Celiac Origin | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin, Postprandial | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Lt Dist MCA M1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid MCA M1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox MCA M1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist MCA M1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid MCA M1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox MCA M1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist MCA M2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid MCA M2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox MCA M2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist MCA M2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid MCA M2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox MCA M2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist CCA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid CCA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox CCA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist CCA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid CCA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox CCA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt CFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist CFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid CFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox CFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt CFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Dist CFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid CFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox CFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist CFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Dist CFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid CFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox CFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
CHA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Dist CHA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid CHA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox CHA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt CIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist CIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid CIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox CIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt CIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Dist CIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid CIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox CIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt CIA (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist Deep Brach A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Deep Brach A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Deep Brach A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Deep Brach A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Deep Brach A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Deep Brach A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt DPA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt DPA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Epididymal A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Epididymal A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist ECA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt ECA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Mid ECA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox ECA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist ECA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt ECA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Mid ECA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox ECA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist EIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt EIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Mid EIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox EIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist EIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt EIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Mid EIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox EIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist EIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Dist Anast 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Prox Anast 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Dist Graft 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Mid Graft 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Prox Graft 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Dist Inflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Inflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Inflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Outflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Outflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Outflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Anast 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Prox Anast 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Dist Graft 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Mid Graft 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Prox Graft 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Dist Inflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Inflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Inflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Outflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Outflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Outflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist LHA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid LHA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox LHA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist RHA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid RHA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox RHA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist IMA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid IMA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox IMA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
IMA Origin | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
IMA Origin (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Inf Thyroid A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Thyroid A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Dist Innom A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Innom A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Mid Innom A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Innom A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid Interlobar A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Inf Interlobar A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Sup Interlobar A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Dist ICA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid ICA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox ICA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist ICA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid ICA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox ICA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt ICA Siphon | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt ICA Siphon | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist ICA-skull entrance | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA-skull entrance | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt IIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Mid IIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox IIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist IIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt IIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Mid IIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox IIA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist IIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt LE Dist Anast 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Prox Anast 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Dist Anast 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Prox Anast 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Dist Graft 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Mid Graft 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Prox Graft 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Dist Graft 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Mid Graft 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Prox Graft 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Dist Inflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Mid Inflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Prox Inflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Dist Inflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Mid Inflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Prox Inflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Dist Outflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Mid Outflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Prox Outflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Dist Outflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Mid Outflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Prox Outflow Vessel 1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Dist Anast 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Prox Anast 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Dist Anast 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Prox Anast 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Dist Graft 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Mid Graft 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Prox Graft 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Dist Graft 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Mid Graft 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Prox Graft 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Dist Inflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Mid Inflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Prox Inflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Dist Inflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Mid Inflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Prox Inflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Dist Outflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Mid Outflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Prox Outflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Dist Outflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Mid Outflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Prox Outflow Vessel 2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Ophthalmic A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Ophthalmic A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Ophthalmic A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Ophthalmic A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Ophthalmic A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Ophthalmic A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Peroneal A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Peroneal A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Peroneal A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Peroneal A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Peroneal A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Peroneal A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Pop A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Pop A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Pop A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Pop A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Pop A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Pop A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Pop A Aneurysm | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop A Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid Pop A Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox Pop A Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist Pop A Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid Pop A Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox Pop A Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Post-inj Cavernosal A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt Post-inj Cavernosal A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Post-inj Dorsal A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt PCA P1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt PCA P1 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt PCA P2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt PCA P2 | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt PCoA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt PCoA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist PTA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid PTA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox PTA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist PTA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid PTA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox PTA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Pre-inj Cavernosal A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt Pre-inj Cavernosal A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Pre-inj Dorsal A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt Dist DFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid DFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt DFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Prox DFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist DFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid DFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt DFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Prox DFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist DFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid DFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox DFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist DFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid DFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox DFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Dist Radial A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Radial A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Radial A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Radial A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Radial A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Radial A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Inf Arcuate A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Dist Main RA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Main RA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Main RA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Main RA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Main RA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Main RA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Main RA Origin | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt Main RA Origin | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Main RA Origin (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt RA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt RA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt Inf Segmental A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Mid Segmental A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Sup Segmental A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Dist Splenic A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Splenic A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Splenic A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Splenic A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Dist Landing Site | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Prox Landing Site | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Subclav A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Subclav A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Subclav A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Subclav A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Dist Subclav A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Subclav A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Subclav A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Subclav A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist Subclav A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Subclav A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Subclav A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Subclav A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Subclav A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Subclav A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist SFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid SFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox SFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist SFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid SFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox SFA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist SFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid SFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox SFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist SFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid SFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox SFA Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Dist SMA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid SMA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox SMA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
SMA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA Origin | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin, Postprandial | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist SMA, Postprandial | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Mid SMA, Postprandial | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Prox SMA, Postprandial | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Lt Sup Thyroid A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Sup Thyroid A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Suprarenal Ao | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Suprarenal Ao (Post Repair) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Terminal ICA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Terminal ICA | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Terminal Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Terminal Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Testicular A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Testicular A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist T-P Trunk | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid T-P Trunk | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox T-P Trunk | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist T-P Trunk | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid T-P Trunk | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox T-P Trunk | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Donor A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Donor A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Donor A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid Arcuate A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Inf Arcuate A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Sup Arcuate A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Art Anast | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Dist RA (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid RA (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox RA (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist RA (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid RA (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox RA (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Donor A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Donor A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Donor A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Donor A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Donor A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Donor A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid Interlobar A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Inf Interlobar A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Sup Interlobar A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Mid Segmental A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Inf Segmental A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Sup Segmental A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A (Tx) | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Dist Ulnar A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Ulnar A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Ulnar A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Ulnar A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Ulnar A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Ulnar A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt AVF Art Anast | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt AVF Art Anast | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist AVF | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid AVF | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox AVF | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist AVF | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid AVF | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox AVF | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist AVF Inflow A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid AVF Inflow A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox AVF Inflow A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist AVF Inflow A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid AVF Inflow A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox AVF Inflow A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist AVF Outflow V | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid AVF Outflow V | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox AVF Outflow V | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist AVF Outflow V | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid AVF Outflow V | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox AVF Outflow V | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt AVF Venous Anast | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt AVF Venous Anast | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt UE Dist Anast | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Prox Anast | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Dist Anast | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Prox Anast | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Dist Graft | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Mid Graft | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Prox Graft | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Dist Graft | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Mid Graft | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Prox Graft | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Dist Inflow Vessel | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Mid Inflow Vessel | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Prox Inflow Vessel | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Dist Inflow Vessel | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Mid Inflow Vessel | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Prox Inflow Vessel | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Dist Outflow Vessel | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Mid Outflow Vessel | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Prox Outflow Vessel | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Dist Outflow Vessel | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Mid Outflow Vessel | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Prox Outflow Vessel | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist UE Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid UE Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox UE Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist UE Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid UE Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox UE Stent | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Dist Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
AAA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
AAA (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt ARA Origin | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Dist ARA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid ARA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox ARA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist ARA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid ARA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox ARA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt ARA Origin | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ACA A1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt ACA A1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt ACA A2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt ACA A2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
ACoA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Anterior Communicating Artery;SRT;T-45530 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Lt Dist ATA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid ATA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox ATA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist ATA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid ATA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox ATA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Aorta | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Aorta | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Aorta | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Aorta (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Ao Stent Graft | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Ao Stent Graft | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Ao Stent Graft | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid Arcuate A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Sup Arcuate A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Vert Atlas Loop | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Vert Atlas Loop | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist Axillary A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Axillary A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Axillary A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Axillary A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Axillary A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Axillary A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Basilar A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Basilar A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Basilar A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Brach A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Brach A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Brach A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Brach A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Brach A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Brach A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Bulb | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Bulb | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Celiac Origin | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin, Postprandial | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Lt Dist MCA M1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid MCA M1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox MCA M1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist MCA M1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid MCA M1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox MCA M1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist MCA M2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid MCA M2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox MCA M2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist MCA M2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid MCA M2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox MCA M2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist CCA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid CCA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox CCA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist CCA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid CCA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox CCA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt CFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist CFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid CFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox CFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt CFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Dist CFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid CFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox CFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist CFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Dist CFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid CFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox CFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
CHA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Dist CHA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid CHA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox CHA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt CIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist CIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid CIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox CIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt CIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Dist CIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid CIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox CIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt CIA (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Mid CIA (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Prox CIA (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA Aneurysm | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA Aneurysm (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Dist CIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid CIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox CIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist CIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Mid CIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Prox CIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Dist Deep Brach A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Deep Brach A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Deep Brach A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Deep Brach A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Deep Brach A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Deep Brach A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt DPA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt DPA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Epididymal A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Epididymal A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist ECA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt ECA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Mid ECA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox ECA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist ECA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt ECA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Mid ECA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox ECA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist EIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt EIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Mid EIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox EIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist EIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt EIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Mid EIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox EIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist EIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Mid EIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Lt Prox EIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Dist EIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Mid EIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Rt Prox EIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Dist Anast 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Prox Anast 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Dist Graft 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Mid Graft 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Prox Graft 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Dist Inflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Inflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Inflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Outflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Outflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Outflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Anast 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Prox Anast 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Dist Graft 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Mid Graft 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Prox Graft 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Dist Inflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Inflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Inflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Outflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Outflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Outflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 2] |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist LHA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid LHA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox LHA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist RHA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid RHA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox RHA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist IMA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid IMA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox IMA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
IMA Origin | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
IMA Origin (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Inf Thyroid A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Inf Thyroid A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Inferior Thyroid Artery;SRT;T-46140 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Dist Innom A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Innom A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Mid Innom A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Innom A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid Interlobar A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Inf Interlobar A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Sup Interlobar A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Dist ICA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid ICA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox ICA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist ICA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid ICA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox ICA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt ICA Siphon | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt ICA Siphon | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist ICA-skull entrance | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA-skull entrance | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt IIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Mid IIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox IIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist IIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt IIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Mid IIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox IIA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist IIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid IIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox IIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist IIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Mid IIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Prox IIA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt LE Dist Anast 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Prox Anast 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Dist Anast 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Prox Anast 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Dist Graft 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Mid Graft 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Prox Graft 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Dist Graft 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Mid Graft 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Prox Graft 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Dist Inflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Mid Inflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Prox Inflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Dist Inflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Mid Inflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Prox Inflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Dist Outflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Mid Outflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Prox Outflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Dist Outflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Mid Outflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Prox Outflow Vessel 1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Dist Anast 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Prox Anast 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt LE Dist Anast 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt LE Prox Anast 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Dist Graft 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Mid Graft 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Prox Graft 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt LE Dist Graft 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt LE Mid Graft 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt LE Prox Graft 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Dist Inflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Mid Inflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Prox Inflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Dist Inflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Mid Inflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Prox Inflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Dist Outflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Mid Outflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Prox Outflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt LE Dist Outflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt LE Mid Outflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt LE Prox Outflow Vessel 2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Ophthalmic A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Ophthalmic A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Ophthalmic A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Ophthalmic A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Ophthalmic A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Ophthalmic A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Peroneal A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Peroneal A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Peroneal A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Peroneal A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Peroneal A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Peroneal A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Pop A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Pop A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Pop A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Pop A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Pop A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Pop A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Pop A Aneurysm | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Pop A Aneurysm | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist Pop A Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid Pop A Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox Pop A Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist Pop A Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Mid Pop A Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Prox Pop A Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Post-inj Cavernosal A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Rt Post-inj Cavernosal A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Post-inj Dorsal A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt PCA P1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt PCA P1 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt PCA P2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt PCA P2 | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt PCoA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt PCoA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist PTA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid PTA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox PTA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist PTA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid PTA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox PTA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Pre-inj Cavernosal A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Rt Pre-inj Cavernosal A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Pre-inj Dorsal A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt Dist DFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid DFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt DFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Prox DFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist DFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid DFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt DFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Prox DFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist DFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid DFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox DFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist DFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Mid DFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Prox DFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Dist Radial A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Radial A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Radial A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Radial A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Radial A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Radial A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Inf Arcuate A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Dist Main RA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Main RA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Main RA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Main RA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Main RA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Main RA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Main RA Origin | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt Main RA Origin | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt Main RA Origin (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Main RA Origin (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt RA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Rt RA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt Inf Segmental A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Mid Segmental A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Sup Segmental A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Dist Splenic A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Splenic A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Splenic A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Splenic A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Dist Landing Site | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Prox Landing Site | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Subclav A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Subclav A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Subclav A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Subclav A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Dist Subclav A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Subclav A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Subclav A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Subclav A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist Subclav A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Subclav A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Subclav A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Subclav A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Subclav A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Subclav A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist SFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid SFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox SFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist SFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid SFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox SFA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist SFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid SFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox SFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Dist SFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Mid SFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Rt Prox SFA Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Dist SMA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid SMA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox SMA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
SMA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
SMA Origin | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
SMA Origin (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
SMA Origin, Postprandial | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist SMA, Postprandial | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Mid SMA, Postprandial | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Prox SMA, Postprandial | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Lt Sup Thyroid A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Rt Sup Thyroid A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000
Suprarenal Ao | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Suprarenal Ao (Post Repair) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Supra-renal Aorta;SRT;T-42510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Terminal ICA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Terminal ICA | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Terminal Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Terminal Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Testicular A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Testicular A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist T-P Trunk | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid T-P Trunk | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox T-P Trunk | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist T-P Trunk | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid T-P Trunk | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox T-P Trunk | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Dist Donor A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Mid Donor A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Prox Donor A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Donor Artery;99SIEMENS;DonorArtery |  | Finding Site;SRT;G-C0E3=Transplantation of liver;SRT;P1-5BD10 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid Arcuate A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Arcuate A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Inf Arcuate A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Arcuate A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Sup Arcuate A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Arcuate A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Art Anast | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Rt Art Anast | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | External Iliac Artery;SRT;T-46910 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00
Lt Dist RA (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid RA (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox RA (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist RA (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid RA (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox RA (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Renal Artery;SRT;T-46600 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Donor A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Donor A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Donor A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Donor A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Donor A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Donor A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Donor Artery;99SIEMENS;DonorArtery | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid Interlobar A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Interlobar A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Inf Interlobar A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Interlobar A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Sup Interlobar A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Interlobar A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Interlobar Artery of Kidney;SRT;T-4667D | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Mid Segmental A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Mid Segmental A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Inf Segmental A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Inf Segmental A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Sup Segmental A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Sup Segmental A (Tx) | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Segmental Artery;SRT;T-46659 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Transplant of kidney;SRT;P1-71D00 |  | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Dist Ulnar A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Ulnar A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Ulnar A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist Ulnar A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Ulnar A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Ulnar A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt AVF Art Anast | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt AVF Art Anast | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Dist AVF | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid AVF | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox AVF | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist AVF | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid AVF | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox AVF | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist AVF Inflow A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid AVF Inflow A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox AVF Inflow A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist AVF Inflow A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid AVF Inflow A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox AVF Inflow A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist AVF Outflow V | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid AVF Outflow V | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox AVF Outflow V | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist AVF Outflow V | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid AVF Outflow V | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox AVF Outflow V | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt AVF Venous Anast | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt AVF Venous Anast | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt UE Dist Anast | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Prox Anast | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Rt UE Dist Anast | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Rt UE Prox Anast | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Dist Graft | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Mid Graft | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Prox Graft | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Rt UE Dist Graft | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Rt UE Mid Graft | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Rt UE Prox Graft | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Dist Inflow Vessel | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Mid Inflow Vessel | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Prox Inflow Vessel | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Dist Inflow Vessel | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Mid Inflow Vessel | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Prox Inflow Vessel | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Dist Outflow Vessel | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Mid Outflow Vessel | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Prox Outflow Vessel | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt UE Dist Outflow Vessel | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt UE Mid Outflow Vessel | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt UE Prox Outflow Vessel | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist UE Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid UE Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox UE Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Dist UE Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid UE Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox UE Stent | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Dist Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Mid Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Prox Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Liver Mass 1 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Shear Wave Velocity | Vs | m/s | 20 | 2 | Shear Wave Velocity;99SIEMENS;ShearVelocity | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Rt ACA A1 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt ACA A1 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt ACA A2 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt ACA A2 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
ACoA | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Dist ICA-skull entrance | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Dist ICA-skull entrance | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Terminal ICA | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Terminal ICA | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt ICA Siphon | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt ICA Siphon | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Prox MCA M1 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Mid MCA M1 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Dist MCA M1 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Prox MCA M1 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Mid MCA M1 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Dist MCA M1 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Prox MCA M2 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Mid MCA M2 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Dist MCA M2 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Prox MCA M2 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Mid MCA M2 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Dist MCA M2 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Prox Ophthalmic A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Mid Ophthalmic A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Dist Ophthalmic A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Prox Ophthalmic A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Mid Ophthalmic A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Dist Ophthalmic A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt PCA P1 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt PCA P1 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt PCA P2 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt PCA P2 | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt PCoA | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt PCoA | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Vert Atlas Loop | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Vert Atlas Loop | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Terminal Vert A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Terminal Vert A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Prox Basilar A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Mid Basilar A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Dist Basilar A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt ACA A1 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt ACA A1 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt ACA A2 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt ACA A2 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
ACoA | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Anterior Communicating Artery;SRT;T-45530 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Rt Dist ICA-skull entrance | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA-skull entrance | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Terminal ICA | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Terminal ICA | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt ICA Siphon | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt ICA Siphon | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Prox MCA M1 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Mid MCA M1 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Dist MCA M1 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Prox MCA M1 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid MCA M1 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Dist MCA M1 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Prox MCA M2 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Mid MCA M2 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Dist MCA M2 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Prox MCA M2 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid MCA M2 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Dist MCA M2 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Prox Ophthalmic A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Mid Ophthalmic A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Dist Ophthalmic A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Prox Ophthalmic A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid Ophthalmic A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Dist Ophthalmic A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt PCA P1 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt PCA P1 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt PCA P2 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt PCA P2 | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt PCoA | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt PCoA | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Vert Atlas Loop | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Vert Atlas Loop | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Terminal Vert A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Terminal Vert A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Prox Basilar A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Mid Basilar A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Dist Basilar A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Prox Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Mid Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Dist Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Prox Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Mid Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Dist Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Prox Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Dist Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Prox Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Dist Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Prox Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox Vert A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Mid Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid Vert A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Dist Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Vert A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Prox Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Mid Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Dist Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Prox Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Mid Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Dist Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Prox Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Dist Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Prox Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Mid Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Dist Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Prox Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Dist Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Prox Vert A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Mid Vert A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Dist Vert A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Prox Vert A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Mid Vert A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Dist Vert A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Prox Vert A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Mid Vert A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Dist Vert A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Prox Vert A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid Vert A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Dist Vert A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Liver Mass 1 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 1 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 2 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 2 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 3 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 3 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 4 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 4 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 5 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 5 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 6 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 6 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 7 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 7 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 8 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 8 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 9 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 9 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 10 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Mass 10 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 1 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 1 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 2 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 2 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 3 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 3 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 4 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 4 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 5 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 5 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 6 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 6 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 7 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 7 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 8 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 8 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 9 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 9 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 10 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Site 10 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 1 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 1 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 2 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 2 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 3 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 3 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 4a | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 4a | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 4b | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 4b | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 5 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 5 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 6 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 6 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 7 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 7 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 8 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Liver Segment 8 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 1 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 1 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 2 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 2 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 3 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 3 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 4 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 4 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 5 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Mass 5 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 1 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 1 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 2 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 2 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 3 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 3 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 4 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 4 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 5 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Thyroid Site 5 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 1 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 1 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 2 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 2 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 3 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 3 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 4 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 4 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 5 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Mass 5 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 1 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 1 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 2 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 2 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 3 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 3 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 4 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 4 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 5 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Rt Thyroid Site 5 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Site;SRT;G-A165 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 1 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 1 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 2 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 2 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 3 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 3 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 4 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 4 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 5 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Lt Mass 5 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 1 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 1 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 2 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 2 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 3 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 3 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 4 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 4 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 5 | Shear Wave Velocity Mean | Mean | m/s | 20 | 2 | Shear Wave Velocity Median: Mean;99SIEMENS;VsMdMean | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Prostate Rt Mass 5 | Shear Wave Velocity Maximum | Max | m/s | 20 | 2 | Shear Wave Velocity Median: Maximum;99SIEMENS;VsMdMax | NUM | Mass;SRT;M-03000 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Finding Site;SRT;G-C0E3=Prostate;SRT;T-92000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Shear Wave Elastography;99SIEMENS;SWE
Lt Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Lt Vert A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Lt Vert A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Rt Vert A | Acceleration Index | Accel | cm/s² m/s² | 5 | 2 | Acceleration Index;LN;20167-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Acceleration Time | Time | ms | 5 | 0 | Acceleration Time;LN;20168-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Acceleration Velocity 1 | V1 | cm/s m/s | 5 | 2 | Velocity 1 of Acceleration;99SIEMENS;AccelVel1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Acceleration Velocity 2 | V2 | cm/s m/s | 5 | 2 | Velocity 2 of Acceleration;99SIEMENS;AccelVel2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Pulsatility Index | PI | unitless | 5 | 2 | Pulsatility Index;LN;12008-9 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Peak Systolic Velocity | PS | cm/s m/s | 5 | 2 | Peak Systolic Velocity;LN;11726-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | End Diastolic Velocity | ED | cm/s m/s | 5 | 2 | End Diastolic Velocity;LN;11653-3 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Peak Systolic to End Diastolic Velocity Ratio | S/D | unitless | 5 | 2 | Systolic to Diastolic Velocity Ratio;LN;12144-2 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Resistivity Index | RI | unitless | 5 | 2 | Resistivity Index;LN;12023-8 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Time Averaged Maximum Velocity | TAMx | cm/s m/s | 5 | 2 | Time averaged peak velocity;LN;11692-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Doppler Angle Property of PS, ED, TAMx | Angle | ° | 5 | 0 | Doppler Angle;DCM;125106 | NUM | Named Attribute
Rt Vert A | Time Averaged Mean Velocity | TAMn | cm/s m/s | 5 | 2 | Time averaged mean velocity;LN;20352-1 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Vert A | Doppler Sample Volume Depth | SV Depth | mm | 5 | 2 | Sample Volume Depth;DCM;125107 | NUM | Named Attribute
Rt Vert A | Mean TCD | Mean TCD | cm/s m/s | 5 | 2 | Mean (TCD);99SIEMENS;MeanTCD | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Prox CCA IMT | Mean Thickness | Mean | mm cm | 5 | 2 | Mean Intima Media Thickness;99SIEMENS;MeanIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA IMT | Max Thickness | Max | mm cm | 5 | 2 | Max Intima Media Thickness;99SIEMENS;MaxIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA IMT | Min Thickness | Min | mm cm | 5 | 2 | Min Intima Media Thickness;99SIEMENS;MinIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox CCA IMT | C10 Thickness | C10 | mm cm | 5 | 2 | C10 Intima Media Thickness;99SIEMENS;IMTC10 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Mid CCA IMT | Mean Thickness | Mean | mm cm | 5 | 2 | Mean Intima Media Thickness;99SIEMENS;MeanIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA IMT | Max Thickness | Max | mm cm | 5 | 2 | Max Intima Media Thickness;99SIEMENS;MaxIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA IMT | Min Thickness | Min | mm cm | 5 | 2 | Min Intima Media Thickness;99SIEMENS;MinIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid CCA IMT | C10 Thickness | C10 | mm cm | 5 | 2 | C10 Intima Media Thickness;99SIEMENS;IMTC10 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Dist CCA IMT | Mean Thickness | Mean | mm cm | 5 | 2 | Mean Intima Media Thickness;99SIEMENS;MeanIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA IMT | Max Thickness | Max | mm cm | 5 | 2 | Max Intima Media Thickness;99SIEMENS;MaxIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA IMT | Min Thickness | Min | mm cm | 5 | 2 | Min Intima Media Thickness;99SIEMENS;MinIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA IMT | C10 Thickness | C10 | mm cm | 5 | 2 | C10 Intima Media Thickness;99SIEMENS;IMTC10 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Prox CCA IMT | Mean Thickness | Mean | mm cm | 5 | 2 | Mean Intima Media Thickness;99SIEMENS;MeanIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA IMT | Max Thickness | Max | mm cm | 5 | 2 | Max Intima Media Thickness;99SIEMENS;MaxIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA IMT | Min Thickness | Min | mm cm | 5 | 2 | Min Intima Media Thickness;99SIEMENS;MinIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA IMT | C10 Thickness | C10 | mm cm | 5 | 2 | C10 Intima Media Thickness;99SIEMENS;IMTC10 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid CCA IMT | Mean Thickness | Mean | mm cm | 5 | 2 | Mean Intima Media Thickness;99SIEMENS;MeanIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA IMT | Max Thickness | Max | mm cm | 5 | 2 | Max Intima Media Thickness;99SIEMENS;MaxIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA IMT | Min Thickness | Min | mm cm | 5 | 2 | Min Intima Media Thickness;99SIEMENS;MinIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA IMT | C10 Thickness | C10 | mm cm | 5 | 2 | C10 Intima Media Thickness;99SIEMENS;IMTC10 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Dist CCA IMT | Mean Thickness | Mean | mm cm | 5 | 2 | Mean Intima Media Thickness;99SIEMENS;MeanIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA IMT | Max Thickness | Max | mm cm | 5 | 2 | Max Intima Media Thickness;99SIEMENS;MaxIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA IMT | Min Thickness | Min | mm cm | 5 | 2 | Min Intima Media Thickness;99SIEMENS;MinIMT | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist CCA IMT | C10 Thickness | C10 | mm cm | 5 | 2 | C10 Intima Media Thickness;99SIEMENS;IMTC10 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Prox ICA IMT | Mean Thickness | Mean | mm cm | 5 | 2 | Mean Intima Media Thickness;99SIEMENS;MeanIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA IMT | Max Thickness | Max | mm cm | 5 | 2 | Max Intima Media Thickness;99SIEMENS;MaxIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA IMT | Min Thickness | Min | mm cm | 5 | 2 | Min Intima Media Thickness;99SIEMENS;MinIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Prox ICA IMT | C10 Thickness | C10 | mm cm | 5 | 2 | C10 Intima Media Thickness;99SIEMENS;IMTC10 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Rt Mid ICA IMT | Mean Thickness | Mean | mm cm | 5 | 2 | Mean Intima Media Thickness;99SIEMENS;MeanIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA IMT | Max Thickness | Max | mm cm | 5 | 2 | Max Intima Media Thickness;99SIEMENS;MaxIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA IMT | Min Thickness | Min | mm cm | 5 | 2 | Min Intima Media Thickness;99SIEMENS;MinIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Mid ICA IMT | C10 Thickness | C10 | mm cm | 5 | 2 | C10 Intima Media Thickness;99SIEMENS;IMTC10 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Rt Dist ICA IMT | Mean Thickness | Mean | mm cm | 5 | 2 | Mean Intima Media Thickness;99SIEMENS;MeanIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA IMT | Max Thickness | Max | mm cm | 5 | 2 | Max Intima Media Thickness;99SIEMENS;MaxIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA IMT | Min Thickness | Min | mm cm | 5 | 2 | Min Intima Media Thickness;99SIEMENS;MinIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA IMT | C10 Thickness | C10 | mm cm | 5 | 2 | C10 Intima Media Thickness;99SIEMENS;IMTC10 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Prox ICA IMT | Mean Thickness | Mean | mm cm | 5 | 2 | Mean Intima Media Thickness;99SIEMENS;MeanIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA IMT | Max Thickness | Max | mm cm | 5 | 2 | Max Intima Media Thickness;99SIEMENS;MaxIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA IMT | Min Thickness | Min | mm cm | 5 | 2 | Min Intima Media Thickness;99SIEMENS;MinIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA IMT | C10 Thickness | C10 | mm cm | 5 | 2 | C10 Intima Media Thickness;99SIEMENS;IMTC10 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid ICA IMT | Mean Thickness | Mean | mm cm | 5 | 2 | Mean Intima Media Thickness;99SIEMENS;MeanIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA IMT | Max Thickness | Max | mm cm | 5 | 2 | Max Intima Media Thickness;99SIEMENS;MaxIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA IMT | Min Thickness | Min | mm cm | 5 | 2 | Min Intima Media Thickness;99SIEMENS;MinIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA IMT | C10 Thickness | C10 | mm cm | 5 | 2 | C10 Intima Media Thickness;99SIEMENS;IMTC10 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Dist ICA IMT | Mean Thickness | Mean | mm cm | 5 | 2 | Mean Intima Media Thickness;99SIEMENS;MeanIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA IMT | Max Thickness | Max | mm cm | 5 | 2 | Max Intima Media Thickness;99SIEMENS;MaxIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA IMT | Min Thickness | Min | mm cm | 5 | 2 | Min Intima Media Thickness;99SIEMENS;MinIMT | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA IMT | C10 Thickness | C10 | mm cm | 5 | 2 | C10 Intima Media Thickness;99SIEMENS;IMTC10 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Bulb IMT | Mean Thickness | Mean | mm cm | 5 | 2 | Mean Intima Media Thickness;99SIEMENS;MeanIMT | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb IMT | Max Thickness | Max | mm cm | 5 | 2 | Max Intima Media Thickness;99SIEMENS;MaxIMT | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb IMT | Min Thickness | Min | mm cm | 5 | 2 | Min Intima Media Thickness;99SIEMENS;MinIMT | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb IMT | C10 Thickness | C10 | mm cm | 5 | 2 | C10 Intima Media Thickness;99SIEMENS;IMTC10 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Lt Bulb IMT | Mean Thickness | Mean | mm cm | 5 | 2 | Mean Intima Media Thickness;99SIEMENS;MeanIMT | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb IMT | Max Thickness | Max | mm cm | 5 | 2 | Max Intima Media Thickness;99SIEMENS;MaxIMT | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb IMT | Min Thickness | Min | mm cm | 5 | 2 | Min Intima Media Thickness;99SIEMENS;MinIMT | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb IMT | C10 Thickness | C10 | mm cm | 5 | 2 | C10 Intima Media Thickness;99SIEMENS;IMTC10 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Liver Mass 1 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 2 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 3 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 4 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 5 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 6 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 7 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 8 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 9 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 10 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Mass;SRT;M-03000 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 1 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 2 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 3 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 4 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 5 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 6 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 7 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 8 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 9 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=9 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Site 10 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Site;SRT;G-A165 |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=10 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 1 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=1 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 2 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=2 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 3 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=3 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4a | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4a | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 4b | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=4b | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 5 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=5 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 6 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=6 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 7 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=7 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Ultrasound Deried Fat Fraction | UDFF | % | 20 | 0 | Ultrasound Derived Fat Fraction;99SIEMENS;UDFF | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Segment 8 | Elasticity ROI Depth | Depth | cm | 20 | 1 | Ultrasound Derived Fat Fraction ROI Depth;99SIEMENS;UDFFROIDepth | NUM | Liver Segment;SRT;T-D053F |  | Finding Site;SRT;G-C0E3=Liver;SRT;T-62000 |  |  |  |  | Identifier;DCM;125010=8 | Measurement Method;SRT;G-C036=Point Shear Wave Elastography;99SIEMENS;pSWE
Liver Mass 1 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Mass 2 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Mass 3 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Mass 4 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Mass 5 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Mass 6 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Mass 7 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Mass 8 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Mass 9 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Mass 10 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Site 1 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Site 2 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Site 3 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Site 4 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Site 5 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Site 6 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Site 7 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Site 8 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Site 9 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Site 10 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Segment 1 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Segment 2 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Segment 3 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Segment 4a | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Segment 4b | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Segment 5 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Segment 6 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Segment 7 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
Liver Segment 8 | Ultrasound Attenuation Coefficient | Atten | dB/cm/MHz | 20 | 2
AAA (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
AAA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Abdominal Aortic Aneurysm;SRT;D3-83305 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
ACoA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Anterior Communicating Artery;SRT;T-45530 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Ao Bif (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aortic Bifurcation;SRT;T-42580 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Ao Bif | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aortic Bifurcation;SRT;T-42580
Celiac Origin | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Celiac Axis;SRT;T-46400 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Celiac Origin, Postprandial | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Celiac Axis;SRT;T-46400 | [6 empty] | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
CHA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103
Dist Aorta (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aorta;SRT;T-42000 | [6 empty] | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Dist Aorta | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Ao Stent Graft | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft
Dist Basilar A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist CHA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Hepatic Artery;SRT;T-46421
Dist IMA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist Innom A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Innominate Artery;SRT;T-46010
Dist Splenic A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Splenic Artery;SRT;T-46460
Dist Landing Site | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Dist SMA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superior Mesenteric Artery;SRT;T-46510
Dist SMA, Postprandial | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 | [6 empty] | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Dist Anast 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070
Prox Anast 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | [Graft 1] |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Dist Graft 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070
Mid Graft 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070
Prox Graft 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070
Dist Inflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel
Mid Inflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel
Prox Inflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft 1 Inflow Vessel;99SIEMENS;Graft1InVessel
Dist Outflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel
Mid Outflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel
Prox Outflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft 1 Outflow Vessel;99SIEMENS;Graft1OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F
Dist Anast 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F
Prox Anast 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F
Dist Graft 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F
Mid Graft 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F
Prox Graft 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F
Dist Inflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F
Mid Inflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F
Prox Inflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft 2 Inflow Vessel;99SIEMENS;Graft2InVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F
Dist Outflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F
Mid Outflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F
Prox Outflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft 2 Outflow Vessel;99SIEMENS;Graft2OutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F
IMA Origin (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 |  |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
IMA Origin | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002
Innom A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020
Lt ARA Origin | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt ACA A1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ACA A2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Anterior Cerebral Artery A2 Segment;99SIEMENS;AntCerebAA2 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Inf Arcuate A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Mid Arcuate A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Sup Arcuate A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Vert Atlas Loop | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Bulb | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt CFA Pseudoaneurysm | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt CIA (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt CIA Aneurysm | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt CIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist ARA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist ATA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist Axillary A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist Brach A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist CCA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist CFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist CFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist CIA (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist CIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist CIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist Deep Brach A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist ECA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist EIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist EIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Dist LHA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist ICA-skull entrance | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist IIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Dist MCA M1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist MCA M2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ophthalmic A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Peroneal A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Pop A Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Dist PTA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist DFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist DFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Dist Radial A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Main RA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Subclav A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist SFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist SFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Dist T-P Trunk | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Ulnar A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist UE Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Dist Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt DPA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Dorsalis Pedis Artery;SRT;T-47741 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Epididymal A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Epididymal Artery;99SIEMENS;EpididymalA |  | Finding Site;SRT;G-C0E3=Testis;SRT;T-94000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt ECA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt EIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Inf Thyroid A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Inferior Thyroid Artery;SRT;T-46140 |  | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Inf Interlobar A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Mid Interlobar A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Sup Interlobar A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Interlobar Artery of Kidney;SRT;T-4667D |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt ICA Siphon | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt IIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Carotid Siphon;SRT;T-45308 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt LE Dist Anast 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Prox Anast 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Dist Graft 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Mid Graft 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Prox Graft 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Dist Inflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Mid Inflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Prox Inflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 1 Inflow Vessel;99SIEMENS;LEGraft1InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Dist Outflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Mid Outflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Prox Outflow Vessel 1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 1 Outflow Vessel;99SIEMENS;LEGraft1OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Dist Anast 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt LE Prox Anast 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt LE Dist Graft 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt LE Mid Graft 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt LE Prox Graft 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt LE Dist Inflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Mid Inflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Prox Inflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 2 Inflow Vessel;99SIEMENS;LEGraft2InVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt LE Dist Outflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt LE Mid Outflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt LE Prox Outflow Vessel 2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Lower Extremity Graft 2 Outflow Vessel;99SIEMENS;LEGraft2OutVes |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Mid ARA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ATA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Axillary A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Brach A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CCA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Mid CIA (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Mid CIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid CIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Mid Deep Brach A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ECA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid EIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Mid LHA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid ICA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid IIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Mid MCA M1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid MCA M2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ophthalmic A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Peroneal A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Pop A Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Mid PTA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid DFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid DFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Mid Radial A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Main RA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Subclav A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid SFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid SFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Mid T-P Trunk | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Ulnar A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid UE Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Mid Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Pop A Aneurysm | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Post-inj Cavernosal A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Lt PCA P1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Posterior Cerebral Artery P1 Segment;SRT;R-10253 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCA P2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Posterior Cerebral Artery P2 Segment;SRT;R-10255 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt PCoA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Posterior Communicating Artery;SRT;T-45320 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Pre-inj Cavernosal A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Cavernosal Artery;99SIEMENS;CavernosalA | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Lt DFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Prox ARA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ATA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Axillary A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Brach A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CCA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Lt Prox CIA (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Prox CIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox CIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Lt Prox Deep Brach A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ECA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox EIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Prox LHA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Left Branch of Hepatic Artery;SRT;T-46427 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox ICA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox IIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Lt Prox MCA M1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox MCA M2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ophthalmic A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Peroneal A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Pop A Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Lt Prox PTA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox DFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox DFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Lt Prox Radial A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Main RA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Subclav A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox SFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superficial Femoral Artery;SRT;T-47403 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox SFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Finding Site;SRT;G-C0E3=Superficial Femoral Artery;SRT;T-47403
Lt Prox T-P Trunk | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Tibioperoneal Trunk;SRT;T-4704A |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Ulnar A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Ulnar Artery;SRT;T-47200 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox UE Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Prox Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Main RA Origin (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Lt Main RA Origin | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Renal Artery;SRT;T-46600 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Lt RA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  |  |  | Finding Site;SRT;G-C0E3=Renal Artery;SRT;T-46600
Lt Inf Segmental A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Lt Mid Segmental A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Lt Sup Segmental A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Segmental Artery;SRT;T-46659 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Lt Subclav A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Subclavian Artery;SRT;T-46100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Sup Thyroid A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superior Thyroid Artery;SRT;T-45210 | Laterality;SRT;G-C171=Left;SRT;G-A101 | Finding Site;SRT;G-C0E3=Thyroid;SRT;T-B6000 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal ICA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Terminal internal carotid artery;SRT;R-102BD |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Terminal Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Terminal Vertebral Artery;99SIEMENS;TerminalVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Testicular A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Testicular Artery;SRT;T-46970 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt AVF Art Anast | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | AVF Arterial Anastomosis;99SIEMENS;AVFArtAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Dist AVF | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid AVF | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox AVF | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arteriovenous Fistula;SRT;M-39390 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist AVF Inflow A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid AVF Inflow A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox AVF Inflow A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | AVF Inflow Artery;99SIEMENS;AVFInflowArtery |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Dist AVF Outflow V | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt Mid AVF Outflow V | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt Prox AVF Outflow V | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | AVF Outflow Vein;99SIEMENS;AVFOutflowVein |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt AVF Venous Anast | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | AVF Venous Anastomosis;99SIEMENS;AVFVenAnast |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt UE Dist Anast | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal anastomosis;SRT;G-D872
Lt UE Prox Anast | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal anastomosis;SRT;G-D871
Lt UE Dist Graft | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, distal section;DCM;122104
Lt UE Mid Graft | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, mid section;DCM;122103
Lt UE Prox Graft | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Findings;DCM;121070 |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Graft to cited segment, proximal section;DCM;122102
Lt UE Dist Inflow Vessel | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Mid Inflow Vessel | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Prox Inflow Vessel | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft Inflow Vessel;99SIEMENS;GraftInVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt UE Dist Outflow Vessel | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Lt UE Mid Outflow Vessel | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Lt UE Prox Outflow Vessel | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Graft Outflow Vessel;99SIEMENS;GraftOutVessel |  | Finding Site;SRT;G-C0E3=Vascular Graft;SRT;T-D000F | Laterality;SRT;G-C171=Left;SRT;G-A101 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Lt Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Left;SRT;G-A101
Lt Vert A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Vertebral Artery;SRT;T-45700 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Left;SRT;G-A101
Mid Aorta (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Mid Aorta | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Ao Stent Graft | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Basilar A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid CHA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid IMA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Innom A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid Splenic A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SMA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188
Mid SMA, Postprandial | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Mid-longitudinal;SRT;G-A188 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Post-inj Dorsal A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Post-injection;99SIEMENS;Postinjection
Pre-inj Dorsal A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Dorsal Artery of Penis;SRT;T-46857 |  | Finding Site;SRT;G-C0E3=Penis;SRT;T-91000 |  |  |  |  | Injection State;99SIEMENS;InjectionState=Pre-injection;99SIEMENS;Preinjection
Prox Aorta (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Prox Aorta | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aorta;SRT;T-42000 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Ao Stent Graft | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aortic Stent Graft;99SIEMENS;AoStentGraft |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Basilar A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Basilar Artery;SRT;T-45800 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox CHA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Hepatic Artery;SRT;T-46421 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox IMA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Inferior Mesenteric Artery;SRT;T-46520 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Innom A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Innominate Artery;SRT;T-46010 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Splenic A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Splenic Artery;SRT;T-46460 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox Landing Site | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent Graft Landing Site;99SIEMENS;StGraftLandSite |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SMA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118
Prox SMA, Postprandial | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Superior Mesenteric Artery;SRT;T-46510 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Unilateral;SRT;G-A103 |  | Topographical modifier;SRT;G-A1F8=Proximal;SRT;G-A118 |  | Temporal period related to eating;SRT;R-41FFC=Post-prandial;SRT;G-A491
Rt ARA Origin | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Origin of vessel;SRT;G-036A
Rt ACA A1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt ACA A2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Anterior Cerebral Artery A1 Segment;99SIEMENS;AntCerebAA1 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Inf Arcuate A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Inferior;SRT;G-A115
Rt Mid Arcuate A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Medial;SRT;G-A109
Rt Sup Arcuate A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Arcuate Artery of the Kidney;SRT;T-4668A |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 | Vessel Branch;DCM;125101=Superior;SRT;G-A116
Rt Vert Atlas Loop | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Atlas Loop Vertebral Artery;99SIEMENS;AtlasLoopVertA |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Bulb | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Carotid Bulb;SRT;T-45170 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt CFA Pseudoaneurysm | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Pseudoaneurysm;SRT;M-32390 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt CIA (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710 | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt CIA Aneurysm | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Aneurysm;SRT;M-32200 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  |  |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt CIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100
Rt Dist ARA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Accessory Renal Artery;SRT;T-46640 |  | Finding Site;SRT;G-C0E3=Vascular Structure of Kidney;SRT;T-71019 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ATA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Anterior Tibial Artery;SRT;T-47700 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Axillary A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Axillary Artery;SRT;T-47100 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Brach A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Brachial Artery;SRT;T-47160 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CCA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Carotid Artery;SRT;T-45100 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Femoral Artery;SRT;T-47400 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Femoral Artery;SRT;T-47400
Rt Dist CIA (Post Repair) | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Temporal periods Relating to Procedure;SRT;R-40FB8=After Procedure;SRT;R-422A4
Rt Dist CIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Common Iliac Artery;SRT;T-46710 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist CIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Common Iliac Artery;SRT;T-46710
Rt Dist Deep Brach A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Deep Brachial Artery;99SIEMENS;DeepBrachialA |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ECA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Carotid Artery;SRT;T-45200 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | External Iliac Artery;SRT;T-46910 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist EIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=External Iliac Artery;SRT;T-46910
Dist RHA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Right Branch of Hepatic Artery;SRT;T-46423 |  | Finding Site;SRT;G-C0E3=Artery of Abdomen;SRT;T-46002 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Carotid Artery;SRT;T-45300 |  | Finding Site;SRT;G-C0E3=Artery of neck;SRT;T-45005 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist ICA-skull entrance | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Carotid Artery, entrance to skull;99SIEMENS;ICASkullEntr |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 |  |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Internal Iliac Artery;SRT;T-46740 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist IIA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Internal Iliac Artery;SRT;T-46740
Rt Dist MCA M1 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Middle Cerebral Artery M1 Segment;SRT;R-1024F |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist MCA M2 | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Middle Cerebral Artery M2 Segment;SRT;R-10251 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Ophthalmic A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Ophthalmic Artery;SRT;T-45400 |  | Finding Site;SRT;G-C0E3=Blood Vessel of Head;SRT;T-40501 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Peroneal A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Peroneal Artery;SRT;T-47630 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Popliteal Artery;SRT;T-47500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist Pop A Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Popliteal Artery;SRT;T-47500
Rt Dist PTA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Posterior Tibial Artery;SRT;T-47600 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist DFA | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Profunda Femoris Artery;SRT;T-47440 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119
Rt Dist DFA Stent | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Stent;SRT;A-25500 |  | Finding Site;SRT;G-C0E3=Artery of Lower Extremity;SRT;T-47040 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119 |  | Finding Site;SRT;G-C0E3=Profunda Femoris Artery;SRT;T-47440
Rt Dist Radial A | Minimum Diastolic Velocity | MD | cm/s m/s | 5 | 2 | Minimum Diastolic Velocity;LN;11665-7 | NUM | Radial Artery;SRT;T-47300 |  | Finding Site;SRT;G-C0E3=Artery of Upper Extremity;SRT;T-47020 | Laterality;SRT;G-C171=Right;SRT;G-A100 |  | Topographical modifier;SRT;G-A1F8=Distal;SRT;G-A119